|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| **S.No.** | **Name Of Equipment** | **Qty** | **C.I.F Price** | | **F.O.R Price** | |
|  |  |  | **Unit price** | **Total Price** | **Unit Price** | **Total Price** |
| **RADIOLOGY** | | | | | | |
|  | **Fluoroscopy Machine 800mA, three phase 380 volts or more complete set**  **with all standard accessories for its installation and commissioning:**  **X-Ray Generator 80kW, three phase , 380 V or more with**  **Single tube for radiography and fluoroscopy. Remote control x-ray table**  **With integrated x-ray tube column stand, image intensifier TV system with operator console.**  **X-Ray Generator**  Microprocessor controlled Generator with touch screen console for radiography and fluoroscopy  Maximum 80 kW or more  X-ray tube voltage 40~150 kV( 1 kv steps )  X-Ray tube current 10~800mA  Exposure time from 0.001 to 10 sec.  High speed starter (9000 rpm)  **Flouoroscopy**  Continuous mA range 0.5 mA up to 4 mA or more (low dose)  up to 8mA (high dose) with auto flouro  Fluoroscopy kV : 40 ~ 120kV  Timer : 5~ 10 minutes  **R&F Remote Controlled Tilting Table**  Remote controlled table with advanced features and micro-processor controlled  Main standard table control console  Motorized lateral movements  Spot filling device or serial changer  Transversal travel : +/- 12cm or more with speed of 4 cm/sec or better  longitudinal movement of table top 45cm or better with speed of 4cm/sec or better  Table tilting speed of 4 degree/sec or better  Maximum tilt angle: + 90° or more  Maximum Trendenlenburg angle : - 15° or more  **Automatic spot film device**  Motorized longitudinal movements of 90cm or more  All Cassettes size  Automatic format detection 1 – 2 – 3 – 4 division  Moving grid ratio 8:1 or better  **Image Intensifier**  Image Intensifier 9’’ (23 cm) 3 field, 9/6/4.5”  Resolution central 48pl/cm  **CCD Camera**  Camera with last image hold  2 nos. LCD TV Monitor 18" or more Medical quality  Trolley for monitors  **X-Ray Tube**  Dual focus 0.6mm/1.2mm - 9000 rpm - 300 kHUor more  **Compression Device**  Maximum compression force : 80N  Angulation operative range : +/- 20° or +/- 5° from starting angl  **AUTOMATIC COLLIMATOR**  **ACCESSORIES:**  All standard accessories including, Lead glass two pair of lead gloves, two Aprons, two goggles,  two thyroid shield, Two gonadal shields,4x1 two local made illuminators and one auto processor  with a set of chemical and with two cassettes of each size including 14x17, 14,14,12x15,10x12,  8x10 with best screen film combination should be provided.  One voltage stabilizer/step up transformer to support the whole system on full load.  Phase reversal devise for fluoroscopy machine will be responsibility of the vendor  Country of originshould be USA/western Europe/Japan withtheirrespectivestandards MDD (CE) compliance/MHLW Japan.  Complete installation shall be responsibility of vendor as per PNRA standards.  *Warranty period shall be of three years starting from date of installation / commissioning* | 1 |  |  |  |  |
|  | **Cassettes with Screen various sized**  08x10”  10x12”  12x15”  14x14”  14x17”  Make: UK/USA/EEC/Japan  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 02  Nos. |  |  |  |  |
|  | **Mobile X-Ray Machine 100mA complete set with all standard accessories for its installation and commissioning:**  High frequency X-Ray inverter type Generator  Having output power of 4kW or more  X-Ray tube 50 KHU or more  KVP range: 40-110kV or more  MA Range: 25-100mA or better,  mAS range: 1.0-200mAs or better  Centering light indicator  Standard power supply cable, standard x-ray exposure switch cable,  Vertical movement of the arm to lower or raise the tube collimator assembly, Rotation of the Collimator with reference to the power module (+/- 90deg) or better.  Country of originshould be USA/western Europe/Japan withtheirrespectivestandards MDD (CE) compliance/MHLW Japan  *Warranty period shall be of three years starting from date of installation / commissioning* | 01 |  |  |  |  |
|  | **X-Ray Illuminator Single**  Illuminator for viewing of x-ray firm  Size: 8x10 ~ 14x17”  Single view acrylic sheet of 2mm.  Fluent tube light with heavy duty switch.  Local/imported  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 108 nos. |  |  |  |  |
|  | **Ultrasound Machine gray scale (black & white)**  High Resolution Portable B/W Ultrasound Machine having Digital Beam Forming System for Abdominal/Obstetrical/Gynae/Urology/small parts with the following Composition:  15” to 12’’ or more TFT LCD monitor having wide viewing.  Choice upto four(04) transducer  STC gain six steps or more  B-mode imaging,harmonic imaging  Depth: 24 cm or more.  B/M scan, B (Single, Dual), B/M, M  3~5 MHz Multi Frequency Convex Probe.  Linear probe 5~10MHz (for small parts)  Comprehensive measurement package/functions for abdomen and obs/Gynae.  Built in battery.  CD/DVD writer or USB port.  B/W Ultrasound printer.  Voltage Stabilizer.  Mobile Trolley  Country of originshould be USA/western Europe/Japan withtheirrespectivestandards MDD (CE) compliance/MHLW Japan  *Warranty period shall be of three years starting from date of installation / commissioning* | 03  nos |  |  |  |  |
|  | **Mammography**  **Mammography Machine film based, complete set with all standard accessories**  **for its installation and commissioning:**  **Film based Mammography system, complete with Microprocessor controlled**  **X-Ray Generator, Automatic voltage compensation. X-Ray tube,**  **control panel ,collimator , compression device ,C-Arm , bucky and**  **Cassettes size: 18 x 24 and 24 x 30cm.**  High frequency microprocessor controlled generator.  X-ray tube voltage: between 20-35 kV with 1 kV step.  Tube current: 35-110mA.  mAs : 10-600mAs or better  Automatic Exposure Control  X-Ray Tube  High Speed Bi-angular, rotating anode.  Small focus: 0.1 mm, large focus: 0.3mm,  Anode head storage capacity: 300KHU or more.  Dual motorized filters: 30 μm Mo and 0.5 mm.  25 μm Rh with automatic change.  Motorized diaphragms  Built-in spot collimator.  C-Arm  SID 65 cm or better.  Single or Dual control panels, one on each side of the unit.  C-arm Angulation from -140° to +180 ° or better.  C-arm vertical travel from 80 to 130 cm or better.  Digital display of projection angle.  Motorized telescopic column with two adjustable speeds.  Motorized beam collimation.  LED/Halogen imaging field illumination.  Compression  Digital display of compression force and breast thickness.  Adjustable compression force limit and digital display.  Motorized compression with selectable speed levels.  Bucky  Manual cassette loading and unloading.  Grid movement should be microprocessor controlled with a grid ratio of at least 5:1  or better for a good contract film.  Complete with following Accessories:   1. Cassettes 18 x 24cm 2 Nos. 2. Cassettes 24 x 30cm 2 Nos. 3. Compatible Stabilizer 4. LCD type X-Ray film viewer to view four films at one time. 5. One automatic dedicated mammography film processor with set of chemicals 6. Lead glass for protection 7. Patient stool/chair.   Country of originshould be USA/western Europe/Japan withtheirrespectivestandards MDD (CE) compliance/MHLW Japan  Complete installation shall be responsibility of vendor as per PNRA standards.  *Warranty period shall be of three years starting from date of installation / commissioning* | 1 |  |  |  |  |
| **LABORATORY** | | | | | | |
|  | **Hematology Analyzer**  3 parts differential  18 or more hematological measured parameters . WBC, RBC, Hb, MCV and PLT.  Colored display .  -WBC-LYM#-MID#- GRN#LYM%-GRN%- HGB-RBC-HCT-MCVRDW-  MCH-MCHCPLT- MPV-PCT-PDW-WBC, RBC and PLT histogram  Alarm system with auto diagnostic capabilities  72 or more samples/ hour capacity .  Blood sample volume upto10ul  CBC test should incorporate within one minute & minimum quantity should be used  RS232 interface  Make: UK/USA/EEC/Japan  *Warranty period shall be of five years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | **Blood Bank Refrigerator**  Capacity 200 Liters or more / 100 Bags/450/500ml  Cabinet to have smooth stainless steel interior.  Microprocessor control panel with digital display.  Integrated Data log system.  04/04 shelves. Refrigerant Gas R404a  Lockable doors automatic defrosting  Interior lighting. Data logger  Audible & Visual alarms.  Door: Single Glass Door  Country of origin should be UK/USA/ EEC./Japan/  Who approved CE & TUV  Installation Should be 100 or more units.  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 01 |  |  |  |  |
|  | **PCR Machine**  Thermal block format 24 well, 96 well test tubes of 0.2 ml  Uniformity of thermal block temperature +/- 0.3oC  Excitation source Light-emitting diode (LED).  Fully Automated & motorized Function  Number of the 5 fluorescence.  Temp Range: OoC-99.9oC  Temp Increment Range: 0.1oC-9.9oC  Max Ramp Rate: 4.5oC or more  Make: UK/USA/EEC/Japan  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 01 |  |  |  |  |
|  | **Precsiosn Calibration Machine with Shaker**  Make: UK/USA/EEC/Japan  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 01 |  |  |  |  |
|  | **Hot Air Oven**  Hot Air Ovens for general purposes applications with double wall and fresh air inlet.  Microprocessor Control  LED digital display for monitoring  Automatic over-temperature safety device  Digital timer. 99 hrs 59 min  Temperature: more than 225 deg.  Capacity: 86 litres or more.  Complete with shelves.  Must be CE Marked  Country of origin should be UK/USA/ EEC./Japan/  .  Installation Should be 100 or more units.  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 2 |  |  |  |  |
|  | **Centrifuge machine 24 head**  Complete with 24-place angle rotor.  Microprocessor-controlled.  Head: 24 nos.  Max g-force:17,000 xg  Max rpm: 13,000  Noise Level: Less than 55 dBA  Timer LED digital, 99 min LED  Brake  Locking latch  Safety shutout  Country of origin should be UK/USA/EEC/ Japan  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 03 |  |  |  |  |
|  | **Ultrasonic water bath**  Built-in stirrer . Microprocessor control  Volume: 22 Liters or better, Uniformity: 5 ºC, Heater: 1 KW  Thermometer readout , Temperature Controller  Must be CE Marked  Make: UK/USA/ EEC/Japan  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 02 |  |  |  |  |
|  | **Autoclave Verticle**  40 liters or more, manual control system. Eco friendly and economically design with over temperature, pressure & elec. Over load protection with emergency exhaust switch with low water buzzer with over heat/temp. indication light chamber. Temperature upto 121oC;Digital PID Controller with Timer and Tunning:  CE Marked  Make: UK/USA/EEC/Japan  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 01 |  |  |  |  |
| **OPERATION THREATRE** | | | | | | |
|  | **Dressing set**  Forceps, dissecting, Straight, 1/2 teeth, 130 mm x 2  Forceps, sinus, 180mm x 1  Forceps, Dissecting, Straight, Plain, 115 mm x 1  Forceps, Sponge holding, 200mm x 1  Scissors, Mayo, Straight, 150mm x1  Scissors, dressing, Straight, 165mm x 1  Scissors, ligature, spencer, 130mm x 1  Bowls, lotion, s/s,         5 liter x 1  Bowl lotion, s/s, 0.2 liter x 2  Kidney Dish, 24cm x 2  Instrument container, s/s with cover, 300x200x50mm x 1  Made from imported stainless steel, Rust resistant and stain finish  Export Quality Material  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 17 |  |  |  |  |
|  | **Portable LED OT light**  Mobile spot Light  Minimum 35,000LUX/0.5m or more  Color temperature: 4000K or more better  Flexible arm with up & down movement and 360 degree rotation.    **Make: UK/USA/EEC/JAPAN**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 6 |  |  |  |  |
|  | **Anesthesia Machine with built in Ventilator**  Specifications:-  Patient type: neonates, Peads& adult  VENTILATION MODES:  The modes of ventilation should include: Volume Ventilation (CMV, Assist CMV), Pressure Support Ventilation (PSV), SIMV, SPONT, MANUAL,SIGH, Apnoea BACK-UP, CPAP, Bilevel, & NIV. Manual breath lock and inspiratory/expiratory hold.  TECHNICAL SPECIFICATIONS:  The patient range should be from neonate, pediatric and adults with easy operation.  Breaths Rate for CMV+, PCV+, P – SIMV+ should be from at least 6 bpm to 140 bpm  PEEP & CPAP from 0 cmH2O up to 35 cmH2O  Tidal Volume from minimum 2 ml to 2000 ml for Volume mode  Inspiratory minute Volume  Expiratory minute volume  The I: E ratio: 1:4 - 4:1  ALARMS:  Operator adjustable alarms should include low/high minute volume, high pressure, low/high tidal volume, apnea rate  Oxygen concentration, disconnection, loss PEEP  Apnoea Alarm  Display:  Screen size not less than 12” Inch or more TFT Color Screen with Loops, Graph & Monitoring on the same window  EVENT LOG:  Memory Storage up to 100 Events  Power Supply:  100-230V  Internal Battery Operation 1 Hour or more  Supplied should be with supporting arm, trolley, breating circuits for adult, paeds and neonates , power cord and operating manual  **Medical Air Compressor or Turbine system**  CE & IEC Certified.  Make :UK/USA /Europe  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 02 |  |  |  |  |
|  | **Basic instrument set**  Metal Ruler 6’’ (1 Nos)  Scalpel Handle #3 (1 Nos)  Scalpel Handle#4 (1 Nos)  TC Wire Cutting Scissors Angled 4 3/4’’ (1 Nos)  Rochester-OchsnerFoersepsCvd 5 ½’’ (1 Nos)  Allies Tenaculum Forceps Delicate 5 ¾’’ (2 Nos)  TC Derf NH Serr 4 ¾’’ (1 Nos)  TC-Crile-Wood NH NarroeSerr 6’’ (1 Nos)  TC Ryder Micro NH Serr 6’’ (1 Nos)  Stevens Tenotomy Scissors Delicate S/S 4 ½’’ (1 Nos)  Stevens Tenotomy Scissors Delicate B/B 4 ½’’ (1 Nos)  Metzenbaum Scissors Cvd 4 ½’’ (1 NOs)  Reynolds Sciss Nar Tip Cvd 6 (1 Nos)  Semkin Dressing Forceps Delicate 5’’ (1 Nos)  Semkin Tissue Forceps 1x2 Teeth Delicate 5’’ (2 Nos)  Miller-Senn Retractor Double End Sharp 6’’ (1 Nos)  Iris Scissors str 4 1/8’’ (1 Nos)  Iris Scissors Cvd 4 1/8’’ (1 Nos)  Mayo Scissors Str 5 ½’’ (1 Nos)  Mayo Scissors Cvd 5 ½’’ (1 Nos)  Metzenbaum Scissors Cvd 6’’ (1 Nos)  Hartmann Mosquito Forceps Str 3 1/2 ‘’ (3 Nos)  Hartmann Mosquito Forceps Cvd 3 1/2 ‘’ (3 Nos)  Halsted Mosquito Forceps Str 5’’ (6 Nos)  Halsted Mosquito Forceps Cvd 5’’ (6 Nos)  Adson Dressing Forceps Serr 4 ¾’’ (1 Nos)  Adson Tissue Forceps Slide grasping 4 ¾’’ ( 1 Nos)  Allis Forceps 4x5 Teeth 6’’ (1 NOS)  Extra Delicate Allis Forceps 4x5 Teeth 5’’ 91 Nos)  Freer Dissector Double End S/B 7 (1 Nos)  Joseph Hook 1 Prong Sharp 6 ¼’’ (2 Nos)  Backhaus Towel Clamp 3 ½’’ (2 Nos)  Backhaus Towel Clamp 5 1/4’’ (6 Nos)  Cushing Nerve and Vein retractor 9 (1 Nos)  Frazier Suction Tube 7 French (1 Nos)  Frazier Suction Tube 8 French (1 Nos)  Crile Retractor D/E 4 ½’’ (1 Nos)  US Army Retractor (Set of 2) 8 1/2 ‘’ (1 Nos)  Senn retractor Double End Sharp 6 ¼’’ (1 Nos)  Ballenger Sponge Forceps StrSerr 7’’ (2 Nos)  Ballenger Sponge forceps CvdSerr 7’’ (2 Nos)  Foerster Sponge Forceps StrSerr 9 ½’’ (2 Nos)  Made from imported stainless steel, Rust resistant and satin finish  Export Quality Material  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 04 |  |  |  |  |
|  | **Hysterotactomy set**  Made from imported stainless steel, Rust resistant and Satin finish  Export Quality Material  Must be CE Marked Product  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 4 |  |  |  |  |
|  | **LSCS set**  Tray kidney shape with Lid S/S 200x30mm (02 Nos)  Tray instrument with Lid 225x125x50mm Deep (02 Nos)  Forceps Artery Spencer Wells Straight BJ.S/S 125mm (12 Nos)  Forceps Artery Spencer Wells Straight BJ.S/S 200mm (02 Nos)  Forceps Dissecting Plain serrated jaws S/S 150 mm ( 02 Nos)  Forceps Dissecting Treves, 1x2 Teeth 150 mm (02 Nos)  Forceps Tissue Allis fine 4x5 Teeth BJ.S/S 150 mm (02 Nos)  Forceps Tissues little wood 2x3 Teeth (bj).s/S 190 mm (03 Nos)  Forceps Towel Holding cross Action S/S 100mm (06 Nos)  Forceps Sponge Holding Rampley BJ S/S 240 mm (04 Nos)  Handle size 3 for Scalple blades size 10-15 (01 No)  Handle size 4 for Scalple blades size 20-24 (01 No)  Needle Holder Mayo HegarBj.S/S 180mm z(02 Nos)  Scissor Mayo Straight SJ. S/S 200mm. (01 No)  Scissor Mc Indoe/Metzenbaum CO Flat SJ.S/S 190mm (01 No)  Retr. Doyen Ring Handle blade 95 mm Length 280 mm (01 No)  Retr.Lange beck H/H 13x44 Blade, 210 mm Long (01 No)  Retr. Morris Double-ended (52x38) & (5264) mm (01 No)  Pin Safety Mayo, forceps holding, 120mm (03 Nos0  Suction Tube Pool’s S/S 225mm (01 No)  Forceps Haemostasis Green Armitage 200 mm (06 Nos)  Sterilizing and storing case St. St. with filters 260x150x50 mm (02 Nos)  Made from imported stainless steel, Rust resistant and satin finish  Export Quality Material  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 6 |  |  |  |  |
|  | **DNC set**  Scalpel Handle #3 x 1  Scalpel Handle #7 x 1  Scalpel Handle #3L x 1  Crile Forceps Str 6 1/4" x 6  Crile Forceps Cvd 6 1/4" x 6  Dressing Forceps 8" x 1  Tissue Forceps 1 x 2 Teeth 8" x 1  TC Crile-Murray NH Serr 6" x 1  TC Mayo-Hegar NH Serr 7"x 1  Backhaus Towel Clamp 5 1/4" x 6  Bozeman Forceps Double Curved 10" x 1  Foerster Sponge Forceps CvdSerr 9 1/2" x 2  Jackson Retractor Small Blade 3" x 1 1/2" x 1  Jackson Retractor Medium Blade 3 1/2" x 1 1/2" x 1  Auvard Weighted Speculum Medium 2 1/2 lbs. x 1  Graves Vaginal Speculum Medium 1 1/4" x 4"x 1  Graves Vaginal Speculum Large 1 3/8" x 4 1/2" x 1  Pederson Vaginal Speculum Small 1/2" x 3" x 1  KoganEndospeculum w/ Gauge and Screws x 1  Goodell Dilator Large 13" x 1  Hegar Dilators (Set of 8) x 1  Sims Sound Plain 12 1/2" x 1  Sims Curette Sharp 11" #2 x 1  Sims Curette Sharp 11" #3 x 1  Sims Curette Sharp 11" #4 x 1  Sims Curette Sharp 11" #5x 1  Kevorkian-YoungeEndocervical Curette w/ Basket 12"x 1  Schroeder Tenaculum Forceps 9 1/2" x1  Javerts Forceps Serr 9 1/2"x 1  Made from imported stainless steel, Rust resistant and Satin finish  Export Quality Material  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 4 |  |  |  |  |
|  | **DNE set**  Scalpel Handle #3 x 1  Scalpel Handle #7 x 1  Scalpel Handle #3L x 1  Crile Forceps Str 6 1/4" x 6  Crile Forceps Cvd 6 1/4" x 6  Dressing Forceps 8" x 1  Tissue Forceps 1 x 2 Teeth 8" x 1  TC Crile-Murray NH Serr 6" x 1  TC Mayo-Hegar NH Serr 7"x 1  Backhaus Towel Clamp 5 1/4" x 6  Bozeman Forceps Double Curved 10" x 1  Foerster Sponge Forceps CvdSerr 9 1/2" x 2  Jackson Retractor Small Blade 3" x 1 1/2" x 1  Jackson Retractor Medium Blade 3 1/2" x 1 1/2" x 1  Auvard Weighted Speculum Medium 2 1/2 lbs. x 1  Graves Vaginal Speculum Medium 1 1/4" x 4"x 1  Graves Vaginal Speculum Large 1 3/8" x 4 1/2" x 1  Pederson Vaginal Speculum Small 1/2" x 3" x 1  KoganEndospeculum w/ Gauge and Screws x 1  Goodell Dilator Large 13" x 1  Hegar Dilators (Set of 8) x 1  Sims Sound Plain 12 1/2" x 1  Sims Curette Sharp 11" #2 x 1  Sims Curette Sharp 11" #3 x 1  Sims Curette Sharp 11" #4 x 1  Sims Curette Sharp 11" #5x 1  Kevorkian-YoungeEndocervical Curette w/ Basket 12"x 1  Schroeder Tenaculum Forceps 9 1/2" x1  Javerts Forceps Serr 9 1/2"x 1  Made from imported stainless steel, Rust resistant and Satin finish  Export Quality Material  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 04 |  |  |  |  |
|  | **General Surgery Set Complete in All Respect**  Bowl Lotion 250 mm (2 Nos)  Gallipot/Cup S/S 180 ml (2 Nos)  Gallipot/Cup S/S 250 ml (2 Nos)  Tray Kidney shape with Lid S/S 200x30 mm (2 Nos)  Tray Kidney shape with Lid S/S 300x50 mm (2 Nos)  Forceps artery Kocher COF BJ.S/S 180 mm (20 Nos)  Forceps artery Spencer Wells CO Flat (bj). S/S 200 mm (15 Nos)  Forceps Dissecting Plain serrated jaws S/S 150 mm (6 Nos)  Forceps Dissecting Plain serrated jaws S/S 250 mm (1 Nos)  Forceps Dissecting narrow , serrated jaws S/S 200 mm (1 Nos)  Forceps Dissecting Lane’s 2x3 teeth 180 mm (3 Nos)  Forceps Dissecting Waugh, fine serrated jaws 1x2 Tth 200 mm (1 Nos)  Forceps Tissue Allis fine 4x5 Teeth BJ.S/S 150 mm (4 Nos)  Forceps Towel Holding cross Action S/S 100 mm (8 Nos)  Forceps Towel Moynihan, Double Teeth (bj) 190 mm (4 Nos)  Forceps Sponge Holding Rampley BJ S/S 240 mm (5 Nos)  Handle size 3 for scalpel blades size 10-15 (1 Nos)  Handle size 4 Scalple blades size 20-24 (1 Nos)  Needle Holder Kilner’s 165 mm (1 Nos)  Needle Holder Mayo Hegar BJ.S/S 180mm (2 Nos)  Scissor Blunt Point Straight SJ. S/S 180mm (1 Nos)  Scissors Mayo’s CO Flat 200 mm S/S (1 Nos)  Scissors Mayo’s Straight 200 mm S/S (3 Nos)  Scissors Mclndoe’s CO Flat 190 mm S/S (1 Nos)  Retr. Deaver #1-Blade 25mm, Length 305mm (1 Nos)  Retr. Lange beck, H/H 26x64 Blade, 210 mm Long (2 Nos)  Retr. . Deaver #5-Blade 75mm, Length 305mm (1 Nos)  Retr. Morris Double-ended (52x38) & (52x64) mm (2 Nos)  Retr. Walton, malleable copper 13mm length 305mm (1 Nos)  Needle Aneurysm S/S. Small 146 mm (1 Nos)  Pin Safety Mayo, forceps holding, 120mm (4 Nos)  Suction Tube Mclndoe (1 Nos)  Suction Tube Poole’s S/S 225 mm  Sterilizing and storing case St.St. with filters 260x150x50 mm (2 Nos)  Wire Cutting, Tungsten. Carbide inserts 180 mm.  Made from imported stainless steel, Rust resistant and Satin finish  Export Quality Material  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 4 |  |  |  |  |
|  | **Instrument Trolley (Fixed)**  Stainless steel trolley with plastic drawer internally divided in 4 sectors. 2 shelves with up stands on four sides. Revolving rubber castors Ø 80 mm (two with brake).  Made from imported stainless steel, Rust resistant and Satin finish  Export Quality Material  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 13 |  |  |  |  |
|  | **Instrument Trolley Adjustable**  With four drawers  Made from imported stainless steel, Rust resistant and Satin finish  Export Quality Material  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 9 |  |  |  |  |
|  | **Surgical Drum Stainless steel**  Drum, domed top and bottom with sliding removable band and drop down-hasps  Made from imported stainless steel, Rust resistant and Satin finish  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 84 |  |  |  |  |
|  | **Instrument Tray SS**  Set of different size of instrument trays  Made from imported stainless steel, Rust resistant and Satin finish  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 11 |  |  |  |  |
|  | **Delivery table**  Three sections delivery table with high density padding, covered with washable, fire proof, leatherette material.  Adjustable Back-rest and leg-rest section  Complete with 2 foam leg holders and stainless steel bowl Ø 32 cm.  Maximum load: 130 kg  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 2 |  |  |  |  |
|  | **Normal delivery set SS**  Sponge holidingforcepe x 2  Cord clamp  Straight big artery forcepe x 1  Dissecting forceps tooth x 1  Dissecting forceps non tooth x 1  Scissor mayo x 1  Scissor cord cutting x 1  Scissor episiotomy x 1  Outlet forceps x 1  Low mid cavity forceps x 1  Silicon cups, 40mm, 50mm and 60mm  Complete with cover bag  Made from imported stainless steel, Rust resistant and Satin finish  Export Quality Material  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 6 |  |  |  |  |
|  | **Epiestomy set**  Episiotomy Set,  1 Dressing Forceps, 14.5 cm - 5 3/4",  1 Tissue Forcepse, 14.5 cm - 5 3/4",  - 1 MATHIEU Needle-Holder, 17 cm - 6 3/4",  - 1 BRAUN Episiotomy Scissors, 15 cm - 6",  - 1 Tissue Scissors, 14.5 cm - 5 3/4", 03.06.14 - 1 KOCHER Hemostatic Forceps, 14 cm - 5 1/2",  - 1 Needle Case, rectangular, 46.37.47 - 6 Suture Needles, in assorted sizes  Export Quality  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 4 |  |  |  |  |
|  | **Vacuum extractor**  Vacuum Extractor with various sizes suction cups.  4 antistatic castors , 2 with brakes.  Piston type, oil free.  Pump flow: 40L/min or more.  overflow prevention.  Hydrophobic PTFE filter.  3 m PVC Suction tube.  Metallic handle, Metallic curved and straight catheter .  Make: UK/USA/EEC/Japan.  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 2 |  |  |  |  |
|  | **CTG Machine**  Dual Channel, 5 inch or more display  Trace annotation facility  automatic fetal movement measurement, Automatic detection of all sensor,  Patient data saving facility  Integrated battery,  USB Connectivity,  Built-in Thermal printer.  Complete unit with 2 ultrasound probes,  Make: UK/USA/EEC/Japan  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 01 |  |  |  |  |
| **EMERGENCY/ICU/CCU** | | | | | | |
|  | **Infusion Pump**  Volumetric peristaltic infusion pump. Flow rate 0.1 to 99.9 ml/h. Unit compatible with most IV set. Memory function incase of power failure to retain the most recent flow rate & total volume. Comprehensive alarm for Air in line, occlusion, door open, low battery, infusion complete, empty container. Built in battery charger. Unit complete with IV sets.  Make: UK/USA/EEC/Japan.  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 14 |  |  |  |  |
|  | **Suction Machine heavy duty**  Suction Machine with two polycarbonate collection Jars of 2000ml each  4 antistatic castors , 2 with brakes.  Piston type, oil free.  Pump flow: Minimum 60L/min and more.  overflow prevention.  Hydrophobic PTFE filter.  3 m PVC Suction tube.  Metallic handle, Metallic curved and straight catheter .  Make: UK/USA/EEC/Japan.  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 04 |  |  |  |  |
|  | **Emergency Resuscitation Trolley Fully Equipped**  Emergency trolley - 3 drawers or more  Structure and drawers in epoxy powder coated steel with antibacterial and antimicrobial treatment, to improve Infection Control  Walls moulded strong polymer  Ample worktop with moulded 4-sides raised rim and ergonomic built-in push handle, scratch resistant  Extendable writing surface  Equipped with:  3 drawers in painted steel mounted on steel ball-bearing fully-extending rails, complete with handles with label holder for quick identification of content:  1 drawer with  Central locking system by knob with security breakable plastic seals.  Upper part:  S/S IV pole holder with 2 S/S hooks, height adjustable  One (R/L) side:  2 transparent tilt-out bins  Shelf for vacuum pump  Back side:  Oxygen tank compartment with holding straps, from 3 to 7 liters  Storage space for catheters / probes  One Side (R/L):  Waste basket without lid, with removable internal  **Defibrillator**  with all Standard Accessories  **Specifications:**  Biphasic technology  Defibrillator biphasic minimum 270 Joules atleast(***Mandatory***)  Charging time: Less than 10 sec for 270 J  ECG­:-6 Channel ECG  Leads: I, II, III, aVR, aVL, aVF  Heart rate: 50-300 per min  Display: TFT LCD 5’ Inches or More  Built-in Printer with Memo Print  Data management:-  Removable storage unit: 01GB or more with voice recording up to 80 Hours or more  Rechargeable external battery with minimum 04years shelf life  Monitoring time 03Hours or more Extendable  Complete with all standard Accessories  External Paddles for Adult & Built-in Pediatrics Paddles  ECG Leads minimum 05 cable set  Ultrasound Gel  Power Cord  (References required from reputable institute like SIUT,Civil,JPMC,Indus,Dow,Tabba,AKUH)  Installation must be more than 50 units atleast  Make: UK/USA/EEC/Japan  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 06  06 |  |  |  |  |
|  | **Glucoemeter**  Measurement Range 20 to 600 mg/dL or 1.1 to 33.3 mmol/L, Sample Sites Finger tips, Forearm or palm  Sample size 1 µL, Test Time Approximately 10 seconds, Power source One 2032.3.0V coin cell battery, Battery Life 12 months or approximately 1,000 tests, Memory plus Up to 300 records with time and date EZ Last 30 results, no averaging, no date/time  Make: UK/USA/EEC/Japan.  Or CE Approved  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 20 |  |  |  |  |
|  | **Ambu Bag Adult with Manometer**  100% latex-free Ambu bags  Built-in pressure limitation  Suitable for adults  with face mask and single shutter valve.  Autoclavable at 134°C.  Make: UK/USA,EEC,Japan  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 12 |  |  |  |  |
|  | **Ambu bag paeds with Manometer**  100% latex-free Ambu bags  Built-in pressure limitation  Suitable for paeds  with face mask and single shutter valve.  Autoclavable at 134°C.  Make: UK/USA,EEC,Japan  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 12 |  |  |  |  |
|  | **Emergency Resuscitation trolley**  Resuscitation trolley with 4 swivel castors with brakes.  Having 4 drawers or more.  Top tray for defibrillator  Side tray for instrument  Complete with adult and child resuscitator with face mask.  Venture suction with collection jar  Container for waste.  Drip stand support  Cardiac board  Oxygen regulator  Flow meter with humidifier  Nebulizer with kit  Make: UK/USA/EEC/Japan  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 19 |  |  |  |  |
|  | **Cardiac Monitor**  Multi Parameter monitor  Colour 12" TFT Monitor  Parameters:  (ECG/RESP - NIBP - SPO2 - 2TEMP - 2IBP)  ECG:  Selectable 3 to 5 wires  Heart rate range: 25-250 bpm  Defibrillator and electro surgery protection  SpO2:  Pulse rate range: 25 to 250bpm  Temperature:  Range 0-50 ° C  Respiration:  Breath rate range: 6-150 BrPM  IBP:  Sensor Impedance: 300-2500Ohm  - ST segment and arrhythmia analysis  - Internal battery  Standard supplied accessories:  **Make: UK/USA/Europe**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 14 |  |  |  |  |
| **OPD/WARD** | | | | | | |
|  | **Drip Stand**  **Stainless steel, double side hooks**  **With castors**  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 154 |  |  |  |  |
|  | **Baby Weighing Scale**  Digital scale with extra large LCD display (65 mm).  Strong and durable  steel comfortable platform.  **Should be CE Approved**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 2 |  |  |  |  |
|  | **Weighting Machine**  Digital scale with extra large LCD display (65 mm).  Strong and durable  steel comfortable platform.  **Should be CE Approved**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 06 |  |  |  |  |
|  | **Portable BP Apparatus**  with fabric cuff with Velcro closure, reliable manometer with clip to be fixed on the cuff.  Scale 0-300 mmHg. Air release at closed tap max 4 mm Hg/minute. Manual setting of deflation possible up to 2/3 mm Hg.  Make: UK/USA/EEC/Japan  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 12 |  |  |  |  |
|  | **Medical Trolley**  Stainless steel trolley with plastic drawer internally divided in 4 sectors. 2 shelves with up stands on four sides. Revolving rubber castors Ø 80 mm (two with brake).  Made from imported stainless steel, Rust resistant and Satin finish  Export Quality Material  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | **Wall Mounted BP apparatus**  with fabric cuff with Velcro closure, reliable manometer with clip to be fixed on the cuff.  Scale 0-300 mmHg. Air release at closed tap max 4 mm Hg/minute. Manual setting of deflation possible up to 2/3 mm Hg. With wall stand  Make: UK/USA/EEC/Japan  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 12 |  |  |  |  |
| **NEPHROLOGY** | | | | | | |
|  | **RO Plant**   |  |  |  |  |  | | --- | --- | --- | --- | --- | | **Specifications:** | |  |  |  | | 1. Operating Pressure: 100~250 PSI | | |  |  | | 2. Salt Rejection: 99% (under TDS 2000 PPM) at 25C\* | | |  |  | | 3. Recovery Rate: 70% | |  |  |  | | 4. Product 10 liter per minute | | |  |  | | 5. System should operate on 220VAC / 50Hz | | |  |  | |  |  |  |  |  | | **Feed/Inlet Water Pressure Pump SS** | | |  |  | | 3HP, Single phase 230V/50HZ | | |  |  | | **Automatic Sand Filter** | |  |  |  | | Flow: 40 Liter/Pm | |  |  |  | | Controller: | |  |  |  | | Automatic 3 processes rinse backwash service | | | |  | | It's adjust to fixed time for renew by 12 days cycle. | | |  |  | | Size: 13" x 54" H, FRP tank, | |  |  |  | |  |  |  |  |  | | **Automatic Activated Carbon Filter** | | |  |  | | Flow: 40 Liter Per Minute | |  |  |  | | Controller: Automatic on 220v/50Hz | |  |  |  | | Automatic 3 processes renew. (Reverse---wash---supply water) | | | |  | | Size: 13" x 54" H, FRP tank, | |  |  |  | |  |  |  |  |  | | **Automatic Water Softener Filter** | | |  |  | | Material: cation resin | | |  |  | | Flow: 40 Liter | |  |  |  | | Controller: Automatic | |  |  |  | | Automatic 5 processes renew. | | | |  | | (Reverse--Inject Salt--wash--Inject Salt Water--supply water) | | | |  | | Size: 13" x 54" H, FRP tank, | |  |  |  | |  |  |  |  |  | | **Brine tank 100 Liters / HDPE** | | |  |  | | **Reverse Osmosis Water Purification System** | | |  |  | | Product Water Ability: 10 Liter per minute ( 15 to 20 Dialysis Machine) | | | |  | | Rejection Rate: 99% | |  |  |  | | Recovery Rate: 60% | |  |  |  | | RO Membrane: Spiral Wound (Filmtec) 3/USA | | |  |  | | RO Pressure Pipe: PVC  RO multistorage pump-single phase | |  |  |  | |  | | | |  | | Flow Rate Indicator meter | |  |  |  | | system Pressure  Control | |  |  |  | | Automatic Operation System. | | |  |  | | Online Conductivity Meter | |  |  |  | | Equipped with Over/Under Voltage Protection device | | |  |  | | Over Loaded Alarm | |  |  |  | | **Auto Ultraviolet (UV) Disinfections Represurise pump with SS 01**  **HP Post pump 220VAC/50HZ** | | | | | | Lamp averagelife: 2000/hour / 125PSI/4Years | | |  |  | | Complete unit should be skid mounted | | |  |  |   Reservevoir tank: 1000L  Make: UK/USA/EEC/Japan  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
| **NICU EQUIPMENT** | | | | | | |
|  | **Phototherapy light Unit**  High density light power with fan less LED technology  5 level of light intensity  LCD graphic screen  Sophisticated base unit  Casters 2/3 with breaks  Targeting feature (red light)  Horizontal angle adjustment 360°  Light Source: Blue LED’s  Wave length: 450 nm or more (±2%)  Max light intensity (40cm): 120 µw/cm²/nm (±10%) or more  Min light intensity (40cm): 30 µw/cm²/nm (±10%)  Effective area (40cm): 40cm x 20cm or more better  LED life time: At max level 20,000 hours or more  Should be CE & TUV.  Make: UK/USA/EEC, Turkey/Japan.  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | **Over head Warmer**  Mobile overhead warmer with integrated baby bassinet along with Cushioned mattresses.  Microprocessor control.  Audible and Visual high and low temperature alarms.  Display of skin Temperature, Air Temperature,  Heater output (%),  APGAR Timer  Make: UK/USA-America/Europe/Japan.  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 6 |  |  |  |  |
| **PHYSIOTHERAPY** | | | | | | |
|  | **Infrared lamp**  Infrared lamp with 250W bulb and dimmer facility. Adjustable heat intensity  complete with protective grilles, stand and table bracket  Make:UK/USA/EEC/Japan.  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 4 |  |  |  |  |
|  | **Wax bath**  Thermostatically control with a safety cut out unit to heat adore free wax up to 18Kg of wax, with water capacity of 17liters. A smaller unit for hand therapy with wax capacity of 2.7kg.  Wax purification unit for with a wax capacity of 10kg and water capacity 10 liters.  Temperature range: 0-100°C.  Unit complete with :  Dust cover  Tapped wax container  Paraffin wax : 25kg  Wax bath thermometer  Make: UK/USA/EEC/Japan.  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 2 |  |  |  |  |
|  | **Two channel vacuum unit**  Pulsed or continuous vacuum.Powerful vacuum with no vibration.  Channel selection : 2 Channels A+B or Channel A only  Vacuum 0-0.5 Bar, Continuous or Pulsed  Unit should complete with:  Four vacuum electrode cups (60mm) each with sponge inserts  Four vacuum electrode leads  One link lead to connect unit to interferential / stimulation unit if require.  Trolley (local)  Make: UK/USA/EEC/Japan.  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | **Hot and Cold Packs**  different size of hot and cold packs  Make: USA/EEC/Japan.  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 6 |  |  |  |  |
|  | **Massager -707 & 717**  Large surface massage for intensive & effective massage.  Complete with  Two snap-on applicators  Carrying case.  Make:UK/USA/EEC/Japan.  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 2 |  |  |  |  |
|  | **Shoulder wheel**  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | **Quardicep drill table**  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | **Walker folding height adjustable**  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 2 |  |  |  |  |
| **DENTAL DEPARTMENT** | | | | | | |
|  | **Dental chair with all accessories**  3 hand-piece positions, Stainless steel instrument tray, Automatic adjustment for seat and backrest, Automatic zero position and three work positions, Multi-articulating headrest, swivel right armrest, Anatomic backrest, One 3-way syringe for doctor  3-Aseptic easy clean, tubing’s and fittings for the hand pieces, Air operated foot switch.  X-Ray Viewer, Switch buttons with joystick drive, Seamless upholstery coat, Calibri system (Continental), Counterbalance flex arm with pneumatic lock, Articulated tray holder, Borden or Midwest connection, X-Ray viewer  **Cuspidor:** Swivel at 90o, Cuspidor bowl in translucent color, cup filler with push button  Detachable water sticks, 2 saliva ejectors Venture (low & high)  **LIGHT:** 12V, 55W, 4500oK day light for perfect, Matching of colors, up, down, back, forward and inverse movement of the light head. Light intensity low: 16000 lux  Light intensity high:\_24000 lux, On/Off switch at chairs base, Light head made of polyester  **Dentist Stool:** Operating stool with adjustable height, Gas spring operated, Revolving on 5 castors, one with back support  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 04 |  |  |  |  |
|  | **Dental Instrument Set**  Probe, Mouth mirror and Forceps  Made from imported stainless steel, Rust resistant and Satin finish  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 16 |  |  |  |  |
|  | **Dental X-Ray**  **Dental X-Ray Machine 8mA complete set with all standard accessories for its installation and commissioning:**  Output rating: 8mA/70 kvp. Time selection: 0.05-0.6 sec, 0.25-3.2 sec. X-ray tube: Stationary anode tube. Focal range: 0.7x0.7mm. Aluminum equivalent: 2.5 mmAL. Supply power allowance:=/- 10% Electric power: 110 or 220  VAC, 50/60 Hz  Length: 1800mm  (Make: UK/USA/EEC/Japan.)  *Warranty period shall be of three years starting from date of supply of installation / commissioning*s | 01 |  |  |  |  |
| **EYE DEPARTMENT** | | | | | | |
|  | **Refraction Set**  set comprises of stable 300 by 200mm base plate; a light source with condensor and slit-holder to obtain several types of light beams, with a 12V 21W lamp; connection via cable and 4mm safety plugs; should supply with two spare lamps, a scaled disc of diameter 240 mm with double scales in degrees: 90-0-90 degree , a plane mirror with a fixing device on the disc; a Plexiglas half-cylinder with 200 mm diameter; a transparent half-cylinder tank with diameter 200 mm; height 2 cm; scaled in degrees 90-0-90 degree on its cylindrical surface, with positioning on the built-in disc; 2 diaphragms: a single-slit and a triple-slit.  Make: UK/USA/EEC/Japan.  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 01 |  |  |  |  |
|  | **Slit Lamp**  Stereoscopic converging binocular microscope.  3 Step magnifications changing by rotating drum. Long life and power-saving LED illumination. Superior function and compact design. Joy stick device, Easy coarse and fine adjustment for vertical and lateral movement of main body by one hand. Microscope Magnification10X, 16X, 25X. Eye piece 12.5X. Diopter adjustment :-5D ~+3D. Total magnification. Field of view:.10X (∅22.5mm), 15.98X (∅14.1mm), 25.53X (∅8.8mm)Slit Illumination Light source:LED(Light Emitting Diode)Aperture diameter : Continuous from 1 to 14, ∅0.3, 1, 5, 10, 14mm. Slit width Continuous from 0 to 14mm. Slit length Continuous from 1 to 14mm. Slit angle 0° to 180°. Filters Blue, Red-free. Primary input: AC 120/220 50-60Hz.  Make: UK/USA/EEC/Japan.  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 01 |  |  |  |  |
|  | **Ophthalmoscope**  Lens range: 0D + 1D – 10 / 15 / 20 / 40 D :  Aperture: Large aperture,small aperture, micro aperture, slit aperture, cobalt blue aperture & fixation aperture.  Filter: red free filter / Polarizing filter.  Operation mode: Intermittent, On-time should not exceed to 2 min with off-time not less than 5 mins.  Includes: Desktop charger & rechargeable handle.  Make: USA/EEC/Japan.  Should be FDA, CE & TUV.  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 02 |  |  |  |  |
|  | **A-Scan**  Large Touch screen operation with high-resolution backlit color LCD.  Complete measurement and calculation record within seconds .  Capture mode automatic and manual.  Ability to store up to five different user profiles.  USB for EMR archiving, SD disk for patient data storage.  A-Scan 10mm Calibration cylinder.  10 Lens Database.  Portable and compact weighing, Built in thermal printer.  Spring loaded Soft touch probes.  A-probe frequency: 10 MHz.  Measurements: ACD, Lens, Vitreous, Axial, Standard  Deviation, Individual Zone Velocity.Examination modes: Direct contact,Immersion,  Dense cataract, Aphakic, Pseudophakic, 4- GateManual.  IOL programs: Regression II, Theoretic T, Binkhorst, Holladay, Hoffer Q, Haigis. Post refractive formulas:Latkany Myopic Regression and Arambersi Double-K. Measurements limits: AXL- 18-45.  Clinical accuracy: ±0.10mm. Electrical accuracy: ±0.023mm  Make:UK/USA/EEC/Japan.  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 01 |  |  |  |  |
|  | **Diagnostic Set**  Diagnostic set including one rechargeable handle, one halogen otoscope and opthalmoscope,  3 autoclavable ear speculum (Ø 2.5, 3.5 and 4.5 mm), 2 laryngeal mirrors, bent arm throat lamp, tongue spatula, adjustable nasal speculum.  Complete with ophthalmoscope and otoscope head, support for tongue depressors and case  *USA/EEC/Japan*  *Should be FDA,CE & TUV*  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 12 |  |  |  |  |
| **MISCELLANEOUS** | | | | | | |
|  | **Centralized Medical gases supply system**  **ITEM 1.**  AUTOMATIC MANIFOLD FOR OXYGEN FOR 2 X 10 CYLINDERS  1 No.  WITH CONNECTION TO VIE TANK.  The Control Panel should handle three sources i.e. Two sided cylinders  battery and one VIE Tank / Oxygen Generation System.  The Manifold shall be a fully automatic type & shall switch  from Tank / Generation System to Cylinders Battery and “Bank in Use” to “Reserve”  bank without fluctuation in delivery supply line pressure  The system should be able to maintain continuous supply in case of power failure.  Capacity: 70-90, M3/h.  Function:  Fully automatic self-contained shuttle-valve with no electrical power  required for switching, Microprocessor based control panel with display.  Units of measure (psi/Pa/Bar).  LCD /LED display for the left bank, the right bank,& VIE Tank / Oxygen Generation System for the supply pressure.  Complete with relief valve, Emergency shut off valve and battery / bank  room alarm for left and right cylinder bank, pressure reducers of Tank Panel.  Service and maintenance works can be carried out without stopping  operation.  In built VIE Tank / Oxygen Generation System connection controlled by Manifold.  **ITEM 2**  **AUTOMATIC MANIFOLD FOR NITROUS OXIDE (2 X 5) 1 No,**  The Manifold shall be a fully automatic type & shall switch from "Bank  in Use" to "Reserve" bank with out fluctuation in delivery supply line  pressure&with out the need for external power after the switch-over  the "Reserve" bank shall then become the “Bank in Use" and the bank  in use shall become the reserve bank. The System should be able  maintain continuous supply in case of Power failure.  Capacity 10 - 50 M3/ h.  Function:  Fully automatic self-contained shuttle-valve with no electrical power  required for switching, Microprocessor based control panel with display.  Units of measure (psi/Pa/Bar) LCD / LED display for the left  bank, the right bank & for the supply pressure.  Complete with relief valve, Emergency shut off valve and battery / bank  room alarm for left and right cylinder bank.  Service and maintenance works can be carried out without stopping  operation.  **ITEM 3**  **WALL OUTLET.**  Outlet should have double lock quick connect and gas specific. The latch /  Locking release mechanism should be incorporated on the gas outlet with  Geometrical key indexing and should have a parking facility so that in the  Event of release of probe by pressing the spring loaded center, the probe  / Flow meter does not fall down.  Gas outlets wall mounted, Complete in Box/casing gas Specific, Self sealing  valve, indexed to eliminate interchangeability/erroneous tapping of gas  and cover plate.  Oxygen Qty.92 Nos.  Nitrous Oxide Qty.6 Nos.  Compressed Air 4/5 bar Qty.92 Nos.  **ITEM 4**  **Air Venturi System with Gauge 25 Nos.**  The injector should be connected to wall source of pressure gas using  Direct probe or rail mounting system. Max. Vacuum 650 to 700 mbar.  System should be supplied with 02liter Un Breakable &Sterilizable  Collection jar.  **ITEM 5(a)**  **Multiple Zone Valve Box with Monitoring Gauge forthree gases 3 No.**  Microprocessor controlled with digital display of pressure.  Each wall type zone valve box shall consist of the following Components.  A steel valve box which can house two to four shut-off-ball valve with tube  extensions.  **ITEM 5(b)**  **Multiple Zone Valve Box with Monitoring Gauge for two gases 4 Nos.**  Microprocessor controlled with digital display of pressure.  Each wall type zone valve box shall consist of the following Components.  A steel valve box which can house two to four shut-off-ball valve with tube  Extensions.  **ITEM 6**  **CEILING PENDANTS:**  Flexible Ceiling Type Outlets:  Oxygen 12 Nos.  Air 12 Nos.  N2O 6 Nos.  **ITEM 7**  **OXYGEN FLOW METER WITH HUMIDIFIER & PROBES: 80 Nos.**  With autoclave able and unbreakable humidifier bottle and probe.  **ITEM 8**  **COPPER PIPING (IMPORTED FROM UK, GERMANY, USA):**  Supply and installation of seamless copper pipe, deoxidized and  for use in **piped medical gas systems**,  Therefore it is guaranteed that the tubes are  leak-tight, free of pores, blisters etc, Tube  is drawn seamless, according to **HMT 2022 (HTM 02-01) Or equivalent**.  Material :-  Copper tubes, including fixing clips for  ceiling and walls, sleeves, reducers, tee's  bends and soldering material.  1350Meter Size 15 x 1.  1500 Meter Size 12 x 1.  **ITEM 9**  **ANESTHESIA SCAVENGING SYSTEM 6 Nos.**  Anesthetic gas scavenging systems designed to remove exhaled anesthetic  gas mixtures from operating theatres driven by Air/Oxygen at 05 bar or  through vacuum plant, anesthetic rooms and recovery area.  **ITEM 10**  Motor 7.5KW (10HP)  Each 10 HP Motor Produces 1200 Liters and 4 motor supplies 4800 Liters min when duplex compressor in use and more than 2400 minimum Liters per minute in each Skid Mount compressor with Hepa filter  Working Pressure: 9.5 KG/CM2G (135 PSI)  **Air Dryers**  2 x Dryer  Volume Flow Rate @ Inlet: 1200 lit/min @80 ⁰C  Max. Operating Pressure: 14 barg.  Max. Ambient Temperature: 45 ⁰C  Pressure Dew Point: +3 ⁰C  Power Source: 230V / 1PH / 50HZ  **Filtration**  2 x Pre-Filter – 1200 lit/min (3 Micron) ½”  4 x After-Filter – 1200 lit/min (0.01 Micron) ½”  1.9) By Pass System (Local with Imported Fittings):  Isolation Valves, Adaptors, Regulators & Pressure Switches.  **Electrical Panel**  Simple Controlled Electrical Panel, Pressure based Automatic Switching between Compressors, Contractors, Overload, Power Indication Lights and Circuit Breakers.  Make: UK/USA/EEC/Japan.  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | **Split Air condition 1.5 Ton**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 148 |  |  |  |  |
|  | **Freezer with Temp recorder -20degree C**  Digital display, Microprocessor controlled.  Platinum resistance sensor in the temperature control  Temperature range: -10oC to -25oC  Capacity: 700 Litres; Shelves: 4 or more  Built in audio visual alarm  Warning system with built in audible and visual alarm.  CO2 back up.  Temperature/test electronic probe/alarm test  Self-closing door folding magnetic gasket on four sides for perfect closing  Make: UK/USA/EEC/Japan.  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 2 |  |  |  |  |
|  | **Deep Freezer**  Branded  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | **Ordinary Fridge 11cft**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | **Spring scale 500ml**  **Should be CE Approved**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | **Air drill with attachment** | 1 |  |  |  |  |
|  | **Appendectomy Set**  Made from imported stainless steel, Rust resistant and satin finish  **Local Manufacturers only can Participate directlyMust be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 2 |  |  |  |  |
|  | **Astinmoor**  Made from imported stainless steel, Rust resistant and satin finish  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 2 |  |  |  |  |
|  | **Auto Fine Lumen Cleaner**  Make: UK/USA/ EEC/Japan  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | **Autoclave**  40 liters or more, manual control system. Eco friendly and economically design with over temperature, pressure & elec. Over load protection with emergency exhaust switch with low water buzzer with over heat/temp. indication light chamber. Temperature upto 121oC;Digital PID Controller with Timer and Tunning:  CE Marked  Make: UK/USA/EEC/Japan  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | **Autoclave steamgenrated. Double door (CSSD)**  150 liters, manual control system. Eco friendly and economically design with over temperature, pressure & elec. Over load protection with emergency exhaust switch with low water buzzer wit hover heat/temp. indication light chamber  CE Marked  Make: UK/USA/EEC/Japan  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 2 |  |  |  |  |
|  | **Autoclave Verticle**  40 liters or more, manual control system. Eco friendly and economically design with over temperature, pressure & elec. Over load protection with emergency exhaust switch with low water buzzer with over heat/temp. indication light chamber. Temperature upto 121oC;Digital PID Controller with Timer and Tunning:  CE Marked  Make: UK/USA/EEC/Japan  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 2 |  |  |  |  |
|  | **Autoclave double steam str. Single door (CSSD)**  40 liters or more, manual control system. Eco friendly and economically design with over temperature, pressure & elec. Over load protection with emergency exhaust switch with low water buzzer with over heat/temp. indication light chamber. Temperature upto 121oC;Digital PID Controller with Timer and Tunning:  CE Marked  Make: UK/USA/EEC/Japan  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | **Baby Running Trolley with Warmer**  Baby running trolley with warmer  Mobile resuscitation trolley with overhead warmer with integrated baby bassinet along with Cushioned mattresses.  Microprocessor control.  Audible and Visual high and low temperature alarms.  Display of skin Temperature, Air Temperature  Heater output (%),  APGAR Timer,  Resuscitation system  Make: UK/USA-America/Europe/Japan.  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 2 |  |  |  |  |
|  | **Surgical instrument set**  Metal Ruler 6’’ (1 Nos)  Scalpel Handle #3 (1 Nos)  Scalpel Handle#4 (1 Nos)  TC Wire Cutting Scissors Angled 4 3/4’’ (1 Nos)  Rochester-OchsnerFoersepsCvd 5 ½’’ (1 Nos)  Allies Tenaculum Forceps Delicate 5 ¾’’ (2 Nos)  TC Derf NH Serr 4 ¾’’ (1 Nos)  TC-Crile-Wood NH NarroeSerr 6’’ (1 Nos)  TC Ryder Micro NH Serr 6’’ (1 Nos)  Stevens Tenotomy Scissors Delicate S/S 4 ½’’ (1 Nos)  Stevens Tenotomy Scissors Delicate B/B 4 ½’’ (1 Nos)  Metzenbaum Scissors Cvd 4 ½’’ (1 NOs)  Reynolds Sciss Nar Tip Cvd 6 (1 Nos)  Semkin Dressing Forceps Delicate 5’’ (1 Nos)  Semkin Tissue Forceps 1x2 Teeth Delicate 5’’ (2 Nos)  Miller-Senn Retractor Double End Sharp 6’’ (1 Nos)  Iris Scissors str 4 1/8’’ (1 Nos)  Iris Scissors Cvd 4 1/8’’ (1 Nos)  Mayo ScissorsStr 5 ½’’ (1 Nos)  Mayo ScissorsCvd 5 ½’’ (1 Nos)  MetzenbaumScissorsCvd 6’’ (1 Nos)  Hartmann Mosquito ForcepsStr 3 1/2 ‘’ (3 Nos)  Hartmann Mosquito ForcepsCvd 3 1/2 ‘’ (3 Nos)  Halsted Mosquito Forceps Str 5’’ (6 Nos)  Halsted Mosquito Forceps Cvd 5’’ (6 Nos)  Adson Dressing Forceps Serr 4 ¾’’ (1 Nos)  Adson Tissue Forceps Slide grasping 4 ¾’’ ( 1 Nos)  Allis Forceps 4x5 Teeth 6’’ (1 NOS)  Extra Delicate Allis Forceps 4x5 Teeth 5’’ 91 Nos)  Freer Dissector Double End S/B 7 (1 Nos)  Joseph Hook 1 Prong Sharp 6 ¼’’ (2 Nos)  Backhaus Towel Clamp 3 ½’’ (2 Nos)  Backhaus Towel Clamp 5 1/4’’ (6 Nos)  Cushing Nerve and Vein retractor 9 (1 Nos)  Frazier Suction Tube 7 French (1 Nos)  Frazier Suction Tube 8 French (1 Nos)  Crile Retractor D/E 4 ½’’ (1 Nos)  US Army Retractor (Set of 2) 8 1/2 ‘’ (1 Nos)  Senn retractor Double End Sharp 6 ¼’’ (1 Nos)  Ballenger Sponge Forceps StrSerr 7’’ (2 Nos)  Ballenger Sponge forceps CvdSerr 7’’ (2 Nos)  Foerster Sponge Forceps StrSerr 9 ½’’ (2 Nos)  Made from imported stainless steel, Rust resistant and satin finish  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | **Centrifugal Cell Washer(Electronic)**  **Must be CE Marked Product Imported / Local (Export Quality)**  *Warranty period shall be of three years starting from date of supply o f equipment to the consignee end* | 1 |  |  |  |  |
|  | **Cervical Traction with Weight.**  **Must be CE Marked Product Imported / Local (Export Quality)**  *Warranty period shall be of three years starting from date of supply o f equipment to the consignee end* | 2 |  |  |  |  |
|  | **Chisel Impical set**  Made from imported stainless steel, Rust resistant and satin finish  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | **Cholecystomy set**  Made from imported stainless steel, Rust resistant and Satin finish  Export Quality Material  Must be FDA & CE Marked Product  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 2 |  |  |  |  |
|  | **Cold Storage (For Morgue)**  Morgue for 4 to 6 body Storage with modular construction.  Complete with 2 bay. 2 tier304 grade stainless steel storage cabinet, 80 mm high density rigid insulation polyurethane insulation chiller, Stainless Steel height adjustable racking & rolling system, 304 grade Stainless Steel body storage trays, 10°C to -5°C cooling system with digital control panel, body transport trolley.  Make: UK/USA/EEC/Japan.  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | **Defibrillator**  with all Standard Accessories  **Specifications:**  Biphasic technology  Defibrillator Biphasic upto270 Joules atleast  (***Mandatory***)  Charging time: Less than 10 sec for 270 J  ECG­:-6 Channel ECG  Leads: I, II, III, aVR, aVL, aVF  Heart rate: 50-300 per min  Display: TFT LCD 5’ Inches or More  Built-in Printer with Memo Print  Data management:-  Removable storage unit: 01GB or more with voice recording up to 80 Hours or more  Rechargeable external battery with minimum 04years shelf life  Monitoring time 03Hours or more Extendable  Complete with all standard Accessories  External Paddles for Adult & Built-in Pediatrics Paddles  ECG Leads minimum 05 cable set  Ultrasound Gel  Power Cord  (References required from reputable institute like SIUT,Civil,JPMC,Indus,Dow,Tabba,AKUH)  Installation must be more than 50 units atleast  Make: USA/EEC/Japan/UK  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 10 |  |  |  |  |
|  | **Digital Blood Shaker**  3D rotation, with 5-40r/min speed. Rotating angle 20°. Max bearing capacity : 0.8kg;Adjustable low & high flow alarms;  (10-990ml) up to 1000 donations can be stored,  Automatic Clamp of tubing at the end of donations, RS232 cable for barcode reader  Make: UK/USA/EEC/Japan.  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end.* | 1 |  |  |  |  |
|  | **Digital video colonoscopy.**  **Flexible**  Magnification 15X  Focus length: 300 mm  Field of view: diameter 22 mm  Depth of field: > 18/4/2 mm  Light intensity: > 28.000 lux  Interchangeable adjustable eyepiece 12.5X  Illumination: longlife 5.9V / 5W LED  Superior optics with individually adjustable eyepiece  Built-in rheostat for brightness adjustment  Easily changeable, LED bulb  User-selectable built-in green filter  Two (gross/fine) focus adjustment handles  Two (gross/fine) height adjustment handles  Make: UK/USA/EEC/Japan.  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | **Laryngoscope**  fiber optic laryngoscope with blades and halogen light source in the handle  sterilizable with re-chargeable batteries  Make: UK/USA/EEC/Japan  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | **Disinfection Instrument washer double door**  Fully automatic operation,fromprewarsh,wash,rinse to dry,high class stainless material transparent glass with lighting east to check chamber size: w520xd530xh600mm, 165 liters spray. 3 rotational cleaning arms with water spray nozzle.dry stage drying tempreture 121\*c and keep 85min for sterilizing.  Make: UK/USA/EEC/Japan  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 2 |  |  |  |  |
|  | **Dry cabinet single door for anesthesia accessories**  Made from imported stainless steel, Rust resistant and Satin finish  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | **Plaster cutter**  Compact, lightweight plaster cutter equipped with a cooling fan  With blade lock and adjustable guide and cutting depth gauge.  wide range of saw blades  Stainless steel blades .  non-stick blades with PTFE coating  Make: UK/USA/ EEC/Japan  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | **Endoscope**  Video Gastro scope with 15” Medical grade LCD Monitor  ·         High Resolution Video CCD chip and latest digital scope  ·         180º axially rotatable connector head improves user safety  ·         Scope should provide the full screen image on LCD  ·         Scope should have the capability to visualize the suspicious tissues and lesions  ·         Field Of View :140º or better  ·         Focal range: 5 – 100 mm  ·         Outer Diameter:9.8 mm or less  ·         Instrument Channel: 2.8 mm or less  ·         Minimum Angulations (210 up , 120 down and 120 right / left or better)  ·         Minimum Working length: 1050 mm or better  ·         Scope should be capable for band ligation procedure  Accessories:  ·         Biopsy forceps  ·         KSnare/Swan Blade/Bipolar Hemostasis forceps  ·         Grasping Basket  ·         Channel Cleaning Brush Set.  ·         Case and Maintenance kit.  High Resolution Digital Video Processor with Xenon Light source  High Resolution Image video processor with digital signal processing technology  Maximum working comfort  Digital Video Processing for effective mucosa evaluations  Contour Enhancement  Brightness control: 11 steps  Easy data management for rapid storage  Data management: 30 patient / 50 users  Light Source  Xenon light source 100 W or more  Xlum Mode  Color Correction: ± 5 steps  Output: 1x RGB, 2x Y/C, 1x Composite  Make: UK/USA/EEC/Japan. | 1 |  |  |  |  |
|  | **Fully automated Chemistry Analyzer**  Automated clinical chemistry analyzer having several different analysis modes including Absorbance, 1 point end, 2 point end, 1 point rate, 2 point rate and Kinetic,immunoturbimatric assay. Random access with reusable plastic or pyrex cuvettes with throughput 100 or more tests per hour, maximum filling reagent capacity.  Should be built in inventory management system automatically tests and reports.Optical design user-selected monochromatic or dichromatic results. 8 wavelengths ranging from 340-700nm; 340,405,505,546,630 nm + Two Optional.Linear ranges-0.2-3.0A  Must have 120 or more programmable tests.  Reaction Volume: 250 µl , Long life IAD filters  Sample type: Serum, plasma,  Sample Volume: 3-30 µl  Reagent position :maximum 37; dilution predilution,dispencing single or multiple  Maximum number of reaction 40  Reaction time: Max 15 min  Should be with PC & Laser printer  US FDA Approved  Country of origin: UK/USA/EEC/JAPAN  *Warranty period shall be of three years starting from date of supply* | 1 |  |  |  |  |
|  | **Cordless battery operated orthopedic Drill with attachment**  Speed: 0—1400 RPM or more  Noise Level: ≤20Db  Power: ≥30W  Sterilization temperature and pressure up to 135°  Complete With:  Cordless Drill (Battery) System  Cordless Drill with end Cap  Three- Jaw Chuck up to dia. 7.0 mm, round and triangular shanks with chuck key  Protection & guidance sleeve for kirschner wires  AO – Chuck  AO - Quick- Action chuck for drill wires up to dia. 2.5 mm  AO - Din Chuck for (Aesculapmeduallary reamer)  AO - Harris Chuck  AO - Hudson Chuck  Desktop charger, dual 240V –  Desktop charger, mono 240V  Reduction gears with AO standard clutch, large, Speed 0-150 rpm  Sternum Saw  Rechargeable battery, with extension  Standard Rechargeable Battery  Funnel for inserting rechargeable battery into Cordless Drill  Hand piece + Foot Switch + Electric Console + Bur Adaptor + Aluminum carrying case  Make: UK/USA/EEC/Japan/China  Should be CE/FDA  References from CHK, JPMC, ASH, AKUH, LNH & AO Clinic.  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | **LED Ceiling Mounted OT Lights b**  with all Standard Accessories  High performance LED illumination for every type of surgery  Cold Light  White Light  Excellent contrast in the operating field  Shadows reduction in the surgical field and total illumination of the surface and at depth –  Frequency 50/60 Hz  Light source technical data  Light Source : LEDs  Average life : 50,000 hrs or more  Lighting at the centre, Ec @ 100cm  Diameter of operating field Min 185mm/Max 340mm  Illumination Minimum 2 x120,000 lux @ 1 meter-  Total Lux @ 240000 or more  Color temperature, CCT 4,500° K ± 2%  Color rendering index, Ra 95±1  Make: UK/USA/EEC/Japan  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 2 |  |  |  |  |
|  | **Humbies knife**  Made from imported stainless steel, Rust resistant and Satin finish  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | Inferential Machine | 1 |  |  |  |  |
|  | **Instrument Cleaner Super Sonic**  Ultrasonic frequency 40,000 Hz  Tank Material Stainless  Tank Capacity: 600ml  Timer 180S,360s,,Digital timer  Power 35W  Unit Size 175x 110x155mm(LxWxH)  Tank Size 150x85x65mm (LxWxH)  Make: UK/USA/Europe/Japan.  Warranty period shall be of two years starting from date of full functional commissioning | 1 |  |  |  |  |
|  | **Surgical instrument set**  Metal Ruler 6’’ (1 Nos)  Scalpel Handle #3 (1 Nos)  Scalpel Handle#4 (1 Nos)  TC Wire Cutting Scissors Angled 4 3/4’’ (1 Nos)  Rochester-OchsnerFoersepsCvd 5 ½’’ (1 Nos)  Allies Tenaculum Forceps Delicate 5 ¾’’ (2 Nos)  TC Derf NH Serr 4 ¾’’ (1 Nos)  TC-Crile-Wood NH NarroeSerr 6’’ (1 Nos)  TC Ryder Micro NH Serr 6’’ (1 Nos)  Stevens Tenotomy Scissors Delicate S/S 4 ½’’ (1 Nos)  Stevens Tenotomy Scissors Delicate B/B 4 ½’’ (1 Nos)  Metzenbaum Scissors Cvd 4 ½’’ (1 NOs)  Reynolds Sciss Nar Tip Cvd 6 (1 Nos)  Semkin Dressing Forceps Delicate 5’’ (1 Nos)  Semkin Tissue Forceps 1x2 Teeth Delicate 5’’ (2 Nos)  Miller-Senn Retractor Double End Sharp 6’’ (1 Nos)  Iris Scissors str 4 1/8’’ (1 Nos)  Iris Scissors Cvd 4 1/8’’ (1 Nos)  Mayo ScissorsStr 5 ½’’ (1 Nos)  Mayo ScissorsCvd 5 ½’’ (1 Nos)  MetzenbaumScissorsCvd 6’’ (1 Nos)  Hartmann Mosquito ForcepsStr 3 1/2 ‘’ (3 Nos)  Hartmann Mosquito ForcepsCvd 3 1/2 ‘’ (3 Nos)  Halsted Mosquito Forceps Str 5’’ (6 Nos)  Halsted Mosquito Forceps Cvd 5’’ (6 Nos)  Adson Dressing Forceps Serr 4 ¾’’ (1 Nos)  Adson Tissue Forceps Slide grasping 4 ¾’’ ( 1 Nos)  Allis Forceps 4x5 Teeth 6’’ (1 NOS)  Extra Delicate Allis Forceps 4x5 Teeth 5’’ 91 Nos)  Freer Dissector Double End S/B 7 (1 Nos)  Joseph Hook 1 Prong Sharp 6 ¼’’ (2 Nos)  Backhaus Towel Clamp 3 ½’’ (2 Nos)  Backhaus Towel Clamp 5 1/4’’ (6 Nos)  Cushing Nerve and Vein retractor 9 (1 Nos)  Frazier Suction Tube 7 French (1 Nos)  Frazier Suction Tube 8 French (1 Nos)  Crile Retractor D/E 4 ½’’ (1 Nos)  US Army Retractor (Set of 2) 8 1/2 ‘’ (1 Nos)  Senn retractor Double End Sharp 6 ¼’’ (1 Nos)  Ballenger Sponge Forceps StrSerr 7’’ (2 Nos)  Ballenger Sponge forceps CvdSerr 7’’ (2 Nos)  Foerster Sponge Forceps StrSerr 9 ½’’ (2 Nos)  Made from imported stainless steel, Rust resistant and satin finish  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | **Inter locking instrument Set**  Made from imported stainless steel, Rust resistant and Satin finish  CE Marked Product/ Export Quality | 1 |  |  |  |  |
|  | **Juster Adjustable**  0.1-2, 1-10, 2-20  **Should be CE Approved**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 2 |  |  |  |  |
|  | **Juster Fixed**  5,10,25,50 ml  **Should be CE Approved**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 4 |  |  |  |  |
|  | **Laproscopic** | 1 |  |  |  |  |
|  | **Laryngo Endoscope (Rigid)**  Make: UK/USA/ EEC/Japan | 1 |  |  |  |  |
|  | **Mastroid and middle ear surgery set**  Made from imported stainless steel, Rust resistant and Satin finish  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 2 |  |  |  |  |
|  | **OT Table**  Multipurpose Hydraulic Operation table  Systems operatable by foot pump for hydraulic function.  Movement for Systems by manual gas pistons & hydraulic.  Table top Board divided in 05 sections  Table Elevation & decrease system, by hydraulic piston commanded by foot panel.  Head Rest Detachable & adjustable by piston  Surface of the table board in radio transparent material, tunnel with guides for X-ray Cassettes  Table Top Dimension L X W X H: 2000x500mm or more From 650 to 850mm or better  Lateral Tilt: -20˚ to + 20˚  Back Section,Adjustable from +80˚ to -40 ˚  Maximum Patient Weight: 200 or more  **Standard Accessories:**  Anesthesia Screen ARC, in Stainless Steel  Head Rest Padded, Detachable  X-Ray Cassette  Mattress  Iv Stand  Make: UK/USA/EEC/Japan  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 3 |  |  |  |  |
|  | **Nebulizer**  Nebulization Rate: 0 t0 70ml/hour  Air Volume Flow rate: 0 to 10 l/min  Mist Density: 105mg/min  HF frequency: 2.4MHz  Timer: 0 to 15min / continuous mode  Includes: mobile stand  Make: UK/USA/EEC/Japan  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 2 |  |  |  |  |
|  | **Otoscope fiber 3.5V**  Illumination form: Fiber optics transmit light  Magnification of lens: 2X  Operation mode: Intermittent, On-time should not exceed to 2 min with off-time not less than 5 mins  Make: USA/EEC/Japan.  Should be FDA, CE & TUV.  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | **PABX**  Panasonic/Sony/equivalent  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | **Phaco machine**  Peristaltic pump, Realistic venture effect. Cool Phaco technology. Micro Incision Cataract Surgery (MICS).Dual linear foot switch.Widest range of MICS instruments, 20 gauge forceps and scissors.Auto Self-Test and tuning..Mini Mega ultrasonic Tip. Separate control for Vacuum and Flow rate.  Cool Phaco:Frequency: 27 – 55 kHz, Pulsation: Cool flash technology  Aspiration:Flow rate: 1 – 50 ml, Vacuum: 1 – 550 mmHg  Vitrectomy:Pneumatic, nominal pressure: 1.4 bar (20 PSI), Stroke: 10 – imp/min  Diathermy:Power: 0 – 15W at 125 Ohms, Frequency: 1.5 MHz, Pulsation: 3 Hz  Make: UK/USA/EEC/Japan.  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | **Plaster Cutter**  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | **Portable LED OT light**  Mobile spot Light  Minimum 35,000LUX/0.5m or more  Color temperature: 4000K or more better  Flexible arm with up & down movement and 360 degree rotation.    **Make: UK/USA/EEC/JAPAN**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 6 9 |  |  |  |  |
|  | **Mobile X-Ray Machine 100mA complete set with all standard accessories for its installation and commissioning:**  High frequency X-Ray inverter type Generator  Having output power of 4kW or more  X-Ray tube 50 KHU or more  KVP range: 40-110kV or more  MA Range: 25-100mA or better,  mAS range: 1.0-200mAs or better  Centering light indicator  Standard power supply cable, standard x-ray exposure switch cable,  Vertical movement of the arm to lower or raise the tube collimator assembly, Rotation of the Collimator with reference to the power module (+/- 90deg) or better.  Country of originshould be USA/western Europe/Japanwiththeirrespectivestandards MDD (CE) compliance/MHLW Japan  *Warranty period shall be of three years starting from date of installation / commissioning* | 1 |  |  |  |  |
|  | **Portable LED OT light**  Mobile spot Light  Minimum 35,000LUX/0.5m or more  Color temperature: 4000K or more better  Flexible arm with up & down movement and 360 degree rotation.    **Make: UK/USA/EEC/JAPAN**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 2 |  |  |  |  |
|  | **Reaming Set**  Made from imported stainless steel, Rust resistant and satin finish  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | **Reduction Forcep Set**  Made from imported stainless steel, Rust resistant and satin finish  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | **Refrigerator**  **Must be CE Marked Product Imported / Local (Export Quality)**  *Warranty period shall be of three years starting from date of supply o f equipment to the consignee end* | 1 |  |  |  |  |
|  | **Refrigerator(Rotary)** | 1 |  |  |  |  |
|  | **Retinoscope**  Rotation: ≥190 degree  Length of streak image: ≥30mm  Width of Length image: ≥1.5mm  Operation mode: Intermittent, On-time should not exceed to 2 min with off-time not less than 5 mins  Includes: Desktop charger & rechargeable handle.  Make: USA/EEC/Japan.  Should be FDA, CE & TUV.  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | Screw Box  Make: USA/EEC/Japan.  Should be FDA, CE & TUV. | 1 |  |  |  |  |
|  | **Bone Small Fragment Fixation Set**  Small Fragment Instrumentation Set  Containing:  Tap for screws Ø 3.5mm, length 110/50mm  Tap for cabcellous bone screws Ø 4.0mm, length 110/85mm.  Drill bit Ø 2.5mm, length 110/85mm, 2-flute, for quick coupling.  Drill bit Ø 3.5mm, length 110/85mm, for quick coupling.  DCP drill guide 3.5, for neutral and load position.  Drill sleeve insert 3.5/2.5, length 42mm.  Double drill guide 3.5/2.5  Triple drill guide 2.0, with 3 holes.  Depth gauge for screws Ø 2.7 to 4.0 mm, measuring range upto 60mm.  Screwdriver, hexagonal, small, with holding sleeves.  Screwdriver shaft, hexagonal, small, Ø 2.5mm.  Handle with quick coupling, self-holding, length 110mm.  Screw forceps, self-holding, length 85mm.  Sharp hook, length 155mm.  Bending iron for plates 2.4 to 3.5, length 145mm.  Bending iron for plates 2.4 to 3.5, length 145mm.  Bending template for LC-DCP 3.5 and DCP 3.5, length 87mm.  Bending template for LC-DCP 3.5 and DCP 3.5, length 114mm.  Bone holding forceps, self centering, speed lock, length 190mm.  Reduction forceps with points, ratchet lock, length 130mm.  Reduction forceps, toothed ratchet lock, length 140mm.  Reduction forceps, with points, wide, soft lock, length 127mm.  Reduction forceps, toothed, soft lock, length 146mm.  Bone holding forceps, self centering, soft lock, length 191mm.  Bone liver, width 15mm, length 160mm.  Bone liver, small, short narrow tip, width 8mm, length 160mm.  Periosteal elevator, slightly curved, straight edge, width 6mm, length 187/87mm.  Bone liver, small, short narrow tip, width 6mm, length 160mm.  Counter sink 3.5  LCP drill guide 3.5, for neutral and load position.  LCP drill sleeves 3.5, for drill bits, Ø 2.8mm.  LCP drill bit Ø 2.8mm with stop, length 165mm, for quick coupling.  Bone holding forceps, self centering, soft lock, length 191mm.  DLS drill guide 3.7, for drill bits Ø 3.1mm, length 165mm, 2-flute, for quick coupling.  Drill bit Ø 3.1 w/stop caliber L250/225.  Locking drill sleeves Ø 3.1 percut.  Made from imported stainless steel, Rust resistant and stain finish  **Uk/USA/EEC/Japan/China also**  **Local Manufacturers can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | **SMR, Septoplasty set**  Made from imported stainless steel, Rust resistant and Satin finish  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 4 |  |  |  |  |
|  | **Spring scale 500ml**  **Should be CE Approved**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | **Suction Machine**  Suction Machine with two polycarbonate collection Jars.  4 antistatic castors , 2 with brakes.  Piston type, oil free.  Pump flow: 40L/min .  overflow prevention.  Hydrophobic PTFE filter.  3 m PVC Suction tube.  Metallic handle, Metallic curved and straight catheter .  Make: UK/USA/ EEC/Japan  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 3 |  |  |  |  |
|  | **Tonsillectomy Set**  Made from imported stainless steel, Rust resistant and satin finish  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 4 |  |  |  |  |
|  | **Total hip set**  Made from imported stainless steel, Rust resistant and satin finish  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | Vacuum Steam Vaporizer  **Local Manufacturers only can Participate directly**  **Must be CE Marked Product/ Export Quality**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | **Ventilator with all Standard Accessories**  Specifications:-  Patient type: neonates, Peads& adult  VENTILATION MODES:  The modes of ventilation should include: Volume :  The modes of ventilation should include : AVM/ASV  Volume Ventilation (CMV, Assist CMV), Pressure Support Ventilation (PSV), SIMV, SPONT, MANUAL,SIGH, Apnoea BACK-UP, CPAP, Bilevel, & NIV. Manual breath lock and inspiratory/expiratory hold.  (CMV, Assist CMV), Pressure Support Ventilation (PSV), SIMV, SPONT, MANUAL,SIGH, Apnoea BACK-UP, CPAP, Bilevel, & NIV. Manual breath lock and inspiratory/expiratory hold.  TECHNICAL SPECIFICATIONS:  The patient range should be from neonate, pediatric and adults with easy operation.  Breaths Rate for CMV+, PCV+, P – SIMV+ should be from at least 6 bpm to 140 bpm  PEEP & CPAP from 0 cmH2O up to 35 cmH2O  Tidal Volume from minimum 2 ml to 2000 ml for Volume mode  Inspiratory minute Volume  Expiratory minute volume  The I: E ratio: 1:4 - 4:1  ALARMS:  Operator adjustable alarms should include low/high minute volume, high pressure, low/high tidal volume, apnea rate  Oxygen concentration, disconnection, loss PEEP  Apnoea Alarm  Display:  Screen size not less than 12” Inch or more TFT Color Screen with Loops, Graph & Monitoring on the same window  EVENT LOG:  Memory Storage up to 100 Events  Power Supply:  100-230V  Internal Battery Operation 4-6 Hour  Supplied should be with supporting arm, trolley, breating circuits for adult, paeds and neonates , power cord and operating manual  **Medical Air Compressor or Turbine system**  CE & IEC Certified.  Make :UK/USA /Europe  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 16 |  |  |  |  |
|  | **Water Distillatordeionzer complete**  Processing capacity: 4 liters or more  Procesing time: 7-8 hours or better  Automatic control thermostat with block system at the end of operation  Complete with 6active caol filters, 1 detergent, 1 container, for produced distilled water  Make: UK/USA/ EEC/Japan  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 1 |  |  |  |  |
|  | **Baby Weighing Scale**  Digital scale with extra large LCD display (65 mm).  Strong and durable  steel comfortable platform.  **Should be CE Approved**  *Warranty period shall be of three years starting from date of supply of equipment to the consignee end* | 02 |  |  |  |  |
|  | Wire Instrument Set  **Must be CE Marked Product Imported / Local (Export Quality)**  *Warranty period shall be of three years starting from date of supply o f equipment to the consignee end* | 1 |  |  |  |  |