## oRIGINAL (hidden)
SPECIFICATIONS
S No. | Description | Qty
Serial No. 3
Intercranial & Transphenoidal Endoscopy With Camera System
1 | Ventriculoscope Wide Angle Straigth Forward Telescope 6°, with angled eyepiece, autoclavable, | 1
O.D.: 6.1 mm, instrument channel diameter 2.9 mm, with suction and irrigation channel 1.6 mm, working length 18 cm, with
irrigation adapter, fiber optic light transmission incorporated, color code: green
2 | Operating sheath for ventriculoscope outer diameter 6.8 mm, working length 13.3 cm | 1
3 | Obturator | 1
4 | Click Line Scissors, single action jaws, pointed, diameter 1.3 mm, working length 30 cm, | 1
Consisting of:
Metal Handle, without ratchet
Outer Tube with insert
5 | Click Line Biopsy Forceps, both jaw parts movable, Ø 2.0 mm, working length 30 cm, | 1
Consisting of:
Metal Handle, without ratchet
LCA Outer Tube with insert
6 | Click Line Ventriculostomy Forceps, diameter 2.0 mm, working length 30 cm, | 1
Consisting of:
Metal Handle, without ratchet
Outer Tube with insert
7 | Click Line Grasping Forceps with teeth, 2.0 mm, working length 30 cm, | 1
Consisting of:
Metal Handle, without ratchet
Outer Tube with insert
8 | Puncture Needle | 1
9 | Coagulating Electrode, bipolar, 5 Fr. | 1
10 | Bipolar High Frequency Cord with 2 x 4 mm banana-plug to Coagulator Valleylab, length 300 cm | 1
11 | Holding System, autoclavable, with fastener: KSLock, | 1
Consisting of:
Rotation Socket to clamp on the operating table, for use with European and United States standard rails, with lateral clamping
element for height and angle adjustment of the articulated stand
Articulated Stand, reinforced version, straight, with one mechanical central clamp for all five joint functions, heigth 30 cm,
operating range 37 cm, with fastener: KSLock (female)
Clamping Jaw, metal, with axial intake, for use with instrument and telescope sheaths, clamping range 4.8 up to 12.5 mm, with
fastener: KSLock (male)
12 | Scissors, pointed, lightly curved jaws, double action jaws, diameter 1,7 mm, length 30 cm | 1
13 | Balloon Catheter, O.D. 0.7 mm, length 40 cm, sterile, single use, 10 pieces | 1
Endoscopic Pituatry Surgery Set
1 | Straight Forward Telescope 0°, enlarged view, diameter 4 mm, length 18 cm, autoclavable. Fiber optic light | 1
transmission incorporated. Color code: green
2 | Suction and Irrigation Sheath 0°, outer diameter 4.8 x 6 mm, working length 14 cm, for use with Telescope | 1
lens irrigation system
3 | Forward-Oblique Telescope 30°, enlarged view, diameter 4 mm, length 18 cm, autoclavable. Fiber optic light | 1
transmission incorporated. Color code: red
4 | Suction and Irrigation Sheath 30°, outer diameter 4.8 x 6 mm, working length 14 cm, for use with Telescope | 1
and lens irrigation system
5 | Punch, upbiting 60° forward, size 2 mm, working length 17 cm | 1
6 | scalpel, with telescopic blade, | 1
Consisting of:
Handle
outer tube
Micro-knife, pointed
7 | Ring- Curette, horizontal, round wire, ID 5 mm, long curved, with round handle, length 25 cm | 1
8 | Ring- Curette, round wire, ID 3 mm, tip angled 90°, with round handle, length 25 cm | 1
9 | Ring-Curette,round wire, ductile, ID 3 mm,tip angled 45°, with round handle, length 25 cm | 1
10 | Curette, round wire, ID 3 mm, distally curved shaft, with round handle, length 25 cm | 1
11 | Dissector, sharp, round spatula, tip angled 45°, size 3 mm, with round handle, length 25 cm | 1
12 | Bipolar Forceps, with fine jaws, width 1 mm, distally angled 45°, vertical closing, diameter 3.4 mm, working length 20 cm, | 1
Consisting of:
Bipolar Ring Handle
Outer Sheath
Inner Sheath
Forceps Insert
13 | Holding System, autoclavable, with fastener: KSLock, | 1
14 | Bipolar Forceps Insert, delicate, 1 mm, distally angled 45°, axial cliosing, axial closing, size 3 mm, length 20 cm | 1
15 | lens irrigation system for endoscopes, | 1
power supply 100 - 240 VAC, 50/60 Hz
Cconsisting of:
Mains Cord
One-Pedal Footswitch
Silicone Tubing Set, sterilizable
High Definition Camera System
1 | HD module, for use with up to 3 link modules, resolution 1920 x 1080 pixels, with integrated and digital Image | 1
Processing Module, power supply 100 - 120 VAC, 50/60 Hz; 200 - 240 VAC, 50/60 Hz
Consisting of:
Mains Cord, length 300 cm
DVI Signal Cable, length 300 cm
Connecting Cable, length 100 cm
USB Flash Drive,
2 | link module, for use with FULL HD threechip camera heads, power supply 100 - 120 VAC/200 - 240 VAC, 50/60 Hz, for | 1
use with HD
Consisting of:
Mains Cord, length 300 cm
Link Cable, length 20 cm
3 | Three-Chip FULL HD Camera Head, progressive scan, soakable, gas- and plasma-sterilizable, with integrated Parfocal | 1
Zoom Lens, focal length f = 15 - 31 mm (2x), 2 freely programmable camera head buttons
4 | Power LED 175 Light Source, with integrated high-performance LED and one light outlet, | 1
power supply 110 - 240 VAC, 50/60 Hz
Including:
Mains Cord
Connecting Cable
5 | Fiber Optic Light Cable, with straight connector, extremely heat-resistant, enhanced light transmission, diameter 4.8 mm, | 1
length 250 cm
6 | 27" FULL HD Monitor, color systems PAL/NTSC, max. screen resolution 1920 x 1080, image format 16:9, | 1
Interface: RS 232, power supply 85 - 264 VAC, 50/60 Hz, wall mount with VESA 100 adaptor
Including:
External 24 VDC Power Supply Mains Cord
7 | Trolley Powder Coated Rides on 4 Antistics, Dual Wheels, 2 Euipped with Locking (Front) 2 Fixed Shelves, 1 Drawer Unit | 1
with Lock (Locally Supply)
Serial No. 10
Endoscopic Spinal System
1 | Puncture Needle, including stylet, diameter 1.8 mm, working length 18 cm, with 1.3 mm opening for guide wire | 1
2 | Guide Wire, unsterile, diameter 1.2 mm, length 31 mm, package of 10 | 1
3 | Dilation Sleeve, OD5,2mm, ID 1.5 mm, graduated, length 23cm, colour code white | 1
4 | Dilation Sleeve, OD8.9 mm, ID 5.3 mm, graduated, length 21 cm, colour code yellow | 1
5 | Dilation Sleeve, graduated, inner diameter 9 mm, outer diameter 12.7 mm, length 19 cm, color code: orange | 1
6 | Dilation Sleeve, OD14.9mm, ID12.9mm, graduated, length 17cm, colour code red | 1
7 | Dilation Sleeve, OD16.9 mm, ID15.1 mm, graduated, length 15cm, colour code green | 1
8 | Dilation Sleeve, OD18.9 mm, ID17.1 mm, graduated, length 14 cm, colour code blue | 1
9 | Dilation Sleeve, OD20.9 mm, ID19 mm, graduated, length 13 cm, colour code black | 1
10 | Trocar, diameter 15 mm, working length 39 mm,for use with Trocar Attachment Telescope Holder and Forward-Oblique | 1
Telescope 25°, 4 mm, 6 cm color code: orange
11 | Trocar, diameter 19 mm, working length 40 mm, for use with Trocar Attachment | 1
12 | Trocar, diameter 23 mm, working length 43 mm, for use with Trocar Attachment Telescope Holder and Forward-Oblique | 1
Telescope 25°, 4 mm, 6 cm color code: black
13 | Trocar Attachment, diameter 15 mm, for use with Trocars | 1
14 | Telescope Holder, diameter 15 mm, for use with Trocar Attachment | 1
15 | Trocar Attachment, diameter 23 mm, for use with Trocars | 1
16 | Telescope Holder, diameter 23 mm, for use with Trocar Attachment | 1
17 | Trocar Attachment, diameter 19 mm, for use with Trocars | 1
18 | Telescope Holder, diameter 19 mm, for use with Trocar Attachment | 1
19 | Forward-Oblique Telescope 25°, eyepiece angled 90°, diameter 4 mm, length 6 cm, autoclavable, fiber optic light transmission | 1
incorporated, for use with 4 cm long Trocars color code: red to use with 4 cm long trocars
20 | Forward-Oblique Telescope 25°, eyepiece angled 90°, diameter 4 mm, length 11 cm, autoclavable, fiber optic light transmission | 1
incorporated, for use with 10 cm long Trocars color code: red to use with 10 cm long trocars
21 | Fiber Optic Light Cable, with straight connector, diameter 3.5 mm, length 230 cm | 1
22 | Bone Punch, dismantling, 90° upbiting, not through-cutting, 2 mm, working length 24 cm | 1
23 | Bone Punch, dismantling, 90° upbiting, not through-cutting, 4 mm, working length 24 cm | 1
24 | Bone Punch, dismantling, 40° upbiting, not through-cutting, 2 mm, working length 24 cm | 1
25 | Bone Punch, dismantling, 40° upbiting, not through-cutting, 4 mm, working length 24 cm | 1
26 | Punch, dismantling, bayonet-shaped, fixed, downbiting 40° forward, 2 mm, working length 17 cm | 1
27 | Punch, dismantling, bayonet-shaped, fixed, upbiting 40° forward, 2 mm, working length 17 cm | 1
28 | Hook Scissors, single action jaws, diameter 2.5 mm, working length 25 cm | 1
29 | Spoon Forceps, dismantling, robust, oval, spoon size 3 x 10 mm, single action jaws, working length 20 cm | 1
30 | Trocar, diameter 19 mm, working length 71 mm, for use with Trocar Attachment Telescope Holder and Forward-Oblique | 1
Telescope 25°, 4 mm, 9 cm color code: green
31 | Dissector n. dead hand, bayonet shaped, 3 mm, curved upwards, with round handle, sharp, working length 16 cm | 1
32 | Dissector, distal end tapered , bayonet-shaped, working length 15 cm | 1
33 | Nerve Hook, distal width 3 mm, bayonet-shaped, working length 16 cm | 1
34 | Nerve Hook, distal width 5 mm, bayonet-shaped, working length 16 cm | 1
35 | scalpel, with telescopic blade, | 1
36 | Suction Tube, with distal nerve retractor, with cut-off hole, LUER-Lock connector, diameter 2.7 mm, working length 15 cm | 1
37 | Nerve Retractor, angled 30°, distal width 5 mm, working length 17 cm | 1
38 | Dissector, sharp, tip angled 15°, with round handle, size 2 mm, length 25 cm | 1
39 | Palpation Hook, straight, distally 10 mm long and angled 90°, with ball, working length 20 cm | 1
40 | Nerve Retractor, hook length 2 mm, diameter 4 mm, angled sheath, working length 20 cm | 1
41 | Bipolar Coagulating Forceps, insulated, bayonet-shaped, tip 0.7 mm, length 23 cm, for use with bipolar high frequency cord | 1
42 | Bipolar Coagulating Forceps, insulated, bayonet-shaped, tip 1.2 mm, length 23 cm, for use with bipolar high frequency cord | 1
43 | Nasal Dressing Forceps, bayonet-shaped, length 20 cm | 1
44 | Plastic Container for sterilizing, suitable for steam, gas, and hydrogen peroxide (Sterrad®) sterilization and storage, | 1
perforated, with lid, external dimensions (w x d x h): 321 x 90 x 45 mm for use with two rigid endoscopes up to max. 20 cm
working length
45 | Rotation Socket to clamp on the operating table with one already mounted butterfly nut 28172 HRS, for use with European | 1
and United States standard rails, with lateral clamping element for height and angle adjustment of the articulated stand
46 | Articulated Stand, reinforced version, only, L-shaped, with one mechanical central clamp for all five joint functions, height 48 | 1
cm, operating range 52 cm, with fastener: KSLock (female)
47 | Bipolar High Frequency Cord, Length 300 cm, | 2
48 | Trocar, diameter 23 mm, working length 94 mm, for use with Trocar Attachment Telescope Holder and Forward-Oblique | 1
Telescope 25°, 4 mm, 11 cm color code: black
Serial No. 06
Neuro Drill
1 | NEURO motor system with color display, touch screen, two motor outputs, integrated irrigation pump | 1
and integrated SCB module, 100 - 240 VAC, 50/60 Hz
Consisting of:
Motor Unit
Mains Cord
Irrigator Rod
Two-Pedal Footswitch, two-stage, with proportional function
Silicone Tubing Set, for irrigation, sterilizable
Clip Set, for use with Tubing Set
Connecting Cable, length 100 cm
Single Use Tubing Set, sterile, package of 3
2 | High-Speed Micro-Motor, max. speed 60,000 rpm, including connecting cable, | 1
3 | High-Speed Handpiece, long, angled, 60,000 rpm, for use with High-Speed Micro-Motor | 1
4 | Craniotome Handpiece, maximum speed 60,000 rpm, including medium | 1
dura protector for use with High-Speed Micro-Motor as well as 3.17 mm Craniotome Burrs and Dura Protector
5 | High-Speed Craniotome Burr, medium, shaft diameter 3.17 mm, for single use, sterile, package of 5, for use with 60,000 rpm Craniotome | 1
Handpiece and Medium Dura Protector
Terms & Conditions
Manufacturer | Total FOR
Country of Origin | Rupees Five Million Seven Hundred Thirty Five Thousand One Hundred & Eighty Only.
Delivery Period
Terms Of Payment :
Prices
Validity
Warranty
Karl Storz GmbH & Co.KG.
Germany
12-14 weeks After receipt of order / Advance.
_____________________________ | 100% Advance.
S.Hussain Ali | In Pak Rs (FOR)
National Sales Manager | : 30 Days
Mob:- 0300-826881 | One year warranty against any manufacturing defect

## Specification (2) (hidden)
SPECIFICATIONS
S No. | Description | Qty
ANNEXURE A
Serial No. 3
Neuro Endoscope set with Vertebroscope.
Intercranial & Transphenoidal Endoscopy With Camera System
1 | Ventriculoscope Wide Angle Straigth Forward Telescope 6°, with angled eyepiece, autoclavable, | 1
O.D.: 6.1 mm, instrument channel diameter 2.9 mm, with suction and irrigation channel 1.6 mm, working length 18 cm, with
irrigation adapter, fiber optic light transmission incorporated, color code: green
2 | Operating sheath for ventriculoscope outer diameter 6.8 mm, working length 13.3 cm | 1
3 | Obturator | 1
4 | Click Line Scissors, single action jaws, pointed, diameter 1.3 mm, working length 30 cm, | 1
Consisting of:
Metal Handle, without ratchet
Outer Tube with insert
5 | Click Line Biopsy Forceps, both jaw parts movable, Ø 2.0 mm, working length 30 cm, | 1
Consisting of:
Metal Handle, without ratchet
LCA Outer Tube with insert
6 | Click Line Ventriculostomy Forceps, diameter 2.0 mm, working length 30 cm, | 1
Consisting of:
Metal Handle, without ratchet
Outer Tube with insert
7 | Click Line Grasping Forceps with teeth, 2.0 mm, working length 30 cm, | 1
Consisting of:
Metal Handle, without ratchet
Outer Tube with insert
8 | Puncture Needle | 1
9 | Coagulating Electrode, bipolar, 5 Fr. | 1
10 | Bipolar High Frequency Cord with 2 x 4 mm banana-plug to Coagulator Valleylab, length 300 cm | 1
11 | Holding System, autoclavable, with fastener: KSLock, | 1
Consisting of:
Rotation Socket to clamp on the operating table, for use with European and United States standard rails, with lateral clamping
element for height and angle adjustment of the articulated stand
Articulated Stand, reinforced version, straight, with one mechanical central clamp for all five joint functions, heigth 30 cm,
operating range 37 cm, with fastener: KSLock (female)
Clamping Jaw, metal, with axial intake, for use with instrument and telescope sheaths, clamping range 4.8 up to 12.5 mm, with
fastener: KSLock (male)
12 | Scissors, pointed, lightly curved jaws, double action jaws, diameter 1,7 mm, length 30 cm | 1
13 | Balloon Catheter, O.D. 0.7 mm, length 40 cm, sterile, single use, 10 pieces | 1
Terms & Conditions
Manufacturer | Total FOR
Country of Origin | Rupees Five Million Seven Hundred Thirty Five Thousand One Hundred & Eighty Only.
Delivery Period
Terms Of Payment :
Prices
Validity
Warranty
Karl Storz GmbH & Co.KG.
Germany
12-14 weeks After receipt of order / Advance.
_____________________________ | 100% Advance.
S.Hussain Ali | In Pak Rs (FOR)
National Sales Manager | : 30 Days
Mob:- 0300-826881 | One year warranty against any manufacturing defect

## Cardic bypass surgery S.1 (2)
CHANDKA MEDICAL COLLEGE HOSPITAL LARKANA
FINANCIAL BID
Supply, Installation, Testing, commissioning & Maintenance Durign maintenance period of Medical Equipment / Instruments under Revenue Component ADP Scheme No 371 Namely Construction of ESTABLISHMENT OF MADAM FARYAL TALPUR CARDIAC SURGERY COMPLEX HAVING FACILITIES OF :i.CORNARY ANGIOGRAPHY CARDIAC CATHETERIZED LAB ii. CARDIAC BYPASS SURGERY iii.CHEST SURGERY AT CMC HOSPITAL LARKANA FOR THE YEAR 2015-2016
1. List of Machinery Equipments (A)
S No. | Description | Qty | Unit Rate on CIF Basis Figure & Words | Unit Rate on FOR Basis Figure & Words | Total Cost (Figure & Words
`
1 | SINGLE PLANE DIGITAL CARDIAC CATH LAB/ANGIOGRAPHY SYSTEM | 1
The single-plane C-arm system for digital acquisition designed to meet the requirements of modern general angiography, cardiac angiography and interventional cardiac procedures with Flat Panel Detector. Consisting of:
2 | POSITIONING ARM:
· C-arm with motorized movements. Real time display of rotation angulations. The C-arm should have the possibility of head to toe coverage of the patient without repositioning the patient.
· Right anterior oblique (RAO) /Left Anterior Oblique (LAO): +/- 130° or more.
· Cranial / Caudal: Minimum + 55 / -45° or more (Capable of achieving Spider position).
· Rotation speed 25° / sec. or more in LAO /RAO
· Motorized parking (+/ - 120 deg. Motorized rotation of the C-arm)
· Integrated, computer-aided collision monitoring / Protection
· Programmable auto positioning of selected angulations. (50 or more programmable positions).
· Variable source-to-detector distance between 90 cm and 120 cm
3 | FLAT DETECTOR 30 x 30 cm
The high resolution dynamic flat panel detector (cardiac size) with integrated detachable grid especially designed to fulfill the requirements for diagnostic and interventional cardiology & radiology.
Size: Dedicated cardiac Flat Detector (30 cm x 30 cm)
Pixel size: 185 μm or less
Spatial resolution: 2.5 - 2.7 LP/mm or more.
Fluoroscopy as well as image acquisition to be done in 14 bit digitization depth.
Removable grid for dose saving in paediatric procedures. The gird to be preferably removed with push of a button:
4 | PATIENT SUPPORT / TABLE
· Floor mounted examination table for angiography and interventions.
· Motorized height adjustment with variable speed. Floating tabletop with longitudinal and transverse movements.
· Left / right pivotal table rotation +/- 90 deg. or more
· Patient weight bearing capacity 200 kg or more. Capability to handle additional load (100 kg) for CPR procedure.
Accessories
- Arm cradles (pair), Unilateral armrest, Infusion bottle holder,
5 | X-RAY GENERATOR
· Microprocessor based high frequency X-ray generator
· Output Power 100 KW. Radiographic rating minimum 1000 mA at 100 KV.
· The system should have capability of digital radiography and fluoroscopy. Should have capability of doing digital pulsed fluoroscopy at 10, 15 & 30 pulses / sec.
· Maximum continuous fluoro output power to be 3 KW or more to ensure good image quality during fluoro at oblique angles.
· Shortest Exposure time of less than 1msec with automatic exposure control.
· Should be capable of providing 225 mA or more current during digital pulsed fluoroscopy
6 | X-RAY TUBE
· Dual focus X-ray tube with anode heat storage capacity of at least 3.0 MHU or better
· Liquid bearing technology for longer durability and quiet operation.
· Dual Focus 0.4 mm, 0.8mm or better.
7 | DIGITAL IMAGING SYSTEM (ACQUISITION / FLUROSCOPY) |  |  |  | .
· Acquisition, storage and display in 1024 X 1024X 12 bits.
· Real time filtering, online edge enhancement, noise reduction (spatial filtration), pixel shift / remasking, and dynamic real time Pan / zoom.
· Hard Disk/Magnetic Disk Capacity for storage of 25,000 frames with 1024X1024X 10 bits matrix
· DICOM 3.0 with standard exchange media.
· The system must have DICOM send & DICOM Query / Retrieve facility
· Digital pulsed fluoroscopy with 7.5 or 10, 12.5 or 15 and 25 or 30 frames per second in 1024X1024X 10 bits or more.
· Display of scene directory for easy selection of any image or scene from the examination room or control room
· Radiation-free positioning of the primary and semi-transparent diaphragms by means of graphic display in the LIH (Last Image Hold).
· Variable Copper filtration during fluoroscopy and acquisition for radiation protection from 0.1 mm to 0.9 mm Cu. The selection of the Cu filters must be automatic by the system based on patient weight / absorption without any user interaction.
· Display of the measured dose-area product and the calculated patient air kerma reference on the flat screen display in ceiling suspension
· Monitoring of the skin entry dose, taking into account the geometric conditions of the system (device angulation, table position) to ensure that the skin entry dose applied to a specific region of the patient's body will not exceed a specified threshold.
· The system to display the dose information in DICOM format after every examination.
· Fluoro loop – specs can be added
· Scientific Coronary / Vessel analysis with determination of degree of stenosis, distance measurement and calibration.
· LVA analysis SW
· Dedicated processing software / Application to display image for enhanced stent visibility with display and control in control room as well as in examination room. Any related hardware, if needed is to be included in the offer. The application should be able to utilize online fluoro, stored fluoro runs or fluoro images for stent boosting. Online table-side processing like windowing, Image invert, Zoom / pan, edge enhancement. The application should be able to view images and fade vessels or stents in and out for vessel / stent comparison at the table side and control room.
· Automatic positioning of the c-arm corresponding to reference image and vice versa.
· Simultaneous display of fluoroscopy and reference images.
· Overlay fade feature i.e., on-line superposition of active fluoroscopy and reference image
· The system must have online image density (gray scale) correction i.e., Automatic online image density correction of dynamic scenes and single images for clear view in the bright and dark areas of the image.
· X-ray optimization before fluoro with built-in intelligence by using c-arm and table position data to calculate before taking an acquisition.
· Fast digital acquisition module for cardiac acquisition / angiography.
· Facility to review previous studies in the examination room from the patient’s old CD.
· The system should have the capability for retrieval of angio/CT/MR images back into the digital imaging system from the CDs and/or the network.
Controls of Digital Imaging System:
All controls of digital imaging system must be available in the examination as well as control room.
8 | CD / DVD recorder
CD/DVD drive for automatic digital image storing on CD-ROM for off-line data exchange in DICOM 3 format. The system must archive the images/scenes onto the CD in background, without blocking ongoing cine acquisition
9 | Networking
Network interface (100 BaseT) with the following integrated DICOM 3 services:
DICOM Send, DICOM Storage Commitment (StC), DICOM Query and DICOM Retrieve
10 | MONITORS FOR THE SYSTEM
· Active Matrix TFT monitors having high resolution
· Monitors should be ceiling-mounted in the operating/examination room.
· Large display of 30 or more inch active matrix digital TFT monitors for live and reference images in the operating/examination room. Resolution: 2560 x 1600 (pixels)
· 19 inch monitor active matrix TFT monitor for live live images in the control room. Maximum brightness to be 250 cd/m3 or more.
· Data for display of system and table geometry, system messages as well as dose is to be displayed either on the imaging monitors or on a separate monitor integrated into the ceiling suspension system.
11 | PHYSIOLOGICAL HAEMODYNAMIC MONITORING SYSTEM
Multichannel (20 channels or more) to record at least 4 channels IBP. Surface ECG in any configuration and simultaneous 12 lead ECG, NIBP and SaO2 measurement.
· Hemodymics system must be CE marked and FDA approved
· The system must be complete with software for pediatric & adult, right / left heart, angio / valvular homodynamic examinations including functionality for homodynamic calculations such as gradients, valve areas, shunt. Including annotations and 12 channel ECG.
· Digital display of all the parameters like IBP, Heart rate etc. It should be possible to print the waveforms simultaneously while acquiring the data in the background.
· It should be possible to store the waveforms on the hard disk / CD of the physiological recording system. The system should include following:
o 2 x invasive blood pressure reusable transducer.
o Reusable interface for 12 lead ECG.
o NIBP cuffs (2 Nos.)
o Holder for mounting the IBP transducer alongside the patient table.
· The system should have 40 GB or more hard disk for permanent storage.
· Two color monitors 19 inch Active matrix TFT inside the control room.
· One 21 inch monitor in the examination room that should be installed on the ceiling suspension with two monitors of imaging system.
· DVD / CD writer for archiving of study data and wave form stored on the hard disk.
· Facility of freezing the homodynamic data and simultaneous recovery of recent data / compare stored data with current waveform.
· Bi-directional communication between Haemodynamic system and Digital System of Main angio machine for transfer of patient demographic- and X-ray data such as Acquisition time, RAO/LAO angle, Cran./Caud. Angle, SID, Time of scene, Frame frequency, Total area dose and Fluoroscope time.
The Haemodynamic system must be from original manufacturer of angio system for smooth worfklow.
12 | Other accessories
Surgical shadow-less light ceiling suspended.
Ceiling suspended Lead Glass for Upper Body Radiation Protection
Lower body radiation protection flaps
1 x Fully programmable latest model contrast medium injector with 50 disposable syringes
13 | Other Allied Accessories
· One Postscript level Network Laser Printer for taking image printouts on paper. This printer is to be connected with the Main Digital Imaging System. Paper for 500 prints should be delivered with the printer.
· Lead Glass Window size 1.5 x 1 meter or more. Pb equivalence 2.1 mm or better
· 4 x Pb Aprons, double sided. Pb equivalence front 0.5 mm; back 0.35 mm
· 4 x Thyroid shields and 4 x Pb Glasses
· Intercom for communication between control and exam room
160 KVA UPS compatible for the whole system with back-up time of 10 minutes for fluoro and cine acquisition.
14 | OTHERS:
· 1 visit of application specialist for one week
· Warranty should be 02 years comprehensive including X-ray tube and Detector
15 | KNOCKOUT CLAUSES :
· QUOTING FIRMS/SYSTEMS MUST FULFILL THESE CONDITIONS OTHERWISE WILL NOT BE CONSIDERED FOR PURCHASE
· ANGIOGRAPHY SYSTEM MUST HAVE FDA & CE APPROVED, BOTH ARE REQUIRED
· QUOTING LOCAL FIRMS MUST HAVE SOLD ATLEAST 03 OTHER ANGIOGRAPHY SYSTEMS AT PAKISTAN AND MUST ATTACH SATISFACTORY LETTER FROM USERS
· SYSTEM MUST CARRY A TWO YEAR COMPREHENSIVE WARRANTY FROM THE MANUFACTURER INCLUDING SPARE PARTS AND SERVICES
Serial # 2
| Pacernmaker | 1
Serial # 3
1 | Ceilling LED OT Light Triple Dome | 1
Compact design with the significantly smaller & High performance LEDs.
Excellent color retention
Color temperature similar to day light
Extremely long life of the LEDs
Highest hygiene standards.
Maintenance and service costs are completely eliminated
Consisting of:
a | Main Dome : 160,000 lx | 1
b | Satalite Dome I : 160,000 lx | 1
c | Satalite Dome II : 160,000 lx | 1
TECHNICAL DATA
Central illuminance (at a distance of 1 m ) 160,000 lx
Turbulence degree (standard value: 37.5%) 0.27
Power input 45 W
Effective Life of light source 40,000 h
Central illuminance (electric diming from/to) 30-100% /5% in endo mode
Light field diameter (d10, distance of 1 m ) 170-220 mm
Color rendring index (CRI) [Ra (1-8)] 95
Permissible ambient temperature range 5° to + 40° C
during operation
Soft-start option (light does not start YES immediately with full brightness)
Electronic system meets requirements YES
acc. to VDE and IEC
Classification acc. to MPG YES
(Medical Devices Act / MDD)
Protection class acc. to IEC 601 I
Degree of protection: suspension system IP 30
Degree of protection: Light Head IP 42
Conformity CE
Standard Accessories are Included
a | LED LIGHT CAMERA | 1
120x zoom
HD-SDI digital video signal
Full HD video output For teleconsultations
No image interference caused by electro-surgery
Maximum image quality with 2 million pixels in full HD resolution
100% interference-free transmission to external storage media
Compatible with LED Ceiling OT Light
Compatible with DVD Recording System
TECHNICAL DATA OF HD CAMERA
Image Sensor 1/3" CMOS
Zoom 10x optical / 12x digita
Effective pixel number Approx. 2,000,000
Number of pixels/Resolution 1920 x 1080i
Focal length (zoom lens) f = 5.1 mm to 51 mm
White balance AUTOMATIC
Protection class I
Type B
Mark of conformity CE
Wall socket connector standard 3 x BNC socket for video signal
3x multiple DIN socket for control signals
Power requirements: control unit / camera 100-240 V, 50/60 Hz / 24 V AC/DC
b | Recording System | 1
Examinations & Minor procedures can be recorded in a cost-effective manner
Functions
Multi functional image & video recorder
Recording images & videos up yo full HD quality via camera
Enter patient data in combination with the keyboard
Storage of data on an external thumb drive or the internal data hard drive
Provision of data in the network
Video Signal input (DVI, SDI, S-VIDEO and Composite)
Recording Indicator
Recording indicator should disply in the upper-left corner of the monitor
Patient Data
Button on the front panel enables shuld direct access to the patient information screen
Image Options
Should be enable the section of image formates (e.g. JPG or TIFF)
Specification
Video Signal inputes: DVI, HD, SDI, Composite, S-Video, RGB, YPbPr
Video Out put resolution: 1920x1080, 1280x720, 1024x768, 800x600 640x480
Internal Hard drive 320 GB
USB Ports USB 2.0
Network RJ45/Connections as network drive (SMB)
Recording Formats Videos: H264mp4
Images: JPG, TIFF, BMP
Patient Data Saved as .txt file and / or in EXIF format
Consisting of:
19 Monitor (Local)
Key Board /Mouse (Local)
| Serial # 4 | 1
IABP (Intra Aortics Balloon Pump
| Serial # 5 | 1
Color Doppler Ultrasound with Imaging Modes System Software
and Transducer.
Main Features of System
15” TFT LCD or more Color Monitor
Operational Modes: B Mode, M Mode, Doppler Mode, Color Flow Maping, BFLOW / e flow
Color Flow Mapping in M Mode, HPRF, Duplex/Triplex modes
Power Doppler, Doppler, Directional Power Doppler.
Frame Rate 745 FPS or more
Scanning Dept: 30 cm or more
System Advance Application Software:
Compound precision Imaging /Cross Beam/M-View
System Should Be Capable Of Following Up gradation:
B flow/eflow
Probes
MULTI-FREQUENCY CONVEX PROBE
MULTI-FREQUENCY LINEAR PROBE
MULTI-FREQUENCY ENDOCAVITY PROBE
Other Third Party Accessories UPS
Imported UPS with8-10 minutes backup.
Digital B&W Thermal Printer
| SIGNATURE CONTRACTOR NAME |  |  | DATE & SEAL

## Neruo Med Others
CHANDKA MEDICAL COLLEGE HOSPITAL LARKANA
Supply of Other Items ADP Scheme No 372 Namely Construction of Neuro Science Complex LIST OF MACHINERY / EQUIPMENTS FOR NEURO SURGERY DEPARTMENT & NEURO MEDICINE, CMC HOSPITAL LARKANA FOR THE YEAR 2015-2016
1.List of Other Items For Neruo Complex
S No. | Description | Qty | Unit Rate on CIF Basis Figure & Words | Unit Rate on FOR Basis Figure & Words | Total Cost (Figure & Words
Serial # 1
| Air Condition Window 1.5 Ton With Steblizer | 20
Serial # 2
| Air Condition Split 1.5 Ton with Stabilizer | 12
Serial # 3
| Electric Water Cooler | 2
Serial # 4
| Refigirator large Size | 4
Serial # 5
| TV Set 32" LED | 2
Serial # 6
| Computer Dual Core Processor with Printer (Latest Model). | 6
Serial # 7
| Multi Media With Screen | 2
Serial # 8
| Lap Top 21" | 2
Serial # 1
| Stand By Generator 100 KV | 1
400 V, 3 Phase+N, 0.8 P.F
50 Hz, 1500 RPM
Water Cooled
Radiator 50 C Tropically desinged
Industrial muffler (Supplied loose)
Exhaust Flexible (supplied loose)
Base Fue Tank (Sufficeint for 8 house capacity)
Engine Sump oil drain tap & coolant Drain Tap
IP21 Protection Alternator classs H insulation
Emergency Push button
Set mounted lead acid batteries
Earth Leads
Sound and weather Proff canopy 100/200 KVA DG set with following specifiication :
Fabricated from Hot Rolled AISI 1010 steel, 12/14 SWG
Modular Design for ease of maintenance
Imported accessories inculding stainless, steel hinges, die cast hinges, galvanized fasteners (class 8.8), etc.
Acoustic, high density, open cellular fire Proof foam.
Standard cable access window
Standard control panel veiwing window
Optional
ATS Panel
Power cable 70 mm / 95 mm Copper 30 meters
Manual change Over
Including:
Installation (Fitting , Fixing & Civil Work)
| SIGNATURE CONTRACTOR NAME |  | DATE & SEAL

## Cardiac Other Item
CHANDKA MEDICAL COLLEGE HOSPITAL LARKANA
Supply of Other Items under Revenue Component ADP Scheme No 371 Namely Construction of ESTABLISHMENT OF MADAM FARYAL TALPUR CARDIAC SURGERY COMPLEX HAVING FACILITIES OF :i.CORNARY ANGIOGRAPHY CARDIAC CATHETERIZED LAB ii. CARDIAC BYPASS SURGERY iii.CHEST SURGERY AT CMC HOSPITAL LARKANA FOR THE YEAR 2015-2016
List of others Items For Cardiac Department
S No. | Description | Qty | Unit Rate on CIF Basis Figure & Words | Unit Rate on FOR Basis Figure & Words | Total Cost (Figure & Words
Serial # 1
| Air Condition (Window) 1.5 Tone With Steblizer | 15
Serial # 2
| Air Condition Split 1.5 ton with Stabilizer | 10
Serial # 3
| Electric Water Cooler | 4
Serial # 4
| Refigirator | 4
Serial # 5
| TV Set 32" | 4
Serial # 6
| Computer Dual Core Processr With Laser Printer (latest) | 6
Serial # 7
| Multi Media with Screen | 1
Serial # 8
| Laptop 21" (Latest) | 1
Serial # 9
| Library Books | Lump Sum
Serial # 10
| Ambulance (Company Fitted) | 1
with Vantilator Oxygen Lan Pulse Oximeter and Full ICU equipment
Serial # 12
| Generator 200 KV (Auto Start) | 1
400 V, 3 Phase+N, 0.8 P.F
50 Hz, 1500 RPM
Water Cooled
Radiator 50 C Tropically desinged
Industrial muffler (Supplied loose)
Exhaust Flexible (supplied loose)
Base Fue Tank (Sufficeint for 8 house capacity)
Engine Sump oil drain tap & coolant Drain Tap
IP21 Protection Alternator classs H insulation
Emergency Push button
Set mounted lead acid batteries
Earth Leads
Sound and weather Proff canopy 100/200 KVA DG set with following specifiication :
Fabricated from Hot Rolled AISI 1010 steel, 12/14 SWG
Modular Design for ease of maintenance
Imported accessories inculding stainless, steel hinges, die cast hinges, galvanized fasteners (class 8.8), etc.
Acoustic, high density, open cellular fire Proof foam.
Standard cable access window
Standard control panel veiwing window
Optional
ATS Panel
Power cable 70 mm / 95 mm Copper 30 meters
Manual change Over
Includings:
Fitting , Fixing & Civil Work
| SIGNATURE |  | DATE & SEAL

## Cardiac Others
CHANDKA MEDICAL COLLEGE HOSPITAL LARKANA
Supply, Installation, Testing, commissioning & Maintenance Durign maintenance period of Medical Equipment / Instruments under Revenue Component ADP Scheme No 371 Namely Construction of ESTABLISHMENT OF MADAM FARYAL TALPUR CARDIAC SURGERY COMPLEX HAVING FACILITIES OF :i.CORNARY ANGIOGRAPHY CARDIAC CATHETERIZED LAB ii. CARDIAC BYPASS SURGERY iii.CHEST SURGERY AT CMC HOSPITAL LARKANA FOR THE YEAR 2015-2016
1.List of Machinery Equipment For Onclolgy Treatment Ward
S No. | Description | Qty | Unit Rate on CIF Basis Figure & Words | Unit Rate on FOR Basis Figure & Words | Total Cost (Figure & Words
Serial # 1
| Ambulance (Company Fitted) | 1
Serial # 2
| Generator 200 KV (Auto Start) | 1
Includings:
Fitting , Fixing & Civil Work
| SIGNATURE CONTRACTOR NAME |  | DATE & SEAL

## Neruo Med (Furniture)
CHANDKA MEDICAL COLLEGE HOSPITAL LARKANA
Supply, Installation, Testing, commissioning & Maintenance Durign maintenance period of Medical Equipment / Instruments under Revenue Component ADP Scheme No 372 Namely Construction of Neuro Science Complex LIST OF MACHINERY / EQUIPMENTS FOR NEURO SURGERY DEPARTMENT & NEURO MEDICINE, CMC HOSPITAL LARKANA FOR THE YEAR 2015-2016
2.List of Furnitur Fixture For Neuro Medicine Department
S No. | Description | Qty | Unit Rate on CIF Basis Figure & Words | Unit Rate on FOR Basis Figure & Words | Total Cost (Figure & Words
Serial # 1
| Fowler Beds | 40
Bed Semi Fowler, Size: 3' x 6 1/2' (Japan Design) Double Frame, Powder Coated, 16 gauge.
Mattress base having semi- fowlder adjustment facility. 2-section, semi folwer position adustable with a screw crank, Lying area made from reinfocrced steel strip 18 gauzge. Mobile on 4-casters dia 125mm (two Casters with brakes) fitted on recessed legs.
Serial # 2
| Patients Examination Table | 4
Examination coach construced with 2 inch foam cushion square mild steel pipe 16 gauge.
Adjustable backrest with esy clutch mechanism.
Finished in epoxy powder coating .
Serial # 3
| Officer Table 4 X 6 | 10
Serial # 4
| Revolving Chair | 10
Serial # 5
| Chairs (Saminor Hall, Library, Offices) | 60
Serial # 6
| Bed Side Lockers | 40
Bed side locker, U type, Mild steel pipe 16 gauge MDF Top Size 16" x 14", Base on Four Legs, with One drawer & One Cupboard, & Open Space Between drawer & Cupboard, Powder coated.
Serial # 7
| Bed Side Benches | 40
Serial # 8
| Iron Cupboard | 10
Serial # 9
| Nursing Counter | 4
Serial # 10
| Stools (S-Steel) | 10
Revolving Stool, stand on four legs,
Frame painted powder coated with S.S Top 13"
Serial # 11
| IV Stands (S-Steel) | 20
Serial # 12
| Foam Mattress | 40
3x 6.5 x 4"
Serial # 13
| Blanket (Red) 60X 90 | 40
Serial # 14
| Library Books Rack | 6
| SIGNATURE |  | DATE & SEAL

## Neruo (Furniture)
CHANDKA MEDICAL COLLEGE HOSPITAL LARKANA
Supply, Installation, ADP Scheme No 372 Namely Construction of Neuro Science Complex LIST OF MACHINERY / EQUIPMENTS FOR NEURO SURGERY DEPARTMENT & NEURO MEDICINE, CMC HOSPITAL LARKANA FOR THE YEAR 2015-2016
1.List of Furnitur Fixture For Neuro Surgery Department
S No. | Description | Qty | Unit Rate on CIF Basis Figure & Words | Unit Rate on FOR Basis Figure & Words | Total Cost (Figure & Words
Serial # 1
| Fowler Beds | 40
Bed Semi Fowler, Size: 3' x 6 1/2' (Japan Design) Double Frame, Powder Coated, 16 gauge.
Mattress base having semi- fowlder adjustment facility. 2-section, semi folwer position adustable with a screw crank, Lying area made from reinfocrced steel strip 18 gauzge. Mobile on 4-casters dia 125mm (two Casters with brakes) fitted on recessed legs.
Serial # 2
| ICU Beds Imported | 6
Bed full fowler, Size: 3' x 6 1/2 '
Double Frame (Japan Design),
Frame made from rectangular steel 1 1/4 " x 2 1/2 x 16 Gauge, Powder coated.
Mattress top having full fowler adjustment with the help of 2- screw cranks,
vascular Position.
Lying area made from reinforced steel strips 3" x 18 Gauge.
Head / foot panels laminated,
Mobile on 4 Casters 125 mm dia, 2-caster with brakes. Fitted on recessed legs.
Serial # 3
| Dressing Table | 2
Serial # 4
| Officer Table 4 X 6 wooden | 8
Serial # 5
| Revolving Chair wooden | 8
Serial # 6
| Simple Chair wooden | 40
Serial # 7
| Bed Size Locker | 46
Bed side locker, U type, Mild steel pipe 16 gauge MDF Top Size 16" x 14", Base on Four Legs, with One drawer & One Cupboard, & Open Space Between drawer & Cupboard, Powder coated.
Serial # 8
| Bed Side Benches | 46
Serial # 9
Iron Cupboards
Serial # 10
| Nursing Counter | 4
Serial # 11
| Iron Stools | 8
Serial # 12
| IV Stand (S-Steel) | 20
Serial # 13
| Foam Matress 3x 6 .5 x 4" | 46
Serial # 13
| Blankets (Red) 60X90 | 46
| SIGNATURE CONTRACTOR NAME |  | DATE & SEAL

## Furniture Cardiac
CHANDKA MEDICAL COLLEGE HOSPITAL LARKANA
Supply, of Furniture Fixture ADP Scheme No 371 Namely Construction of ESTABLISHMENT OF MADAM FARYAL TALPUR CARDIAC SURGERY COMPLEX HAVING FACILITIES OF :i.CORNARY ANGIOGRAPHY CARDIAC CATHETERIZED LAB ii. CARDIAC BYPASS SURGERY iii.CHEST SURGERY AT CMC HOSPITAL LARKANA FOR THE YEAR 2015-2016
1. List of Furniture Fixture For Cardiac Department
S No. | Description | Qty | Unit Rate on CIF Basis Figure & Words | Unit Rate on FOR Basis Figure & Words | Total Cost (Figure & Words
Serial # 1
| Fowler Beds (Surgical) | 26
Bed Semi Fowler, Size: 3' x 6 1/2' (Japan Design) Double Frame, Powder Coated, 16 gauge.
Mattress base having semi- fowlder adjustment facility. 2-section, semi folwer position adustable with a screw crank, Lying area made from reinfocrced steel strip 18 gauzge. Mobile on 4-casters dia 125mm (two Casters with brakes) fitted on recessed legs.
Serial # 2
| ICU Beds (Imported) | 6
Bed full fowler, Size: 3' x 6 1/2 '
Double Frame (Japan Design),
Frame made from rectangular steel 1 1/4 " x 2 1/2 x 16 Gauge, Powder coated.
Mattress top having full fowler adjustment with the help of 2- screw cranks,
vascular Position.
Lying area made from reinforced steel strips 3" x 18 Gauge.
Head / foot panels laminated,
Mobile on 4 Casters 125 mm dia, 2-caster with brakes. Fitted on recessed legs.
Serial # 3
| Foam Mettress | 32
3x 6.5 x 4".
Serial # 4
| Blankets (Red) 60X 90 | 32
Serial # 5
| Chair For Saminor Hall | 30
Serial # 6
| Table For Saminor Hall (Large Size) | 2
Serial # 7
| Library Chair | 20
Serial # 8
| Library Books Rack | 8
Serial # 9
| Laboratory Chair | 8
Serial # 10
| Laboratory Tables (Medium Size) | 2
Serial # 11
| Dressing Table | 4
Aseptic / surgical / dressing Table, Frame P.C Top /S/S (Grade 202, Table Size 78" x 32" head portion pull with s/s handle, with 1 s/s (202 grade) Drawer with lock, mobile on 4 " wheel 2 with breaks.
Serial # 12
| Officer Table 4 X 6 | 10
Serial # 13
| Revolving Chair | 10
Serial # 14
| Office Chairs | 60
Serial # 15
| Bed Side Lockers | 32
Bed side locker, U type, Mild steel pipe 16 gauge MDF Top Size 16" x 14", Base on Four Legs, with One drawer & One Cupboard, & Open Space Between drawer & Cupboard, Powder coated.
Serial # 16
| Bed Size Benches | 32
Serial # 17
| Iron Cupboards | 10
Serial # 18
| Nursing Counter | 4
Serial # 19
| Stools (S-Steel) | 20
Stool Stand on four legs,
Frame Painted powder coated with S.S Top 13"
Serial # 20
| IV Stands (S-Steel) | 40
Serial # 21
| Patient Stretures | 4
Stretcher Trolley, Powder Coated,
Main Frame made of Rectangular Mild Steel Tube,
with removable stretcher Top, (Frame : Powder coated & Top
Sheet S/S ).
With Foam Mattress 3".
Collapisble Guard,
with Cylinder cage, with I.V Pool,
Easy Mobility on four 175 mm (7") Dia wheels.
| SIGNATURE CONTRACTOR NAME |  | DATE & SEAL

## Furniture G.S
CHANDKA MEDICAL COLLEGE HOSPITAL LARKANA
Supply, Installation, ADP Scheme No 373 Namely Up gradation Surgery Department Up gradation Surgery Department O.T, Modernization of Laboratory, Blood Bank & Establishment of Oncology Treatment Ward at CMCH LARKANA, FOR THE YEAR 2015-2016
1.List Of Furniture Fixter For Up gradation Surgery Department
S No. | Description | Qty | Unit Rate on CIF Basis Figure & Words | Unit Rate on FOR Basis Figure & Words | Total Cost (Figure & Words
Serial # 1
| Patient Examination Coach | 5
Examination coach construced with 2 inch foam cushion square mild steel pipe 16 gauge.
Adjustable backrest with esy clutch mechanism.
Finished in epoxy powder coating .
Serial # 2
| Officer Table 4 X 6 | 5
Serial # 3
| Sofa Set (Professor Office) | 2
Serial # 4
| Revolving Chair | 5
Serial # 5
| Simple Chair | 30
Serial # 6
| Iron Cupboards | 5
Serial # 7
| Iron Stool (S-Steel) | 4
Serial # 8
| Air Condition 1.5 ton window with stabilizers | 8
Serial # 9
| IV Stand (Iron) | 8
Serial # 10
| Wheel Chair | 4
| SIGNATURE CONTRACTOR NAME |  | DATE & SEAL

## Furniture Oncology
CHANDKA MEDICAL COLLEGE HOSPITAL LARKANA
Supply, Installation, of FURNITURE & FIXTURE ADP Scheme No 373 Namely Up gradation Surgery Department Surgery Equipment For O.T, Modernization of Laboratory, Blood Bank & Establishment of Oncology Treatment Ward at CMCH LARKANA, FOR THE YEAR 2015-2016
1.List Of Furniture Fixter For Onclolgy Treatment Ward
S No. | Description | Qty | Unit Rate on CIF Basis Figure & Words | Unit Rate on FOR Basis Figure & Words | Total Cost (Figure & Words
Serial # 1
| Fowler Beds | 40
Fowler Beds
Bed Semi Fowler, Size: 3' x 6 1/2' (Japan Design) Double Frame, Powder Coated, 16 gauge.
Mattress base having semi- fowlder adjustment facility. 2-section, semi folwer position adustable with a screw crank, Lying area made from reinfocrced steel strip 18 gauzge. Mobile on 4-casters dia 125mm (two Casters with brakes) fitted on recessed legs.
Serial # 2
| Foam Mettress | 40
3x 6.5 x 4"
Serial # 3
| ICU Beds | 6
ICU Beds Imported
Bed full fowler, Size: 3' x 6 1/2 '
Double Frame (Japan Design),
Frame made from rectangular steel 1 1/4 " x 2 1/2 x 16 Gauge, Powder coated.
Mattress top having full fowler adjustment with the help of 2- screw cranks,
vascular Position.
Lying area made from reinforced steel strips 3" x 18 Gauge.
Head / foot panels laminated,
Mobile on 4 Casters 125 mm dia, 2-caster with brakes. Fitted on recessed legs.
Serial # 4
| Blankets (Red) 60X 90 | 40
Serial # 5
| Patient Examination Coach | 4
Examination coach construced with 2 inch foam cushion square mild steel pipe 16 gauge.
Adjustable backrest with esy clutch mechanism.
Finished in epoxy powder coating .
Serial # 6
| Officer Table 6 X 4 | 4
Serial # 7
| Revolving Chair | 4
Serial # 8
| Simple Chairs wooden | 50
Serial # 9
| Bed Side Lockers | 40
Bed side locker, U type, Mild steel pipe 16 gauge MDF Top Size 16" x 14", Base on Four Legs, with One drawer & One Cupboard, & Open Space Between drawer & Cupboard, Powder coated.
Serial # 10
| Bed Side Benches | 40
Serial # 11
| Iron Cupboards | 4
Serial # 12
| Chairs For Library & Saminor | 40
Serial # 13
| Library Tables | 2
Serial # 14
| Library Books | 1
Serial # 15
| Nursery Courter | 4
| SIGNATURE CONTRACTOR NAME |  | DATE & SEAL

## Oncology Equipments
CHANDKA MEDICAL COLLEGE HOSPITAL LARKANA
Supply, Installation, Testing, commissioning & Maintenance Durign maintenance period of Medical Equipment / Instruments under Revenue Component ADP Scheme No 373 Namely Up gradation Surgery Department Surgery Equipment For O.T, Modernization of Laboratory, Blood Bank & Establishment of Oncology Treatment Ward at CMCH LARKANA, FOR THE YEAR 2015-2016
1.List of Machinery Equipment For Onclolgy Treatment Ward
S No. | Description | Qty | Unit Rate on CIF Basis Figure & Words | Unit Rate on FOR Basis Figure & Words | Total Cost (Figure & Words
Serial # 1
| Centeral Oxygen Supply System | 1
Serial # 2
| Suction Machine | 4
Serial # 3
| Nebulizer | 2
Serial # 4
| Oncobay | 1
Serial # 5
| Drug Tray | 6
Serial # 6
| Air Condition (Window 1 1/2 With Steblizer | 8
Serial # 7
| Electric Water Cooler | 2
Serial # 8
| Refrigerator | 1
Serial # 9
| Computer With Laser Printer | 1
Serial # 10
| Stand By Generator 100 KV for O.T Complex Separately | 1
Including:
Generator Installation
Fitting , Fixing & Civil Work
| SIGNATURE CONTRACTOR NAME |  | DATE & SEAL

## Annexure A Section (A)
CHANDKA MEDICAL COLLEGE HOSPITAL LARKANA
FINANCIAL BID
Supply, Installation, Testing, commissioning & Maintenance Durign maintenance period of Medical Equipment / Instruments under Revenue Component ADP Scheme No 372 Namely Construction of Neuro Science Complex LIST OF MACHINERY / EQUIPMENTS FOR NEURO SURGERY DEPARTMENT & NEURO MEDICINE, CMC HOSPITAL LARKANA FOR THE YEAR 2015-2016
(1). List of Neuro Surgery Equipments (A)
S No. | Description | Qty | Unit Rate on CIF Basis Figure & Words | Unit Rate on FOR Basis Figure & Words | Total Cost (Figure & Words
Serial No.01
| Craniotomy Instrument Sets. | 2 Set
1 | FORCEPS, FOERSTER, SERR., STR., 25 CM | 4
2 | TOWEL FORCEPS, BACKHAUS, SHARP, 13 CM | 1
3 | SCALPEL HANDLE, NO. 3, 12 CM | 1
4 | SCALPEL HANDLE, NO. 4, 13.5 CM | 1
5 | SCALPEL HANDLE, NO. 7, SOLID, 16 CM | 1
6 | TC-SCISSORS, MAYO, STR., 17 CM | 1
7 | TC-SCISSORS, MAYO, CVD., 17 CM | 1
8 | TC-DISS. SCISSORS, FINE, CVD., 18 CM | 1
9 | DURA SCISSORS, SCHMIDEN-TAYLOR, 16.5 CM | 1
10 | OPERATING SCISSORS, SH/BL, STR., 15,5CM | 2
11 | DRESSING FORCEPS, ADSON, 12 CM | 2
12 | FORCEPS, MINI-ADSON, 1X2 T., 12 CM | 2
13 | FORCEPS, KOCHER, 1X2 T., STR., 14 CM | 2
14 | DRESSING FORCEPS, CUSHING, STR., 17.5 CM | 2
15 | TISSUE FORCEPS, CUSHING, 1X2 T., 18 CM | 1
16 | DRESSING FORCEPS, POTTS-SMITH, 18 CM | 2
18 | FORCEPS MICRO-MOSQUITO, STR., 12 CM | 2
19 | FORCEPS MICRO-MOSQUITO, CVD., 12 CM | 3
20 | FORCEPS, DANDY, CVD. SIDEW., 14.5 CM | 1
21 | BONE RONGEUR, LUER, CVD., 18 CM | 1
22 | BONE RONGEUR, ZAUFAL-JANSEN, 18 CM | 1
23 | BONE RONGEUR, STR., 22.5 CM | 1
24 | BONE RONGEUR, CVD., 22.5 CM | 1
25 | PITIUTARY FCPS., OLDBERG, Ø 8 MM, 25 CM | 1
26 | WOUNDSPREADER, SHARP, 3X4 T., 20 CM | 1
27 | SPREADER, JANSEN, BLUNT, 3X3 T., 10 CM | 1
28 | HOOKLET, CUSHING, 14 MM, 24 CM | 2
29 | PERIOSTEAL ELEVATOR DR.IGLHAUT, 18 CM | 2
30 | BONE CURETTE, BRUNS, OVAL, NO. 1, 23 CM | 2
31 | BONE CURETTE, BRUNS, OVAL, NO. 2, 23 CM | 1
32 | BONE CUR., BRUNS, OVAL, NO. 000, 23 CM | 1
33 | BURR, CONE SHAPE, HUDSON, NO. 3 | 1
34 | WIRE SAW CONDUCTOR, DE MARTEL | 2
35 | GALEA RETRACTOR, (HOOK) YASARGIL, 41 CM | 2
36 | DURA HOOKLET, FRAZIER, SHARP, 13 CM | 1
37 | SUCTION TUBE, FRAZIER, 2.0 MM, 19.5 CM | 1
38 | DURA DISSECTOR, PENFIELD, NO. 1, 17.5 CM | 1
39 | DURA DISSECTOR, PENFIELD, NO. 3, 19.5 CM | 1
40 | NEEDLEHOLDER, ADSON, FENESTR., 18.5 CM | 1
41 | CONTAINER MS, 60X30X14 CM, HANDLE GREY | 1
42 | COLOR-TAG, RED | 1
43 | CODING LABEL, MS2, W. TEXT, W/O HOLE | 1
44 | TRAY 1/1, 535x255x73 MM
Serial No.02
| Laminectomy Instrument Sets. | 2 Set
1 | SCALPEL HANDLE, NO. 3, 12 CM | 1
2 | SCALPEL HANDLE, NO. 4, 13.5 CM | 1
3 | TOWEL FORCEPS, BACKHAUS, SHARP, 13 CM | 5
4 | SPONGE FORCEPS, RAMPLEY, STR., 24.5 CM | 2
5 | WATSON-CHEYNE DISSECTOR 18 CM | 1
6 | DRESSING FORCEPS, POTTS-SMITH, 18 CM | 2
7 | TISSUE FORCEPS, 1X2 T., SLIM, 18 CM | 1
8 | Diss. forceps, Waugh, 1x2 T., 20 cm | 1
9 | DRESSING FORCEPS, CUSHING, STR., 17.5 CM | 1
10 | DISSECT. SCISSORS, BL/BL, STR., 14.5 CM | 1
11 | DISSECT. SCISSORS, BL/BL, CVD., 14.5 CM | 1
12 | DISSECTING SCISSORS, CVD., 18 CM | 1
13 | DISSECTING SCISSORS, CVD., 20.5 CM | 1
14 | SPREADER, ADSON-ANDERSON, 4X4 T., 19 CM | 2
15 | WOUNDSPREADER, TRAVERS, BLUNT, 21.5 CM | 1
16 | SPREADER, BECKM., SEMISH., 4X4 T., 31 CM | 1
17 | BONE RONGEUR, STILLE-RUSKIN, 23.5 CM | 1
18 | BONE RONGEUR, STR., 22.5 CM | 1
19 | BONE RONGEUR, CVD., 22.5 CM | 1
20 | BONE RONGEUR, STILLE, CVD., 22.5 CM | 1
21 | BONE RONGEUR, WIDE, 24.5 CM | 1
22 | Punch, 40°, 2mm, 18cm | 1
23 | Punch, 40°, 3mm, 18cm | 1
24 | Punch, 40°, 4mm, 18cm | 1
25 | Punch, 40°, 5mm, 18cm | 1
26 | LAMINECT. FORCEPS, 2X10 MM, STR., 15 CM | 2
27 | LAMINECT. FORCEPS, 4X10 MM, STR., 15 CM | 2
28 | LAMINECT. FORCEPS, 3X10 MM, UPW., 15 CM | 1
29 | LAMINECT. FORCEPS, 4X10 MM, UPW., 15 CM | 1
30 | LAMINECT. FORCEPS, 3X10 MM, DOWN, 15 CM | 1
31 | LAMINECT. FORCEPS, 4X10 MM, DOWN, 15 CM | 1
32 | LAMINECT. FORCEPS, 2X10 MM, STR., 18 CM | 1
33 | LAMINECT. FORCEPS, 4X10 MM, STR., 18 CM | 2
34 | LAMINECT. FORCEPS, 3X10 MM, UPW., 18 CM | 1
35 | LAMINECT. FORCEPS, 4X10 MM, UPW., 18 CM | 1
36 | LAMINECT. FORCEPS, 3X10 MM, STR., 18 CM | 1
37 | LAMINECT. FORCEPS, 4X10 MM, DOWN, 18 CM | 1
38 | OSTEOTOME, STILLE, FINE, 25 MM, 20.5 CM | 1
39 | SUCT. TUBE, FERGUSSON, WL:110 MM, 2.0 MM | 1
40 | SUCT. TUBE, FERGUSSON, WL:110 MM, 2.5 MM | 1
41 | SUCT. TUBE, FERGUSSON, WL:110 MM, 3.0 MM | 1
42 | SUCT. TUBE, FERGUSSON, WL:110 MM, 4.0 MM | 1
43 | SUCT. TUBE, FERGUSSON, WL:110 MM, 4.0 MM | 1
44 | NEEDLEHOLDER, MAYO-HEGAR, 16 CM | 1
45 | NEEDLEHOLDER, MAYO-HEGAR, 18.5 CM | 1
46 | NEEDLEHOLDER, MAYO-HEGAR, 20.5 CM | 1
47 | TC-NEEDLEHOLDER, CRILE-WOOD, 15 CM | 1
48 | TC-Needleholder, Crile-Wood, str., 18 cm | 1
49 | TC-NEEDLEHOLDER, CRILE-WOOD, 20.5 CM | 1
50 | TC-NEEDLEHOLDER, DE BAKEY, 18 CM | 1
51 | TRAY DIN, 480x255x73 MM | 1
52 | CONTAINER MS, 60X30X14 CM, HANDLE GREY | 1
53 | COLOR-TAG, RED | 1
54 | CODING LABEL, MS2, W. TEXT, W/O HOLE | 1
55 | Tray 1/1, 535x255x73 mm | 1
Serial No.03
Neuro Endoscope with Vertebroscope.
Endoscopic Pituatry Surgery Set
1 | Straight Forward Telescope 0°, enlarged view, diameter 4 mm, length 18 cm, autoclavable. Fiber optic light | 1
transmission incorporated.
2 | Suction and Irrigation Sheath 0°, outer diameter 4.8 x 6 mm, working length 14 cm, for use with Telescope | 1
lens irrigation system
3 | Forward-Oblique Telescope 30°, enlarged view, diameter 4 mm, length 18 cm, autoclavable. Fiber optic light | 1
transmission incorporated.
4 | Suction and Irrigation Sheath 30°, outer diameter 4.8 x 6 mm, working length 14 cm, for use with Telescope | 1
and lens irrigation system
5 | Punch, upbiting 60° forward, size 2 mm, working length 17 cm | 1
6 | scalpel, with telescopic blade, ( Handle, Outer, Micro-Knife Pointed ) | 1
7 | Ring- Curette, horizontal, round wire, ID 5 mm, long curved, with round handle, length 25 cm | 1
8 | Ring- Curette, round wire, ID 3 mm, tip angled 90°, with round handle, length 25 cm | 1
9 | Ring-Curette,round wire, ductile, ID 3 mm,tip angled 45°, with round handle, length 25 cm | 1
10 | Curette, round wire, ID 3 mm, distally curved shaft, with round handle, length 25 cm | 1
11 | Dissector, sharp, round spatula, tip angled 45°, size 3 mm, with round handle, length 25 cm | 1
12 | Bipolar Forceps, with fine jaws, width 1 mm, distally angled 45°, vertical closing, diameter 3.4 mm, working length 20 cm, | 1
( Bipolar Ring Handle, Outer Sheath, Inner Sheath, Forcef Insert )
13 | Holding System, autoclavable, with fastener: KSLock, | 1
14 | Bipolar Forceps Insert, delicate, 1 mm, distally angled 45°, axial cliosing, axial closing, size 3 mm, length 20 cm | 1
15 | lens irrigation system for endoscopes, | 1
power supply 100 - 240 VAC, 50/60 Hz
Cconsisting of:
Mains Cord
One-Pedal Footswitch
Silicone Tubing Set, sterilizable
Intercranial & Transphenoidal Endoscopy With Camera System
1 | Ventriculoscope Wide Angle Straigth Forward Telescope 6°, with angled eyepiece, autoclavable, | 1
O.D.: 6.1 mm, instrument channel diameter 2.9 mm, with suction and irrigation channel 1.6 mm, working length
18 cm, with irrigation adapter, fiber optic light transmission incorporated
2 | Operating sheath for ventriculoscope outer diameter 6.8 mm, working length 13.3 cm | 1
3 | Obturator | 1
4 | Click Line Scissors, single action jaws, pointed, diameter 1.3 mm, working length 30 cm, | 1
( Metal Handle, without ratchet, Outer Tube with Insert )
5 | Click Line Biopsy Forceps, both jaw parts movable, Ø 2.0 mm, working length 30 cm, | 1
Metal Handle, without ratchet LCA Outer Tube with Insert )
6 | Click Line Ventriculostomy Forceps, diameter 2.0 mm, working length 30 cm, | 1
( Metal Handle, without ratchet, Outer Tube with Insert )
7 | Click Line Grasping Forceps with teeth, 2.0 mm, working length 30 cm, | 1
(Metal Handle, without ratchet, Outer Tube with insert )
8 | Puncture Needle | 1
9 | Coagulating Electrode, bipolar, 5 Fr. | 1
10 | Bipolar High Frequency Cord with 2 x 4 mm banana-plug to Coagulator Valleylab, length 300 cm | 1
11 | Holding System, autoclavable, with fastener: KSLock, | 1
Consisting of:
Rotation Socket to clamp on the operating table, for use with European and United States standard rails, with
lateral clamping element for height and angle adjustment of the articulated stand
Articulated Stand, reinforced version, straight, with one mechanical central clamp for all five joint functions,
heigth 30 cm, operating range 37 cm, with fastener: KSLock (female)
Clamping Jaw, metal, with axial intake, for use with instrument and telescope sheaths, clamping range 4.8 up to
12.5 mm, with fastener: KSLock (male)
12 | Scissors, pointed, lightly curved jaws, double action jaws, diameter 1,7 mm, length 30 cm | 1
13 | Balloon Catheter, O.D. 0.7 mm, length 40 cm, sterile, single use, 10 pieces | 1
High Definition Camera System
1 | HD module, for use with up to 3 link modules, resolution 1920 x 1080 pixels, with integrated and digital Image | 1
Processing Module, power supply 100 - 120 VAC, 50/60 Hz; 200 - 240 VAC, 50/60 Hz
Consisting of:
Mains Cord, length 300 cm
DVI Signal Cable, length 300 cm
Connecting Cable, length 100 cm
USB Flash Drive,
2 | link module, for use with FULL HD threechip camera heads, power supply 100 - 120 VAC/200 - 240 VAC, 50/60 Hz, for | 1
use with HD
Consisting of:
Mains Cord, length 300 cm
Link Cable, length 20 cm
3 | Three-Chip FULL HD Camera Head, progressive scan, soakable, gas- and plasma-sterilizable, with integrated Parfocal | 1
Zoom Lens, focal length f = 15 - 31 mm (2x), 2 freely programmable camera head buttons
4 | Power LED 175 Light Source, with integrated high-performance LED and one light outlet, | 1
power supply 110 - 240 VAC, 50/60 Hz
Including:
Mains Cord
Connecting Cable
5 | Fiber Optic Light Cable, with straight connector, extremely heat-resistant, enhanced light transmission, diameter 4.8 mm, | 1
length 250 cm
6 | 26" FULL HD Monitor, color systems PAL/NTSC, max. screen resolution 1920 x 1080, image format 16:9, | 1
Interface: RS 232, power supply 85 - 264 VAC, 50/60 Hz, wall mount with VESA 100 adaptor
Including:
External 24 VDC Power Supply Mains Cord
7 | Trolley Powder Coated Rides on 4 Antistics, Dual Wheels, 2 Euipped with Locking (Front) 2 Fixed Shelves, 1 Drawer Unit | 1
with Lock (Locally Supply)
Serial No.05
| HEAD HOLDER | 2 Set
SPATULA HOLDER, LEYLA,2 ARMS, COMPLETE
Consisting of
| FIXATION ARM, FLEXIBLE | 2
| FIXATION BASE F. 2 ARMS | 1
| SPATULA HOLDER, F. FLAT SPATULAS | 2
Serial # 4
| C-Arm Image Intensifier | 1
Image intensifier on “C” arm of wide depth and with orbital/skull-caudal rotation
Generator: High Frequency Monoblock, power rating 5kW, 110kV, 35mA, 125mAs. Continuous and pulsed
fluoroscopy 8mA. Snap shot 10mA.
X-ray tube: Stationary anode, 0.5/1.5mm focus
9” Image Intensifier: Triple field 9” – 6” – 4”,
useful input diameter 215/160/120mm
TV Camera: CCD, high resolution, 625/50 TV lines. Image rotation 360°, automatic return to 0°.
L.I.H. + 330 frames non volatile memory + 2° 19” LCD monitors – L/R image inversion – edge enhancement –
Real time Digital Rotation – Grey scale inversion, with keyboard for editing data patient
Collimator: Motorised with iris diaphragm
Serial No.05
| HEAD HOLDER | 2
SPATULA HOLDER, LEYLA,2 ARMS, COMPLETE
Consisting of
| FIXATION ARM, FLEXIBLE | 2
| FIXATION BASE F. 2 ARMS | 1
| SPATULA HOLDER, F. FLAT SPATULAS | 2
Serial No. 6
High Speed Neurosurgical Drill System.
1 | NEURO motor system with color display, touch screen, two motor outputs, integrated irrigation pump | 1
and integrated SCB module, 100 - 240 VAC, 50/60 Hz
Consisting of:
Motor Unit
Mains Cord
Irrigator Rod
Two-Pedal Footswitch, two-stage, with proportional function
Silicone Tubing Set, for irrigation, sterilizable
Clip Set, for use with Tubing Set
Connecting Cable, length 100 cm
Single Use Tubing Set, sterile, package of 3
2 | High-Speed Micro-Motor, max. speed 60,000 rpm, including connecting cable, | 1
3 | High-Speed Handpiece, long, angled, 60,000 rpm, for use with High-Speed Micro-Motor | 1
4 | Craniotome Handpiece, maximum speed 60,000 rpm, including medium | 1
dura protector for use with High-Speed Micro-Motor as well as 3.17 mm Craniotome Burrs and Dura Protector
5 | High-Speed Craniotome Burr, medium, shaft diameter 3.17 mm, for single use, sterile, package of 5, for use with | 1
60,000 rpm Craniotome. Handpiece and Medium Dura Protector
Serial No. 7
| VP Shunt Passer With Complete Set | 1
Serial No. 8
| Spinal Stabilization System | 1
Serial No. 9
Matrix System for Spinal Surgery Set
Endosccopic Spinal Surgery Set
1 | Puncture Needle, including stylet, diameter 1.8 mm, working length 18 cm, with 1.3 mm opening for guide wire | 1
2 | Guide Wire, unsterile, diameter 1.2 mm, length 31 mm, package of 10 | 1
3 | Dilation Sleeve, OD5,2mm, ID 1.5 mm, graduated, length 23cm, colour code white | 1
4 | Dilation Sleeve, OD8.9 mm, ID 5.3 mm, graduated, length 21 cm, colour code yellow | 1
5 | Dilation Sleeve, graduated, inner diameter 9 mm, outer diameter 12.7 mm, length 19 cm, color code: orange | 1
6 | Dilation Sleeve, OD14.9mm, ID12.9mm, graduated, length 17cm, colour code red | 1
7 | Dilation Sleeve, OD16.9 mm, ID15.1 mm, graduated, length 15cm, colour code green | 1
8 | Dilation Sleeve, OD18.9 mm, ID17.1 mm, graduated, length 14 cm, colour code blue | 1
9 | Dilation Sleeve, OD20.9 mm, ID19 mm, graduated, length 13 cm, colour code black | 1
10 | Trocar, diameter 15 mm, working length 39 mm,for use with Trocar Attachment Telescope Holder and Forward-Oblique | 1
Telescope 25°, 4 mm, 6 cm color code: orange
11 | Trocar, diameter 19 mm, working length 40 mm, for use with Trocar Attachment | 1
12 | Trocar, diameter 23 mm, working length 43 mm, for use with Trocar Attachment Telescope Holder and Forward-Oblique | 1
Telescope 25°, 4 mm, 6 cm color code: black
13 | Trocar Attachment, diameter 15 mm, for use with Trocars | 1
14 | Telescope Holder, diameter 15 mm, for use with Trocar Attachment | 1
15 | Trocar Attachment, diameter 23 mm, for use with Trocars | 1
16 | Telescope Holder, diameter 23 mm, for use with Trocar Attachment | 1
17 | Trocar Attachment, diameter 19 mm, for use with Trocars | 1
18 | Telescope Holder, diameter 19 mm, for use with Trocar Attachment | 1
19 | Forward-Oblique Telescope 25°, eyepiece angled 90°, diameter 4 mm, length 6 cm, autoclavable, fiber optic light transmission | 1
incorporated, for use with 4 cm long Trocars color code: red to use with 4 cm long trocars
20 | Forward-Oblique Telescope 25°, eyepiece angled 90°, diameter 4 mm, length 11 cm, autoclavable, fiber optic light transmission | 1
incorporated, for use with 10 cm long Trocars color code: red to use with 10 cm long trocars
21 | Fiber Optic Light Cable, with straight connector, diameter 3.5 mm, length 230 cm | 1
22 | Bone Punch, dismantling, 90° upbiting, not through-cutting, 2 mm, working length 24 cm | 1
23 | Bone Punch, dismantling, 90° upbiting, not through-cutting, 4 mm, working length 24 cm | 1
24 | Bone Punch, dismantling, 40° upbiting, not through-cutting, 2 mm, working length 24 cm | 1
25 | Bone Punch, dismantling, 40° upbiting, not through-cutting, 4 mm, working length 24 cm | 1
26 | Punch, dismantling, bayonet-shaped, fixed, downbiting 40° forward, 2 mm, working length 17 cm | 1
27 | Punch, dismantling, bayonet-shaped, fixed, upbiting 40° forward, 2 mm, working length 17 cm | 1
28 | Hook Scissors, single action jaws, diameter 2.5 mm, working length 25 cm | 1
29 | Spoon Forceps, dismantling, robust, oval, spoon size 3 x 10 mm, single action jaws, working length 20 cm | 1
30 | Trocar, diameter 19 mm, working length 71 mm, for use with Trocar Attachment Telescope Holder and Forward-Oblique | 1
Telescope 25°, 4 mm, 9 cm color code: green
31 | Dissector n. dead hand, bayonet shaped, 3 mm, curved upwards, with round handle, sharp, working length 16 cm | 1
32 | Dissector, distal end tapered , bayonet-shaped, working length 15 cm | 1
33 | Nerve Hook, distal width 3 mm, bayonet-shaped, working length 16 cm | 1
34 | Nerve Hook, distal width 5 mm, bayonet-shaped, working length 16 cm | 1
35 | scalpel, with telescopic blade, | 1
36 | Suction Tube, with distal nerve retractor, with cut-off hole, LUER-Lock connector, diameter 2.7 mm, working length 15 cm | 1
37 | Nerve Retractor, angled 30°, distal width 5 mm, working length 17 cm | 1
38 | Dissector, sharp, tip angled 15°, with round handle, size 2 mm, length 25 cm | 1
39 | Palpation Hook, straight, distally 10 mm long and angled 90°, with ball, working length 20 cm | 1
40 | Nerve Retractor, hook length 2 mm, diameter 4 mm, angled sheath, working length 20 cm | 1
41 | Bipolar Coagulating Forceps, insulated, bayonet-shaped, tip 0.7 mm, length 23 cm, for use with bipolar high frequency cord | 1
42 | Bipolar Coagulating Forceps, insulated, bayonet-shaped, tip 1.2 mm, length 23 cm, for use with bipolar high frequency cord | 1
43 | Nasal Dressing Forceps, bayonet-shaped, length 20 cm | 1
44 | Plastic Container for sterilizing, suitable for steam, gas, and hydrogen peroxide (Sterrad®) sterilization and storage, | 1
perforated, with lid, external dimensions (w x d x h): 321 x 90 x 45 mm for use with two rigid endoscopes up to
max. 20 cm working length
45 | Rotation Socket to clamp on the operating table with one already mounted butterfly nut 28172 HRS, for use | 1
with European and United States standard rails, with lateral clamping element for height and angle adjustment
of the articulated stand
46 | Articulated Stand, reinforced version, only, L-shaped, with one mechanical central clamp for all five joint functions | 1
height 48 cm, operating range 52 cm, with fastener: KSLock (female)
47 | Bipolar High Frequency Cord, Length 300 cm, | 2
48 | Trocar, diameter 23 mm, working length 94 mm, for use with Trocar Attachment Telescope Holder and | 1
Forward-Oblique Telescope 25°, 4 mm, 11 cm.
Serial # 10
| Cervical Stabilization Support System. | 10
| Serial # 11 | 10
Patient Monitor
Parameter: Pulse, EEG, ECG, NIBP, SpO2, Resp., Temp with Battery Backup.
Serial # 12
| ICU Ventilaor | 6
Serial # 13
| Operating Microscope | 2
Serial # 14
Neuro Navigation System
A | ASSEMBLED STAFF (220-240V @ 50-60 Hz) | 1
Including:
Navigation camera with extended camera volume
Staff visualization Monitor
USB Stick transfer of exams to the StealthStation® S7™ system
LINUX operating system/ Computer 4GB Ram / 500GB HD
High Fidelity MP3 Audio Speaker System (MP3 player not included)
DICOM query/retrieve compatible
AxiEM™ compatible
O-ARM® Intra Operative Imaging System compatible
Microscope interface integration with most manufacturers
Technical features:
Cart Footprint:60.3 x 57.1 cm (23.75 x 22.5“)
Monitor viewable size: 39 cm (15.25”)
Monitor resolution: 1440 x 900, 60 Hz
Maximum camera height: 240 cm (94.5“)
Maximum camera reach: 125 cm (49.2“)
B | Touch SURGEON LCD ASSEMBLED
| 61cm high resolution wide screen. Reflection free for best viewing from all angles. Flexible arm to move close to surgical | 1
field to provide all information the surgeon needs where he needs it.
Technical features:
Surgeon Monitor Resolution: 1920 x 1200, 60Hz
Surgeon Monitor Viewable size: 61cm (24“)
Cart Footprint 58.7 x 60.3 cm (22.75” x 23.75 “)
C | SOFTWARE S7 SYNERGY SPINE UNIFIED:
| The Synergy™ EXPERIENCE StealthStation® System is the spine navigation product offering from Medtronic Navigation. | 1
Synergy™ EXPERIENCE software is customized for the surgeon to simplify and standardize the entire navigation process.
Surgeon’s preferences allow each surgeon to work in his preferred and optimized navigation environment and to
maximize his work flow.
Synergy™ Experience StealthStation System was designed to streamline the integration of surgical navigation into an OR
Comprising both hardware and software, the Synergy™ Experience design includes a user interface for preplanning and
surgery, a sterile mouse for interaction, a surgeon monitor for dynamic feedback during surgery, and a wireless tracker to
facilitate image capture.
Synergy™ Experience StealthStation® System is a proprietary image calibration and navigation program allowing
seamless interface to a variety of C-arms. Images from the C-arm are automatically imported and can be calibrated for
navigation use. Real time tool positions are overlaid onto previously acquired images.
Spine Instrumentaion
D | StealthStation® Spine Referencing Set | 1
E | CONTAINER 9734448 REFERENCING SET | 1
F | StealthStation® Navlock Instrument Set | 1
G | StealthStation® CD Horizon Legacy Taps & Drivers | 1
Cranial Software & Instruments
H | New Cranial Instrument Set, Passive | 1
I | SFTWR KIT 9733937 SYNERGY CRANIAL S7 SOF | 1
Passive Biopsy
J | Vertek Passive Biopsy Instrument Set | 1
K | Passive Biopsy Needle Kit | 10
N | Navigation Software
| StealthMerge™ Software + Market leading image fusion application. + Automated merging algorithms for multiple data sets (CT, MRI-T1, MRI-T2, flare MRI, PET etc.) for the StealthStation® S7™ System and Synergy® Cranial Application. + Support merging and visualization of multiple data sets simultaneously. + Imaging matching or panning techniques are available for manual backup. | 1
| Stealth3D™ Software + Advanced Visualization package for the Synergy® Cranial Application. + Optimizes display of specific anatomical regions with preset color maps. + Select from variety of display modes including: volume rendering, shaded volume rendering, silhouette and contours. + Visualization Overlay and blend settings for optimized viewing of multiple image modalities simultaneously. *Requires StealthMerge™ Software | 1
| SFTWR KIT 9733933 DICOM SOFTWARE Q/R: | 1
With DICOM Query / Retrieve, you can search, view and download patient
exams directly over the hospital network to the StealthStation® S7TM.
No more waiting for radiology to push exams to your Stealth system.
No more trips to radiology to load them on a CD.
Query of PACS systems for patient data
Push and pull patient data from PACS or single scanner location
Conform to industry accepted DICOM standards
DICOM receive from any DICOM storage class provider
Patient data viewing capability within the navigation application
| Suretrak | 1
O | SureTrak 2 Universal Instrument Adapter System, Combo
Disposables
P | 5 pack 12/box (60 spheres per box) | 10
Q | Fiducial Markers Box (100) and Pen | 2
Microscope Interface
R | Software (supports the pointer and HUD function) | 1
Tracking Hardware
Serial # 15
| Cortical Mapping Electrods with Monitor | 10

## Annexure A Section (B)
LIST OF MACHINERY / EQUIPMENTS FOR NEURO SURGERY DEPARTMENT, CMC HOSPITAL LARKANA FOR THE YEAR 2015-2016
(1). List of Neuro Surgery Equipments (B)
S No. | Description |  |  | Qty | Unit Rate on CIF Basis Figure & Words | Unit Rate on FOR Basis Figure & Words | Total Cost (Figure & Words
Serial # 1
| Neurosurgical Electro Hydraulic OT Table. |  |  | 2
1 | Five sections Electro-hydraulic operating table,
Adjustment possibilities and ranges meets the requirments
demanded by a modern universal operating tables
The operating table made of stainless, acidproof steel.
Features
Modren Design
Battery powered, with facility for connection to
mains electricity for immediate use.
Unique software providing smooth and precise table movements.
Superior electro- hydraulic system, with technical solutions of the world
leading micro-hydraulics providers.
Controls
Optional override control system, mechanic-hydraulic, used in case of
corded hand control damage, failure of drive or main control system.
Range of table movements controlled via the hand control
Height adjustment
Table top tilts: lateral and longitudinal
(Trendelenburg and Rev- Trendelenburg)
Back rest section and kidney bridge (option)
Longitudinal slide (option)
Adjustment angles of 40° Trendelenburg/Trendelenburg and and
30° tilt make it ideal solution for minimally invasive procedures.
Range of table movements controlled mechanically or gas spring supported:
- Backrest section and kidney bridge,
- Footrest sections,
- Headrest sections,
- Longitudinal slide
TECHNICAL PARAMETERS
| Total length of the table with 5 sections tabletop |  | 2100 mm
| Tabletop width |  | 500 mm
| Total tabletop width (with side rails) |  | 560 mm
| Standard height adjustment, with mattress |  | 730 mm –1080 mm
(electro-hydraulic)
| Extended height adjustment with mattress |  | 700 mm –1150 mm*
(electro-hydraulic)
| Lateral tilts adjustment (hydraulic) |  | ± 30°
| Trendelenburg / reverse-Trendelenburg |  | ± 40°
adjustment (electro-hydraulic)
| Backrest adjustment (pneumatic) |  | - 45° ; +85°
| Optional backrest adjustment (electro-hydraulic) |  | - 35° ; +75° *
| Footrest adjustment (pneumatic) |  | - 90° ; +25°
| Footrest abduction (mechanic) |  | 180°
| Headrest adjustment (pneumatic) |  | ± 45°
| Kidney bridge adjustment (mechanic), |  | ~120 mm
optional (hydraulic)
| Longitudinal slide(mechanic),Optional(hydraulic |  | 350 mm
| Table mass |  | 250 kg
| Power |  | 24 V
| Batteries (sealed, maintenance-free) |  | 12V, 7Ah, 2 pcs
| Battery charging time |  | Maximally to 5 h
| Working time between battery charging s |  | Up to 65 operation
| Type of operations |  | Interrupted operation 2/18
| Battery charger |  | Built-in
| Charger power |  | 230V ~ 50/60Hz
| Power consumption from mains |  | 1 A / 230 V
| Maximum working load |  | 250 kg
| Lifetime |  | 10 years
| Degree of protection |  | IP-X4
| Class of protection against electric paralysis |  | I
| Application part type |  | B
Following Accessories are Included
| Specialistic head rest with adjustable track |  |  | 1
| arm support /with one-position clamp (One Pair) |  |  | 2
| Hand strap (One Pair) |  |  | 2
| Anesthetic frame with adjustable width |  |  | 1
| Knee support /with multiposition clamp (One Pair) |  |  | 2
| Abdominal Strap |  |  | 1
| Infusion Holder |  |  | 1
| bowl |  |  | 1
| Wireless hand controller |  |  | 1
Serial # 2
| Antierior Cervical Surgery Set |  |  | 1
Serial # 3
1 | Electrosurgery Unit 400 watts |  |  | 1
Functional Test
Progressive power controll
Endo Mode" Time Controll cuting
Bipolar coagulation with foot switch
Connector for Monopolar hand switches
Multifunctional connector for Bipolar instruments
Monopolar cutting & coagulation by foot Switch
Technical Specification of the Unit
| Supply voltage: | 220-240 V; 50-60 Hz or 100-127 V; 50-60 Hz
Selectable via dummy plug located inside the unit by Technical Service
| Power input: | with no HF power output: approx.16 VA,
at max. power output: approx. 800 VA
| Class of protection : | I
| Classified acc. to MDD : | II b
| Leakage currents LF and HF : | in acc. with EN IEC 60601, Part 2-2
| Type of equipment : | CF; defibrillator-proof
| Nominal frequency : | 450 kHz
| Pulse frequency : | 30 kHz
HF output power:
| Type of current | Power | Crest factor Voltage
| Cutting 1 : | max. 400 W at 300 Ω | 1.6 at 300 Ω max. 2300 Vpp
| Cutting 2 | max. 300 W at 300 Ω | 1.9 at 300 Ω max. 2500 Vpp
| Endo-Mode | max. 100 W at 200 Ω | 1.6 at 200 Ω max. 2800 Vpp
| Contact coagulation | max. 250 W at 200 Ω | 3.4 at 200 Ω max. 3200 Vpp
| Spray coagulation | max. 120 W at 300 Ω | 5.6 at 300 Ω max. 6000 Vpp
| Bipolar coagulation | max. 100 W at 100 Ω | 2.1 at 100 Ω max. 600 Vpp
| Duty type | intermittent INT 10 s/ 30 s, equivalent to a duty factor of 25%
| Mains fuses | 220-240 V: T 4 A (slow-blow.)
100-127 V: T 8 A (slow-blow.)
| Signal level : | HF activation: 55 dB (A) (adjustable between 50 dB and
60 dB by Technical Service)
Alarm: 65 dB(A)
| Weight : | 8.6 kg
Interference suppression : Limits in conformity with EN 55011,
Interference immunity in conformity with IEC 801
| Dimensions : | 405 mm x 135 mm x 380 mm (W x H x D)
CE marking conform with 93/42/EEC
Accessories: (for each set)
2 | Double-pedal foot switch |  |  | 1
3 | Rubber neutral electrode for adults single 15X26 cm |  |  | 1
4 | Non stick Red forceps, Straight, Blunt. 0.6 mm, 20 cm |  |  | 1
5 | Bipolar Cable |  |  | 1
6 | Reusable Electrode Handle with two push buttons |  |  | 1
7 | Knife Electrode, 30mm |  |  | 1
8 | Trolley (Local - Best Quality) |  |  | 1
Serial # 4
| MayField Head Securing Clip with Applicator |  |  | 1
Serial # 5
| Central Oxygen Supply System |  |  | 6
Serial # 6
| Suction Machine |  |  | 10
Serial # 7
| Oxygen Cylinder Large Size |  |  | 10
Serial # 8
| Nitrogen Gas Cylinder |  |  | 5
Serial # 9
| EEG MACHINE with Grids (Epilepsy Surgery) |  |  | 2
EEG 32 Channel Portable w/Ethernet Interface Portable Base System
Base Laptop with 10/100 Ethernet Network Support Windows XP
Professional Operating System DVD Recorder Archive Device with
blank media.
Portable System Network
Interface Card PCMCIA 1000Base‐T
EEG Module with 32 channel amplifier EEG acquisition and review application
software Basic Report Generation Software 32 channel Amplifier System
supply amplifier interface cable, and amp holder 32 channel Gold EEG
Electrode Starter Kit, Includes:
Includes amplifier with built‐in electrode connections,amplifier power
36 Gold 10mm cup electrodes with 1.5m cables, Nu‐Prep (4oz),
Lemon‐Prep Cups 24Pk
Ten20 Paste (4oz), 1 inch Surgical Tape (1 Roll) and 1 Tape Measure
Operator's Manuals Remote Support Services
Patient Administrator Software (Access)
Lamp Photic Stimulator Lamp Photic and power supply, arm and Amp/Photic
Wheeled Stand, plus cables for fixed and mobile system use.
Includes non‐ isolated interface cable Isolated transformer.
Cart for the portable laptop EEG System Includes: Power outlet with
IEC320 connectors and a foot switch bracket.
Serial # 10
| Multiparameter Monitor |  |  | 10
5 basic parameter ECG, NIBP, SpO2, Resp., Temp., and Li-ion Battery.
TECHNICAL SPECIFICATIONS
| Display Type: |  | 12" or more olor LED backlight LCD
| Waveforms: |  | up to 8
| Temperature range: |  | 0 to 50 °C (32 to 122 F)
| Battery Type: |  | Chargeable Lithium-ion
| Run time: |  | 2 hrs
| ECG input |  | 5 lead ECG cables
| Sweep speed |  | 6.25mm/s, 12.5mm/s, 25mm/s, 50mm/s
| HR display range |  | Adu: 30 to 300 bpm
| Analysis |  | ARR and ST analysis
| Respiratory range |  | 0 - 120rpm
| Sweep speed |  | 6.25mm/s, 12.5mm/s or 25mm
NIBP
| Operation mode |  | Manual/Auto
| Parameter |  | Systolic, Distolic
SPO2
| Range |  | 0 - 100%
| Pulse rate |  | 20 - 250 bpm
| Alarm limit setting |  | Upper limit/lower limit
1 | 12.1" ECG, NIBP, Temp, SpO2.
2 | Thermal Recorder
4 | Adult, 5-lead ECG ELECTRODES
5 | Adult Finger SPO2 probe
6 | Adult cuff with tubing
7 | Reusable Temp probe, Adult
8 | Reusable Temp probe, Adult,
9 | Li-ion Battery
10 | Power cord
Serial # 11
| Digital Cerebral Angiography |  |  | 1
Serial # 12
| Double Dome Ceiling LED OT Light |  |  | 2
Mechanical Design - Ceiling-mounted operating light consisting of a ceiling tube, horizontal
tracking arm and spring arm with cardanic joint and a light head. The spring arm shall allow
full adjustability in height and unrestricted movability.
| TECHNICAL DATA |  | LED
Central illuminance (at a distance of 1 m ) 160,000 lx
Turbulence degree (standard value: 37.5%) 27%
| Power input |  | 45 W
| Color Temperature |  | 4300 K
Effective Life of light source
Soft-start option (light does not start immediately with full brightness) Yes
Electronic system meets requirements acc. To Yes
| VDE and IEC |  | Yes
Central illuminance (electric diming from/to) 30-100% /5% in endo mode
Light field diameter (d10, distance of 1 m ) 170-220 mm
Color rendring index (CRI) [Ra (1-8)] 95
| Color temperature |  | 3,600 / 4,300 / 4,600 K
Illumination depth, without refocusing) L1/L2 180 / 200 mm
Permissible ambient temperature range during operation 5° to +40° C
Power module incl. mounting plate, 200 x 130 x 105 mm
Classification acc. to MPG (Medical Devices Act / MDD) I
Protection class acc. to IEC 601 I
Degree of protection: light head IP 42
| Conformity |  | CE
Serial # 13
a | LED Light Camera |  |  | 2
120x zoom
HD-SDI digital video signal
Full HD video output For teleconsultations
No image interference caused by electro-surgery
Maximum image quality with 2 million pixels in full HD resolution
100% interference-free transmission to external storage media
Compatible with LED Ceiling OT Light
Compatible with DVD Recording System
TECHNICAL DATA OF HD CAMERA
| Image Sensor | 1/3" CMOS
| Zoom | 10x optical / 12x digita
| Effective pixel number | Approx. 2,000,000
| Number of pixels/Resolution | 1920 x 1080i
| Focal length (zoom lens) | f = 5.1 mm to 51 mm
| White balance | automatic
| Protection class | I
| Type | B
| Mark of conformity | CE
| Wall socket connector standard | 3 x BNC socket for video signal
1 x multiple DIN socket for control signals
| Power requirements: control unit / camera | 100-240 V, 50/60 Hz / 24 V AC/DC
b | Recording System |  | 1
Examinations & Minor procedures can be recorded in a cost-effective manner
Functions
Multi functional image & video recorder
Recording images & videos up yo full HD quality via camera
Enter patient data in combination with the keyboard
Storage of data on an external thumb drive or the internal data hard drive
Provision of data in the network
Video Signal input (DVI, SDI, S-VIDEO and Composite)
Recording Indicator
Recording indicator should disply in the upper-left corner of the monitor
Patient Data
Button on the front panel enables shuld direct access to the patient information screen
Image Options
Should be enable the section of image formates (e.g. JPG or TIFF)
Specification
| Video Signal inputes: | DVI, HD, SDI, Composite, S-Video, RGB, YPbPr
| Video Out put resolution: | 1920x1080, 1280x720, 1024x768, 800x600 640x480
| Internal Hard drive | 320 GB
| USB Ports | USB 2.0
| Network | RJ45/Connections as network drive (SMB)
| Recording Formats | Videos: H264mp4
Images: JPG, TIFF, BMP
| Patient Data | Saved as .txt file and / or in EXIF format
Consisting of:
19 Monitor (Local)
Key Board /Mouse (Local)
Serial 14
| Electical AutoClave (Size 250 to 400 Litres) |  |  | 2
Serial 15
| DEFIBRILLATOR with Trolley |  |  | 5
TECHNICAL SPECIFICATIONS
DISPLAY:
| Screen | Active LCD screen, with visible viewing area
| Waveform display | should display waveform sweep speed for ECG and SPO2
| Manual defibrillation mode: | Should have energyy select range from 2 to 360 with user
configurable sequence
| External defibriilator | AED with shock advisory system
| Pacing | Available with pacing option, with pacing rate of 40 - 180ppm
| Output waveform | Monophasic output waveform
| Calibration: | Unit should be able to measure impedance of a patient and
calibrate defibrillation waveform and amplitude
| waveform | Type: Biphasic Truncated Exponential
| Modes: | In-service mode, Archieve mode
| ECG lead selection | 3 wire / 5 wire ECG cable selection option
| Patinet's Data management | and transmission Should be able to record/store patient's data
| Heart rate range | 20-350bpm
| Heart rate display | Digital
| Battery | Rechargeable internal battery
Standard Accessories:
Hard paddles (detachable, Adu/Ped)
3-lead ECG cable
Adult ECG electrodes
Pacing electrodes
Therapy cables
ECG printer papers ( 50 x 30 mm)
Trolley
Serial 16
| Anesthesia Machine with Ventilator |  |  | 2
Tidal Volume setting20-1500ml,including: 8.4 inch color screen, ventilator, glass flowmeters
(hypoxic guard),work surface, 1 drawer, 2 sets of flowsensor, 1 battery (90 min), 20pcs of
foam pad, pipeline pressure gauges, user manual, quick user guide, packing material
System software, including: VCV mode, Pressure controlled ventilation mode (PCV) SIMV
mode (SIMV-VC and SIMV-PC) Pressure supported ventilation mode (PSV,
w/apnea backup) TV compensation, pressure monitoring, volume monitoring, PEEP
European (ID) style NIST-2 (European) O2 & N2O & Air - inlet with hoses O2 drive
2 positions of vaporizer (One vaporizer included). PSU Breathing circuit
(with manual ventilation arm)
O2 monitor and accessory kit, including:
O2 sensor, Medicel MOX-2, 1
O2 sensor cable Disposable accessory kit (Adult), including:
Patient circuit (breathing bag 3L, Adapter Wye, Adapter elbow,
Airway filter), adult, 1pcs
Large adult mask, Air-Soft, w/valve, 22 mm, 1pcs Passive AGSS system,
including accessory kit: Tube, 32mm, 1pc
Adaptor, 1pc Auxiliary O2 flowmeter PIN index cylinder yokes
(O2 and N2O) Auxiliary sockets LED lights Power cord (Europe)
PHYSICAL CHARACTERSITICS
Dimenions
| Height |  | 1300mm or more
| Width |  | 700mm or more
| Depth |  | 600mm or more
| Castors: |  | 125 mm with brakes
| Drawers |  | 2 or more drawers with mounting shelf
Modes of Ventilation
VCV, PCV, SIMV-VOLUME, SIMV-PRESSURE, PSV (OPTIONAL).
Ventialtor parameters:
| tidal volume |  | 1500ml
| Pressure: |  | 0 - 60cmH2O
| Rate: |  | 4 - 100bpm
| PEEP |  | Integrated electronic controlled
| FiO2 |  | 0 - 100%
Vaporizer:
2 positions, selectatec, Halothane, EnfluraneSevoflurane, Isoflurane
Pipeline inlet pressure
Gas Type: O2, NO2, Air
| pipeline input range: |  | 0 to 6 bar
Manometer Gauges
| N2O, Air, O2 |  | 0 – 10 bar
Rotameter:
| O2 tube |  | 1 – 10 L/min
| N2O tube |  | 1 – 10 L/min
| Air |  | 1 – 10 L/min
| Volume of CO2 absorber: |  | 1.5 liters (1500ml)
| Audible and Visual alarm: |  | Low/high airway pressure, Apnea alarm,
O2 alarm, Tidal volume alarm,
upper/lower limit alarm, user and technical alarm,
major leak or patient disconnection.
| Battery |  | Li-ion (built-in)
Power Requirement:
Power input: 100-240 VAC, 50/60 Hz
Serial # 17
| Stereo-Taxy Computerized Set. |  |  | 1
1 | The stereotactic system should be (Complete Arc) Arc centered with a 190mm radius, and be
based on a Cartesian coordinate system.
Open Stereotactic system that can be used with CT and MRI scanners including 3T systems.
The Cartesian coordinate system shall conform to the X-, Y-, and Z- nomenclature used in
CT- and MR –scanning.
The total accuracy of the system should be minimum 0.7mm (with accuracy certification)
and certification in this regard to be provided along with FDA and CE certificates.
There shall be sterilization trays tailored for the frame and arc system included in the delivery.
The principal components of the stereotactic system shall include a Cartesian coordinate
frame and a full semicircular arc with counter scale.
There should be three options of lengths of the posterior posts (long, medium, short)
The guide and stop inserts should be able to split into two parts to enable effective cleaning.
The stereotactic system should include CT and MR adapters to secure and support the
patient’s head during scanning ensuring accurate imaging.
Should be supplied with all reusable instruments for biopsy, hematoma evacuation and catheter
placement and other functional neurosurgery accessories.
The base frame of stereotactic system should have the capability to be utilized as a head holder
for micro-surgical applications (adapter for Microsurgical Accessories, Fork and Clamp type.)
Frame must have the ability for 3-point fixation to the patient to enable fixation even in cases
where a bone flap is limiting a 4-point fixation to the skull (Slotted front piece.)
The stereotactic system should provide the ability for rigid 3-point fixation of the frame to
the operating table (Frame Fixation)
The CT and MR Indicator box should not be a limitation for how low the frame may be mounted
(Open MR Indicator, Open CT Indicator).
System should be supplied with USFDA, CE certificate and system accuracy certificate
1 | (SURGICAL PLANNING SYSTEM) for use in Stereotactic and
Functional Neurosurgery.
The system shall allow the user to virtually place entry and target points directly within images of
the patient’s brain acquired from CT, MR and Angiography.
The system shall allow to interactively visualize and examine the tissue that the surgical pathway
will affect.
The system shall allow to avoid critical structures such as blood vessels, cranial nerves or other
critical structures by adjusting the virtual surgical pathway prior surgery.
It shall be possible to simulate any number of surgical pathways.
The system shall be able to import and display CT, MR, PET and Angiographic images.
Images should be possible to import directly from CT, MR and PET scanners via hospital
networks or CD.
Imported images shall be automatically scaled and quality controlled.
Target should be possible to outline semi-automatically or manually, and non-continuous regions
should be allowed.
3D visualization of the images shall be possible using different 3D renderings, including open
book and cut box, and with all defined objects displayed.
It must possible to define the AC-PC line whether or not it lies with a single image plane.
The system should enable co-registration of frame-bases and frameless images.
The system should be Linux based.
It should be possible to design a workspace that displays more than 4 image window.
System should be supplied with USFDA, CE certificate and system accuracy certificate
Technical specification for the Microdrive
Microdrive should have working distance channel of 40mm to integrate completely with the
stereotactic frame with a mechanical resolution of 50µm and be designed for optimal, accurate
positioning of electrodes in the brain when performing MER, macro stimulation and lead
electrode implantation. For the optimal accuracy, the system should have the guide and stop holders.
The system should be able to hold up to five electrode(s) simultaneously and enable to record
Multiple single tracks using any one of the five positions, without adjusting the arc.
Should enable the surgeon to operates with micro electrodes, combined micro-macro electrodes,
macro electrodes and lead electrodes.
Should have manual control which allows continuous recording during movement of the electrode(s).
The system should be compatible with all MER systems available in the market.
System should be fully autoclavable and should be supplied with all necessary dedicated
sterilization trays and components.
All necessary universal Guide tubes, microdrive bipolar lesion electrode kit, and other DBS
guide tubes should also be provided.
serial # 18
| Ultrasonic Aspirator |  |  | 1

## Annexure B Neuro
CHANDKA MEDICAL COLLEGE HOSPITAL LARKANA
FINANCIAL BID
Supply, Installation, Testing, commissioning & Maintenance Durign maintenance period of Medical Equipment / Instruments under Revenue Component ADP Scheme No 372 Namely Construction of Neuro Science Complex LIST OF MACHINERY / EQUIPMENTS FOR NEURO SURGERY DEPARTMENT & NEURO MEDICINE, CMC HOSPITAL LARKANA FOR THE YEAR 2015-2016
(1). List of Neuro MEDICINE Equipments
S No. | Description | Qty | Unit Rate on CIF Basis Figure & Words | Unit Rate on FOR Basis Figure & Words | Total Cost (Figure & Words
Serial # 1
| EMG MACHINE | 1
Specifications:
Number of Channels: 2
Viking Quest Amp Arm
5-pin DIN connector on all channels, LED Indicator Light
4 to 8 electrode inputs
2 to 4 recording channels in the differential mode.
Sensitivity: 1 μV/division to 10 mV/division in 13 steps, 2V peak-peak max., full
scale output 16 bit Analog-to-Digital Converter (ADC) with 100KHz smapling rate
Input Impedance >100 MOhm
Common Mode Rejection Ratio: 110 dB, typical; >105 dB at 50 to 60 Hz, typical
Notch Filter 50 Hz, 60 Hz, On or Off, in all tests
Electrical Stimulator Specifications:
Footswitch
S403 Electrical Stimulator Probe with remote controls and adapter head
One Internal Electrical Stimulator
Removable isolated electrical stimulus unit for placement close to patient
Independent Outputs: 1 or 2 channel
Stimulus Intensity: 0 to 100 mA or 0 to 400 V continuously adjustable
intensity within user selectable maximum range into a 4 kW load
Stimulus Duration: 0.01 - 1 ms
Stimulus Modes Single, train, recurrent and non-recurrent operation
Stimulus Rate: 0.1 - 100 per second, depending on test
Isolation Features: Fully isolated outputs
Level Control Console/Remote mode for adjusting stimulus intensity either from console
or remotely from electrical stimulator probe
Stimulus Type Choice of stimulus type: Selectable constant current or constant voltage
Auditory Simulator Features;
Signal Types Click, tone pip or tone burst
Stimulus Rates: 0.1 - 91.1 per second in 20 steps
Stimulus Intensity: 0 to 139 dB pSPL or -31 to 109 dB nHL, depending on stimulus
type and frequency and transducer type
Stimulus Attenuators Keyboard controlled, separate units for right and left signal
channels, each with 140 dB dynamic range selectable in programmable step size
Click Polarity: Condensation, rarefaction and alternating
Click Duration: 100 μsec
Visual Simulator Features;
LED Flash Rate: 0.1 to 100 per second
LED Flash Duration: .01 to 1 msec in .05 steps under software control
Clinical software
NCS/EMG Software Package: MNC, SNC, ANS, F-waves, H-waves, Blink Reflex, Repetitive Stimulation,Inching studies, Anomalous innervation, Heart rate, Variation (R-R interval), Sympathetic skin response, Signal Enhancer, Spontaneous EMG, Maximum Volitional EMG, Quantitative EMG.
EP Software (Includes SEP, AEP, and VEP).
Power Supply 100/120 V, 230 V± 10%, 50 Hz or 60 Hz
Power Consumption Approximately 50-80 W, depending on
Temperature: 60 to 90° F (15.6 to 32.2° C)
Relative Humidity: 20-80%, non-condensing
Serial # 2
| EEG MACHINE | 2
EEG 32 Channel Portable w/Ethernet Interface Portable Base System
Base Laptop with 10/100 Ethernet Network Support Windows XP
Professional Operating System DVD Recorder Archive Device with
blank media.
Portable System Network
Interface Card PCMCIA 1000Base‐T
EEG Module with 32 channel amplifier EEG acquisition and review application
software Basic Report Generation Software 32 channel Amplifier System
supply amplifier interface cable, and amp holder 32 channel Gold EEG
Electrode Starter Kit, Includes:
Includes amplifier with built‐in electrode connections,amplifier power
36 Gold 10mm cup electrodes with 1.5m cables, Nu‐Prep (4oz),
Lemon‐Prep Cups 24Pk
Ten20 Paste (4oz), 1 inch Surgical Tape (1 Roll) and 1 Tape Measure
Operator's Manuals Remote Support Services
Patient Administrator Software (Access)
Lamp Photic Stimulator Lamp Photic and power supply, arm and Amp/Photic
Wheeled Stand, plus cables for fixed and mobile system use.
Includes non‐ isolated interface cable Isolated transformer.
Cart for the portable laptop EEG System Includes: Power outlet with
IEC320 connectors and a foot switch bracket.
Serial # 3
| ICU Ventilator | 3
| Serial # 4 | 8
BP PATIENT MONITOR
5 basic parameter ECG, NIBP, SpO2, Resp., Temp., and Li-ion Battery.
TECHNICAL SPECIFICATIONS
Display Type: 12" or more olor LED backlight LCD
Waveforms: up to 8
Temperature range: 0 to 50 °C (32 to 122 F)
Battery Type: Chargeable Lithium-ion
Run time: 2 hrs
ECG input 5 lead ECG cables
Sweep speed 6.25mm/s, 12.5mm/s, 25mm/s, 50mm/s
HR display range Adu: 30 to 300 bpm
Analysis ARR and ST analysis
Respiratory range 0 - 120rpm
lp 6.25mm/s, 12.5mm/s or 25mm
NIBP
Operation mode Manual/Auto
Parameter Systolic, Distolic
SPO2
Range 0 - 100%
Pulse rate 20 - 250 bpm
Alarm limit setting Upper limit/lower limit
12.1" ECG, NIBP, Temp, SpO2.
Thermal Recorder
Adult, 5-lead ECG ELECTRODES
Adult Finger SPO2 probe
Adult cuff with tubing
Reusable Temp probe, Adult
Reusable Temp probe, Adult,
Li-ion Battery
Power cord
Serial # 5
| TENS Machine (Transcutaneous Electrical Nerve Stimulation) | 1
Serial # 6
| Ultrasonic Therapy Machine | 1
Serial # 7
| Shortwave Diathermy Machine | 1
Rapid Reduction of wounds
Rapid re absorption of haematomas and oedemas
Powerful Stimulation of the Peripheral Circulation
Post-Traumatic and post operative disorders such as arthropathy, confusion, distortion, haematom.
Other Disorders such as ankylopoietic ,Spondylosis , Bursitis, coccygodinia, myalagia(acute ) and chronic sinusitis.
Specification
General frequency 27.12 MHz
Power output 500 W (Adjustable in 5 steps
Timer digital timer 0-99 min. with acoustic signal and automatic cut-off.
Serial # 8
| DL (Nerve Stimulator) Machine | 1
Serial # 9
| Execise Instrument Set | 1
Serial # 10
Library Books
| SIGNATURE CONTRACTOR NAME |  |  | DATE & SEAL

## Annexure A Surgery dept
CHANDKA MEDICAL COLLEGE HOSPITAL LARKANA
Supply, Installation, Testing, commissioning & Maintenance Durign maintenance period of Medical Equipment / Instruments under Revenue Component ADP Scheme No 373 Namely Up gradation Surgery Department Surgery Equipment For O.T, Modernization of Laboratory, Blood Bank & Establishment of Oncology Treatment Ward at CMCH LARKANA, FOR THE YEAR 2015-2016
1.List of Machinery Equipment (A)
S No. | Description |  |  | Qty | Unit Rate on CIF Basis Figure & Words | Unit Rate on FOR Basis Figure & Words | Total Cost (Figure & Words
Serial # 1
Video Gastroscope & Colonoscope with Endovision System
1 | Video Gastroscope, color system PAL, |  |  | 1
sheath outer diameter 9.3 mm,
working channel diameter 2.8 mm,
working length 1100 mm,
deflection up/down 210°/100°,
deflection left/right 120°/120°,
field of vision 140°,
depth of field 2 - 100 mm
Following accessories are included in delivery:
| 1x | Case
| 1x | Pressure compensation cap
| 1x | Suction valve
| 1x | Air/water valve
| 1x | Leakage tester
| 1x | Cap, for working channel, package of 10
| 1x | Irrigation tube, for suction/biopsy channel
| 1x | Y-irrigation tube, for suction/irrigation
| 1x | Cleaning brush, for single use
| 1x | Valve brush, for video endoscopes
| 1x | Protective Film, for transporting flexible endoscopes
Serial # 2
Video Colonoscope
2 | Standard Colonoscope, color system PAL, |  |  | 1
sheath outer diameter 12.9 mm,
working channel diameter 3.8 mm,
working length 1400 mm,
deflection up/down 180°/180°,
deflection left/right 160°/160°,
field of view 160°,
depth of field 2 - 100 mm
Following accessories are included in delivery:
| 1x | Irrigation Tube, reusable
| 1x | Y-Irrigation Tube, reusable
| 1x | Pressure Compensation Cap, ETO cap
| 1x | Leakage Tester
| 1x | Sealing Cap, for working channel, package of 10
| 1x | Suction Valve
| 1x | Air/Water Valve
| 1x | Valve Brush, for video endoscopes
| 1x | Protective Film
| 1x | Case
3 | XENON 100 with integrated including an integrated insufflation pump and 100 Watt Xenon bulb, |  |  | 1
power supply 100 -240 VAC, 50/60 Hz, for use with video endoscopes
consisting of:
XENON 100
Mains Cord
Water Bottle
Irrigation Bottle Holder
Irrigation Adaptor, for water bottle System requirements for use with SCB-PC: SCB-R-UI
Software Release
4 | connect module, for use with up to 3 link modules, resolution 1920 x 1080 pixels, with integrated |  |  | 1
Processing Module, power supply 100 - 120 VAC, 50/60 Hz; 200 - 240 VAC, 50/60 Hz and digital Image
Consisting of:
Mains Cord, length 300 cm
DVI Signal Cable, length 300 cm
SCB Connecting Cable, length 100 cm
USB Flash Drive,
5 | link module, for use with flexible video endoscopes, 50/60 Hz, power supply 100 - 120 |  |  | 1
VAC/200 - 240 VAC, 50/60 Hz
Mains Cord, length 300 cm
Link Cable, length 20 cm
6 | Video Endoscope Adaptor 0°, color systems PAL/NTSC, length 60 cm, for use with Link Module |  |  | 1
7 | 26" FULL HD Monitor, color systems PAL/NTSC, max. screen resolution 1920 x 1080, image format |  |  | 1
16:9, video inputs: DVI, 3G-SDI, VGA, S-Video, Composite, video outputs: DVI, 3G-SDI, Composite,
power supply 100 - 240 VAC, 50/60 Hz, 5 V DC output (1 A), wall-mounted with VESA 100 adaption,
including:
External 24 VDC Power Supply Mains Cord
8 | Equipment Cart narrow, small, rides on 4 antistatic dual wheels equipped with locking brakes |  |  | 1
mains switch on cover, energy beam with integrated electrical subdi stributors with 6 sockets,
grounding plugs, Dimensions in mm (w x h x d): Equipment cart: 660 x 1265 x 730,
Shelf: 450 x 25 x 510, Caster diameter: 150 mm
Consisting of:
Base Module, equipment cart, narrow
Cover, equipment cart, narrow
Beam Package, equipment cart, small
Shelf, narrow
Drawer Unit with Lock, narrow
Equipment Rail, long
9 | Monitor Holder, height adjustable, swiveling and tilting, central mount, swivel range approx. 360°, |  |  | 1
loading capacity max. 18 kg, with monitor mount VESA 75/100, for use with Equipment Carts
Serial # 3
Laparoscope with HD Camera System
1 | Straight Forward Telescope 0° enlarged view, diameter 10 mm, length 31 cm, autoclavable, |  |  | 1
fiber optic light transmission incorporated.
2A | Forward-Oblique Telescope 30° enlarged view, diameter 10 mm, length 31 cm, autoclavable, |  |  | 1
fiber optic light transmission incorporated.
2B | Straight Forward Telescope 0°, enlarged view, diameter 5 mm, length 29 cm, autoclavable, |  |  | 1
fiber optic light transmission incorporated.
2C | Forward-Oblique Telescope 30°, diameter 5 mm, length 29 cm, autoclavable, fiber |  |  | 1
optic transmission light incorporated.
3 | Trocar, size 11 mm, color code: green, |  |  | 2
Trocar only, with pyramidal tip
Cannula without valve, with insufflation stop- cock, length 10.5 cm
Multifunctional Valve
4 | Trocar, size 6 mm, |  |  | 2
Trocar only, with pyramidal tip
Cannula without valve, with insufflation stop- cock,length 10.5 cm
Multifunctional Valve
5 | Reduction Sleeve, 11/5 mm |  |  | 1
6 | VERESS Pneumoperitoneum Needle with spring loaded blunt stylet, LUER-lock, length 13 cm |  |  | 1
7 | KELLY Dissecting and Grasping Forceps, rotating, with connector pin for unipolar coagulation, |  |  | 1
size 5 mm, length 36 cm, double action jaws,
Plastic Handle, without ratchet, Outer Tube, insulated, Forceps Insert.
8 | REDDICK-OLSEN Dissecting and Grasping Forceps, rotating, with connector pin for unipolar |  |  | 1
coagulation, size 5 mm, length 36 cm, heavy, double action jaws,
Plastic Handle, without ratchet, Outer Tube, insulated, Forceps Insert.
9 | Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, |  |  | 1
length 36 cm, atraumatic, fenestrated, single action jaws,
Plastic Handle, with hemostat style ratchet, Outer Tube, insulated, Forceps Insert.
10 | Claw Grasping Forceps rotating, with connector pin for unipolar coagulation, size 5 mm, |  |  | 1
length 36 cm, 2 x3 teeth, single action jaws,
Plastic Handle, with hemostat style ratchet, Outer Tube, insulated, Forceps Insert.
11 | Grasping Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, |  |  | 1
length 36 cm, atraumatic, fenestrated, single action jaws,
Plastic Handle, with hemostat style ratchet, with larger contact area.
Outer Tube, insulated, Forceps Insert
12 | METZENBAUM Scissors, rotating, with connector pin for unipolar coagulation, size 5 mm, |  |  | 1
lenght 36 cm, blades curved, double action jaws, length of blades 12 mm,
Insulated Handle, without ratchet, Outer Tube, insulated, Scissor Insert
13 | Suction and Irrigation Tube, anti-reflex surface with two-way stopcock, for single hand control, |  |  | 1
size 5 mm, length 36cm
14 | Coagulating and Dissecting Electrode, L-shaped, with connector pin for unipolar coagulation, |  |  | 1
size 5 mm, working length 36 cm
15 | Monopolar High Frequency Cord with 4 mm plug for HF-unit, models and Erbe type T, |  |  | 1
older models; length 300cm
16 | Clip Applicator, for use with Ethicon Titanium-Clips LT-300 (medium large), dismantling, rotating, |  |  | 1
with ratchet to lock the jaw part holding the clip, size 10 mm, length 36 cm,
Metal Handle, with ratchet, Metal Outer Tube, Appiler Insert
17 | KOH Macro Needle Holder, ergonomic handle with ratchet, left curved jaws, with tungsten carbide |  |  | 1
inserts, ø 5mm, length 33 cm for use with suture material 0/0-7/0
HD Camera System
18 | Digital FULL HD videoprocessor module with integrated documentation |  |  | 1
function for image and video capturing for use with three-chip-FULL HD
camera heads, one-chip-FULL-HD camera heads and video endoscopes in
combination with compatible camera modules.
Technical data:
Image sensor: 3 x 1/3“ CCD chip
Pixels: 1920 (h) x 1080 (v) pixels per chip.
CCD chip supports 16:9 input format
Progressive scan (full frame) for Integrated parfocal zoom lens, f = 15-31 mm (2x optical zoom).
display of 50 or 60 frames per second.Lens:
Minimum light sensitivity: 1.17 Lux (f = 1.4 mm).
Control buttons: 3 (2 of them freely programmable).
Dimensions (wxhxd): 38x100x48 mm.
Weight: 246 g.
Integrated connection cable with plug to CCU and color coded cable nut to indicate
compatibility to Sterrad® 100S, NX 100NX Steris® System1, V-Pro1 sterilisation methods.
Special features:
Modular design: For use in combination with at least one camera module.
Parallel live display of visualization modes besides white light mode (picture-in-picture).
Up to three different camera modules can be connected to the FULL HD video processor module.
Integrated picture-in-picture mode of two different camera modules in five different display
sizes available.
Primary and secondary signal source change in picture-in-picture mode can be performed
easily via camera head button.
Integrated communication bus for device control and information display of connected devices.
Highest possible FULL HD resolution of 1920 x 1080 pixels.
Integrated documentation unit via USB.
Still image capturing in FULL HD quality (JPEG format).
Still image capturing can be adjusted by 4 different, timewise variable image freeze selections
including picture-in-picture mode.
Video capturing in FULL HD quality (MPEG4 format).
Still image and video capturing in FULL HD quality in picture-in-picture mode for both signal
sources simultaneously.
Medical USB printer compatible (plug& play).
Data capturing functions can be released by optional USB footswitch.
Integrated digital imaging processing module for a 5 level brightness regulation and
2 electronic anti-moirée filter for fiberscopes.
Automatic adjustment of light intensity of light source via communication bus.
White balance button on front plate of video processor module.
Control of complete camera system can be realized from camera head from sterile area.
Short learning curve due to familiar handling, short starting time and customizable parameter
adjustment.
Grid and pointer can be displayed for improved orientation and communication during surgery.
Grid and pointer can be displayed individually and together.
2 x digital zoom, adjustable in 5 levels.
Possibility of 180° image rotation.
Possibility of vertical and horizontal image mirroring.
Storage of up to 20 individual presets.
Storage of up to 20 individual patient data.
System overview is individually configurable andsetup status can be directly displayed with
intelligent icons.
Parameter setup can be adjusted during surgery.
Number of menu icons can be customized individually for optimal system adjustment for the user.
Modular design for integration of up to 3 camera modules and easy upgradeability with
new technologies.
Future ready thanks to integrated and extendable architecture.
Including:
FULL HD video processing module
Power cable, length 300 cm
Silicone keyboard with touchpad
DVI-D cable, length 300 cm
USB flash drive, 32GB
Communication bus cable
19 | Digital FULL HD camera module for use with Three-Chip-FULL-HD camera |  |  | 1
heads.
Technical Data:
Camera input: 1 x three-chip-FULL HD camera head.
Maximal resolution: 1920 x 1080 pixels.
Automatic shutter. yes
Automatic Gain Control: Microprocessor controlled.
| Connection: |  | Module link cable to video processor module.
Dimension (w x h x d): 30.5 cm x 5.4 cm x 32.0 cm.
| Weight: |  | 1.86 kg.
Power supply: 100 - 120 VAC, 50/60 Hz, 200 - 240 VAC, 50/60 Hz.
Power consumption: 62 VA.
| Certified to: |  | IEC601-1, 601-2-18, CSA 22.2 No. 601, UL2601-1 and CE
according to MDD, protection class 1/CF defib.
Special features:
Modular design: For use with a video processor module and a compatible three-chip FULL HD
camera head. In combination with a compatible three-chipFULLHD camera head the following
modes can be activated without special
light sources or filters:
- Color inversion by spectral color shift.
- Brightening of dark areas in the endoscopic image.
- Contrastenhancement.
Compatible with special FULL HD microscopy camera head.
Highest possible FULL HD resolution of 1920 x 1080 pixels.
Progressive scan for an extremely smooth image without flickering and noise.
Up to 4 functions can be assigned to 2 freely programmable camera head buttons of
compatible camera heads. Changes in visualization modes, device control, digital zoom,
brightness, video capture, still image capture and direct print orders, picture-in-picture mode,
image direction, white balance and setup settings can be performed in sterile area via camera
head buttons.
Integrated digital image processing module for brightness adjustment in 5 levels and 2
electronic anti-moirée filters for fiberscopes.
Future ready thanks to integrated and extendable architecture.
Backward compatible with selected existing three-chipFULLHD camera heads.
Including:
FULL HD camera module
Power cable, length 300 cm
Link cable, length 20 cm
20 | Digital three-chip camera head for full HD with 2x parfocal optical zoom and |  |  | 1
2 freely programmable camera head buttons, compatible with systems with
integrated, innovative visualization technology for surgery by shifting the
color spectrum and via homogeneous illumination and contrast
enhancement.
Technical data:
Image sensor: 3 x 1/3“ CCD chip
| Pixels: |  | 1920 (h) x 1080 (v) pixels per chip.
CCD chip supports 16:9 input format
Progressive scan (full frame) for Integrated parfocal zoom lens, f = 15-31 mm (2x optical zoom).
display of 50 or 60 frames per second.Lens:
Minimum light sensitivity: 1.17 Lux (f = 1.4 mm).
Control buttons: 3 (2 of them freely programmable).
Dimensions (wxhxd): 38x100x48 mm.
| Weight: |  | 246 g.
Integrated connection cable with plug to CCU and color coded cable nut to indicate
compatibility to Sterrad® 100S, NX 100NX Steris® System1, V-Pro1 sterilisation methods.
Special features:
Three-chip technology for brilliant images.
Image acquisition in format 16:9, with 1920 x 1080 pixels and progressive scan 50 Hz or 60Hz.
Digital through and through: Conversion of optical images to digital signals at the earliest
possible site: the CCD sensing chip.
Digital input guarantees best possible image quality and eliminates need to convert analog
input signals to digital output formats, which could result in a loss of image quality.
Captured digital images in format 16:9 can be displayed on WideView monitors in the same
full HD format without being converted. This prevents a loss if image quality caused by image
ovalization (stretching).
Image acquisition with an input resolution of more than 2 million pixels – 5 to 6 times higher
than with SD (depending on the SD color system used, i.e. PAL or NTSC).Ergonomic design.
2 freely programmable camera head buttons.
2x parfocal optical zoom.
Compatible with systems with integrated, innovative visualization technology for surgery by
shifting the color spectrum and via homogeneous illumination and contrast enhancement.
Suitable light sources are controllable by camera head buttons via communication bus.
Suitable insufflators are controllable by camera head buttons via communication bus.
Suitable camera control units can store HD images and HD video clips to USB mass storage
devices in conjunction with the camera head.
Suitable camera control units can print HD images in conjunction with the camera head and a
compatible USB printer.
The camera head is rated CF-Defibrillation according to EN 60601 allowing cardiac surgery.
The camera head is compatible with the following sterilization methods:
– Sterrad® 100S, NX 100NX
– Steris® System1, V-Pro1
– ETO-Gas Sterilisation
A suitable sterilization tray is available to safely store the camera head during sterilization.
Including:
Digital three-chip Full HD camera head with 2x parfocal optical zoom and 2 freely
programmable camera head buttons, with integrated connecting cable with plug to CCU,
compatible with systems with integrated, innovative visualization technology for surgery by
shifting the color spectrum and via homogeneous illumination and contrast enhancement.
21 | Cold Light Fountain Power LED 175 with integrated , high-performance LED and one light outlet, |  |  | 1
power supply 110 - 240 VAC, 50/60 Hz
Specification:
| Lamp Life: |  | 30,000 Hours
| Temperature: |  | 6400K
| Light Out Let |  | 1
| Lamp Tyoe |  | High-Performance LED
| Certified to |  | TEC -601-1 Protection Class 1/CF
Features:
Low Noise
Can be controlable with camera head button
| operation mode: | Auto/ Manual
| Display | Digital
Including:
Cold Light Fountain Power LED 175
Mains Cord
Connecting Cable
22 | 27" FULL HD Monitor, with VESA 100 adaption, With LED Back Light For Low Energy, |  |  | 1
Picture-in-Picture (PiP) And Picture-by-Picture With Medical Grade FDA, UL & CE Approval
Special Features:
| color systems | PAL/NTSC,
max. screen resolution 1920 x 1080,
| image format |  | 16:09
| Interface: |  | RS 232,
power supply 85 - 264 VAC, 50/60 Hz,
| Video inputs: |  | DVI, VGA, S-Video, Composite /FBAS
Video outputs: DVI, S-Video, Composite/FBAS
including: External 24 VDC Power Supply Mains Cord
23 | Plastic Container for sterilizing, applicable for hydrogen peroxid- (Sterrad-) sterilizing and |  |  | 1
for storage, perforated, with lid, for use with HD-Cameraheads. External dimensions
(wxdxh): 385 x 255 x 75 mm
24 | ENDOFLATOR® 40 Set, with integrated module, |  |  | 1
power supply 100 - 240 VAC, 50/60 Hz
Special Features:
High Degree of patient safety
Easy to use
Touch keys for precise preselection of set values
Optical & acoustic alarm signals in the event of patient overpressure
Fully automatic, electronically controlled gas refill (e.g in case of gas loss when changing
instruments)
Specification:
Operative Mode:
| Gas Flow |  | Operating Mode
| Pressure: |  | 0-40l/min
| Gas |  | 0-40l mmHg (3990Pa)
Measuring/ Control System. Electronic
Parameter Display. Insufflation Presuure, Intraabdominal Pressure: 0-50 mmHg (6650Pa),
Gas Flow:0-40l/min, Gas Consumption :0-999 L
Power Supply. 100-240 VAC, 50/60 Hz
Dimensionswxhxd. 305x164x233 mm
| Weight |  | 6 KG
| Certified to : |  | IEC 601-1,CE acc.to MDD
Including:
ENDOFLATOR® 40
SCB Connecting Cable, length 100 cm
Universal Wrench
Insufflation Tubing Set, with gas filter, sterile, for single use, package of 5
25 | High Pressure Hose. (Local Supply) |  |  | 1
26 | Equipment Cart LC, rides on 4 antistatic dual wheels equipped with locking brakes, |  |  | 1
central beam with integrated electrical subdistributors with 6 sockets, grounding plugs,
Dimensions in mm (w x h x d): Equipment cart: 830 x 1474 x 730, Shelf: 630 x 25 x 510,
Caster diameter: 125 mm
Including:
Base Module, equipment cart LC
Cover, equipment cart LC
Beam Package, equipment cart LC
Shelf, wide
Drawer Unit with Lock, wide
Cameraholder
27 | Monitor Holding Arm, long, height and side adjustable, tilting, swivel range up to 320°, |  |  | 1
reach 760 mm,loading capacity max. 15 kg, with monitor holder VESA 75/100, for use with
Equipment Carts UGxxx
28 | Duomat Set |  |  | 1
| Power Supply | 100-120, 230-240V~, 50/60 Hz
Including:
Duomat
Power Cord
VACUsafe Pack Suction
29 | Silicone Tubing Set for suction, sterilizable |  |  | 4
30 | Suction Bottle, 5l |  |  | 1
31 | Bottle Stand, for suction bottle |  |  | 1
32 | Irrigation Liquid Bottle, 1l, sterilizable |  |  | 1
33 | Silicone Tubing Set, sterilizable for use with pressure cuff |  |  | 4
34 | Bottle Cap, for suction bottle 1.5 and 5l, sterilizable |  |  | 1
35 | Bottle Stand Holder |  |  | 1
36 | Pressure Infusion Cuff, 1.5 l |  |  | 1
Serial # 04
1 | Laparoscopic Simulator |  |  | 1
Serial # 05
1 | Ventilator |  |  | 3
| SIGNATURE |  | CONTRACTOR NAME |  | DATE & SEAL

## Surgery A (B)
CHANDKA MEDICAL COLLEGE HOSPITAL LARKANA
LIST OF MACHINERY EQUIPMENTS FOR OPERATING OF SURGERY DEPARMENT, CMC HOSPITAL LARKANA
1.List of Machinery Equipment (B)
S No | Description | Qty | Unit Rate on CIF Basis Figure & Words | Unit Rate on FOR Basis Figure & Words | Total Cost (Figure & Words
Serial # 01
1 | Pelvic Retractor, size 10 mm,length 36cm | 1
Serial # 02
| EndoTrainer System | 1
1 | Training Model, for laparoscopic surgery
Including:
Base Plate, Top Plate, EAGS model, Foam Rubber Diaphragm, 8 pcs, Case, Drape
2 | TELE PACK X LED, endoscopic video unit for use with all TELECAM one-chip camera heads and
video endoscopes, incl. LED-light source similar to Xenon technology, with integrated digital
Image Processing Module,
Special Features:
15" LED Backlight Display
Image Rotation
24 bit color depth for natural color rendition
DVI-D video output for connecting external monitors
Compaitable with all camera
Compaitable with medical grade USB color printer 549M
USB port for external hard drives USB flash drives
SD card slot for high storage capacity
High-performace LED light source similar to power LED 175
Color temparature of 6400 K similar to daylight guarantee color fidelity
Average lamp llife up to 30,000 hours
Specifications:
Angle of view:
Memory module,
Power input
Pwer supply
Dimensionswxhxd
Weight
Interfaces
Light source
Image format
Video codec
Video format
Storage interface
LED backlight monitor
including:
USB Silicone Keyboard with Touchpad, with US character set
USB Stick 8GB
Mains Cord
3 | TELECAM 1-Chip Camera Head with 2 freely programable Camera Head buttons, Color System
PAL, with integrated Parfocal-Zoom focal length f = 25 - 50mm (2x)
4 | Fiber Optic Light Cable, with straight connector, diameter 3.5mm, length 230 cm
Serial # 3
1 | Electrosurgery Unit 400 watts | 4
Functional Test
Progressive power controll
Endo Mode" Time Controll cuting
Bipolar coagulation with foot switch
Connector for Monopolar hand switches
Multifunctional connector for Bipolar instruments
Monopolar cutting & coagulation by foot Switch
Technical Specification of the Unit
Supply voltage: 220-240 V; 50-60 Hz or 100-127 V; 50-60 Hz
Selectable via dummy plug located inside the unit by Technical Service
Power input: with no HF power output: approx.16 VA,
at max. power output: approx. 800 VA
Class of protection : I
Classified acc. to MDD : II b
Leakage currents LF and HF : In acc. with EN IEC 60601, Part 2-2
Type of equipment : CF; defibrillator-proof
Nominal frequency : 450 kHz
Pulse frequency : 30 kHz
HF output power:
Type of current Power Crest factor Voltage
Cutting 1 : max. 400 W at 300 Ω 1.6 at 300 Ω max. 2300 Vpp
Cutting 2 max. 300 W at 300 Ω 1.9 at 300 Ω max. 2500 Vpp
Endo-Mode max. 100 W at 200 Ω 1.6 at 200 Ω max. 2800 Vpp
Contact coagulation max. 250 W at 200 Ω 3.4 at 200 Ω max. 3200 Vpp
Spray coagulation max. 120 W at 300 Ω 5.6 at 300 Ω max. 6000 Vpp
Bipolar coagulation max. 100 W at 100 Ω 2.1 at 100 Ω max. 600 Vpp
Duty type intermittent INT 10 s/ 30 s, equivalent to a duty factor of 25%
220-240 V: T 4 A (slow-blow.)
100-127 V: T 8 A (slow-blow.)
Signal level : HF activation: 55 dB (A) (adjustable between 50 dB and
60 dB by Technical Service)
Alarm: 65 dB(A)
Weight : 8.6 kg
Interference suppression : Limits in conformity with EN 55011,
Interference immunity in conformity with IEC 801
Dimensions : 405 mm x 135 mm x 380 mm (W x H x D)
CE marking conform with 93/42/EEC
Accessories: (for each set)
2 | Double-pedal foot switch | 1
3 | Rubber neutral electrode for adults single 15X26 cm | 1
4 | Non stick Red forceps, Straight, Blunt. 0.6 mm, 20 cm | 1
5 | Bipolar Cable | 1
6 | Reusable Electrode Handle with two push buttons | 1
7 | Knife Electrode, 30mm | 1
8 | Trolley (Local - Best Quality) | 1
Serial # 4
| Advance Image Data Archiving System | 1
AIDA Documentation System, for recording still images and videos, dual channel up to FULL
HD, 2D/3D, including SmartScreen (touch screen), power supply 100 - 240 VAC, 50/60 Hz
Including:
AIDA Appliance Manual Mains Cord License Microsoft WES7 (WS7P)
License DICOM
USB Silicone Keyboard, with touchpad
ACC Connecting Cable
DVI Connecting Cable
HDMI to DVI-CABLE 1x
Serial # 5
| Video Choledochoscope | 1
Flexible Video Choledochoscope
deflection of distal tip up/down 270°/270°,
direction of view
angle of view
working channel inner diameter 3.6 Fr.,
distal tip outer diameter 8.5 Fr.,
working length
Following accessories are included:
Case
Pressure Compensation Cap
Leakage Tester
Cleaning Brush, diameter 2 mm, 150 cm
Applicator and Guide Tube
Biopsy Forceps, flexible, 3 Fr.
Grasping Forceps, flexible, 3 Fr.
Stone Basket, 2.5 Fr., 120 cm
LUER-Adaptor, with seal
Seal, for working channel, package of 10
Serial # 6
| Vessel Sealing System | 2
General Characteristics
Output configuration Isolated output
Cooling Natural and forced convection
Display System status indicator - circle LED (12 blocks) shows system status
System error indicator - triangle LED with exclamation mark
Instrument status indicator - LED above the LigaSure™ receptacle
Usage limit indicator - 2 with a line through it
Mounting A Covidien cart (UC8009) or a stable, flat surface
Dimensions and Weight
Width 300 mm (11.81 inches)
Depth 377 mm (14.84 inches)
Height 105 mm (4.13 inches)
Weight 5 Kg (11 lbs)
LigaSure™ Output Characteristics
Mode: P-P Voltage, Rated Load (Ohms), Maximum Power (Watts), Duty Cycle
LigaSure™ LigaSure™
Input Power
100-120 Volt: 220-240 Volt
Maximum VA – 400 VA: Maximum VA – 400 VA
Input mains voltage, full regulation range 90-130 VAC: Input mains voltage, full regulation range 180-264 VAC
Input mains voltage, operating range 100-120 VAC: Input mains voltage, operating range 210-240 VAC
Mains line frequency range (nominal) 50 Hz to 60 Hz: Mains line frequency range (nominal) 50 Hz to 60 Hz
Fuses (2) – 5 mm x 20 mm 8 A, 250 V fast blow: Fuses (2) – 5 mm x 20 mm 8 A, 250 V fast blow
Serial # 7
| Disposible Trocar Size 5mm | 60
Serial # 8
| Disposible Trocar Size 10/11mm | 30
Serial # 9
| Disposible Trocar Size 12mm | 30
Serial # 10
| Disposible Trocar Size 15mm | 30
Serial # 11
| 5mm Sicssor use with Vessel Sealing System | 15
Serial # 12
| 60(EC 60), endo gia unviersal | 15
Serial # 13
| 60 Reload, endo gia | 60
Serial # 14
| Endoscopic Cutter | 30
Serial # 15
| Endoscopic Cutter Reload | 60
Serial # 16
| Counter Gun | 15
Serial # 17
| Counter Gun Reload | 60
Serial # 18
| Liga Max | 30
Serial # 19
| Multi Fire Stapler | 15
Serial # 20
| Cutter 75mm | 15
Serial # 21
| Cutter 100mm | 15
Serial # 22
| Cutter 75mm Reload | 120
Serial # 23
| Cutter 100mm Reload | 120
Serial # 24
| Haemorrhoidal Gun | 60
Serial # 25
| Haemorrhoidal Gun | 30
Serial # 26
| Open Share | 30
Serial # 27
Laparoscpic Share
Serial # 28
| Book Waller aAbdominal Retractor | 1
Serial # 29
| LED Headlight "with Surgical Loops | 4
Special Features
With USB charging cable
Continuously digital dimming
Continuously adjustable focus
Intigrated battrey status indicators
Rapid USB chargiing from wall adapter or PC
Forehead strap two-dimensionally adjustable with one hand
Technical Data:-
Luminous intensity: max. 200,000 lx
Color temperature : 5,000 k
Battrey Service Life up to 6.5 hrs
Weights less than 310 g
Charging time: 6 hours
Consisting of :
| Headlight | 1
| Magni- Loop Glasses 2x | 1
| Carrying case, | 1
| USB charging cable | 1
| Carrying belt | 1
Serial # 30
| Double Dome Ceiling LED OT Light | 4
Mechanical Design - Ceiling-mounted operating light consisting of a ceiling tube, horizontal
tracking arm and spring arm with cardanic joint and a light head. The spring arm shall allow
full adjustability in height and unrestricted movability.
TECHNICAL DATA LED
Central illuminance (at a distance of 1 m ) 160,000 lx
Turbulence degree (standard value: 37.5%) 27%
Power input 45W
Color Temperature 4300K
Effective Life of light source
Soft-start option (light does not start immediately with full brightness) Yes
Electronic system meets requirements acc. To Yes
VDE and IEC Yes
Central illuminance (electric diming from/to) 30-100% /5% in endo mode
Light field diameter (d10, distance of 1 m ) 170-220 mm
Color rendring index (CRI) [Ra (1-8)] 95
Color temperature 3,600 / 4,300 / 4,600 K
Illumination depth, without refocusing) L1/L2 180 / 200 mm
Permissible ambient temperature range during operation 5° to +40° C
Power module incl. mounting plate, 200 x 130 x 105 mm
Classification acc. to MPG (Medical Devices Act / MDD) I
Protection class acc. to IEC 601 I
Degree of protection: light head IP 42
Conformity CE
HD-SDI digital video signal
Full HD video output For teleconsultations
No image interference caused by electro-surgery
Maximum image quality with 2 million pixels in full HD resolution
100% interference-free transmission to external storage media
Compatible with LED Ceiling OT Light
Compatible with DVD Recording System
TECHNICAL DATA OF HD CAMERA
Image Sensor 1/3" CMOS
Zoom 10x optical / 12x digita
Effective pixel number Approx. 2,000,000
Number of pixels/Resolution 1920 x 1080i
Focal length (zoom lens) f = 5.1 mm to 51 mm
White balance Automatic
Protection class I
Type B
Mark of conformity CE
Wall socket connector standard 3 x BNC socket for video signal
1 x multiple DIN socket for control signals
Power requirements: control unit / camera 100-240 V, 50/60 Hz / 24 V AC/DC
Serial # 31
| Electro Hydrulic OT Table | 4
Five sections Electro-hydraulic operating table,
- Adjustment possibilities and ranges meets the requirments
demanded by a modern universal operating tables
- The operating table made of stainless, acidproof steel.
Features
Modren Design
Battery powered, with facility for connection to
mains electricity for immediate use.
Unique software providing smooth and precise table movements.
Superior electro- hydraulic system, with technical solutions of the world
leading micro-hydraulics providers.
Controls
Optional override control system, mechanic-hydraulic, used in case of
corded hand control damage, failure of drive or main control system.
Range of table movements controlled via the hand control
- Height adjustment
- Table top tilts: lateral and longitudinal
(Trendelenburg and Rev- Trendelenburg)
- Back rest section and kidney bridge (option)
- Longitudinal slide (option)
Adjustment angles of 40° Trendelenburg/Trendelenburg and and
30° tilt make it ideal solution for minimally invasive procedures.
Range of table movements controlled mechanically or gas spring supported:
- Backrest section and kidney bridge,
- Footrest sections,
- Headrest sections,
- Longitudinal slide
TECHNICAL PARAMETERS
Total length of the table with 5 sections tabletop
Tabletop width
Total tabletop width (with side rails)
Standard height adjustment, with mattress
(electro-hydraulic)
Extended height adjustment with mattress
(electro-hydraulic)
Lateral tilts adjustment (hydraulic)
Trendelenburg / reverse-Trendelenburg
adjustment (electro-hydraulic)
Backrest adjustment (pneumatic)
Optional backrest adjustment (electro-hydraulic)
Footrest adjustment (pneumatic)
Footrest abduction (mechanic)
Headrest adjustment (pneumatic)
Kidney bridge adjustment (mechanic),
optional (hydraulic)
Longitudinal slide(mechanic),Optional(hydraulic
Table mass
Power
Batteries (sealed, maintenance-free)
Battery charging time
Working time between battery charging s
Type of operations
Battery charger
Charger power
Power consumption from mains
Maximum working load
Lifetime
Degree of protection
Class of protection against electric paralysis
Application part type
Following Accessories are Included
| Specialistic head rest with adjustable track | 1
| arm support /with one-position clamp (One Pair) | 2
| Hand strap (One Pair) | 2
| Anesthetic frame with adjustable width | 1
| Knee support /with multiposition clamp (One Pair) | 2
| Abdominal Strap | 1
| Infusion Holder | 1
| bowl | 1
| Wireless hand controller | 1
Serial # 32
| Emergency Light LED Mobile | 4
Light system
High-performance LEDs. Computer-calculated optical systems boasting a special
microstructure enable the color shadow-free illumination of small or large surgical sites.
Operation
All operating functions can be controlled via an ergonomic,flush-integrated
membrane keypad.Brightness control and light field adjustment are fully electronic.
setting is indicated via an LED segment display on the control panel. High-precision
positioning of the light head in any surgical situation is guaranteed by integrated handle
elements that are accessible from all sides
TECHNICAL DATA
Central illuminance at a distance of 1 m
Effective service life of light sources
Color rendering index [Ra (1 – 8)]
Red rendering index [R9]
Light head dimensions
Color temperature
Light field diameter (d10, distance 1 m)
Light field diameter (d50, distance 1 m)
Central illuminance, electrically dimmable from / to
Illumination depth (20 % on Ec max. without refocusing)
Illumination depth (60 % on Ec max. without refocusing)
Permissible ambient temperature range during operation
Soft-start option (light does not start immediately with full brightness)
Electronic system meets requirements acc. to VDE and IEC
Battery recharging time
Operating time with battery
Classification acc. to MPG (Medical Devices Act / MDD)
Protection class acc. to IEC 601
Conformity
| Sterilizable Handel | 2
| Local UPS (Best Quality) | 1
Standard Accessories are Included :-
Serial # 33
| General Surgery Instrument Set | 4
1 | SCALPEL HANDLE, NO. 3, 12 CM | 1
2 | SCALPEL HANDLE, NO. 4, 13.5 CM | 2
3 | TOWEL FORCEPS, BACKHAUS, SHARP, 13 CM | 9
4 | SPONGE FORCEPS, RAMPLEY, STR., 24.5 CM | 2
5 | FORCEPS, DANDY, CVD. SIDEW., 14.5 CM | 25
6 | HAEM. FORCEPS, PROV.-HOSP., STR., 14 CM | 25
7 | HAEM. FORCEPS, MOSQUITO, CVD., 14 CM | 12
8 | HAEMOST. FORCEPS, CRILE, CVD., 14 CM | 12
9 | TISSUE FORCEPS, ALLIS, 5X6 T., 15.5 CM | 24
10 | LENGENBECK-MANNERFELD RETR | 2
11 | RETRACTOR, RICHARDSON, 48X22 MM, 24 CM | 2
12 | OPERATING SCISSORS, SH/BL, STR., 16.5 CM | 1
13 | RETRACTOR, VOLKMANN, SH., 4-PR., 22.5 CM | 2
14 | NEEDLEHOLDER, MAYO-HEGAR, 14.5 CM | 1
15 | NEEDLEHOLDER, MAYO-HEGAR, 18.5 CM | 1
16 | TC-NEEDLEHOLDER, MAYO-HEGAR, 16 CM | 1
17 | TC-NEEDLEHOLDER, MAYO-HEGAR, 18.5 CM | 1
18 | TISSUE FORCEPS, 1X2 T., SLIM, 18 CM | 2
19 | DRESSING FORCEPS, POTTS-SMITH, 18 CM | 2
20 | SPREADER, ADSON, SHARP, 3X4 T., 24.5 CM | 2
21 | SPREADER, PLESTER, BLUNT, 2X2 T., 13 CM | 2
22 | SPREADER, MOLLISON, SHARP, 4X4 T., 15 CM | 2
23 | BOWL, METAL, H = 75, Ø 167 MM, 0.9 L | 2
24 | BOWL, METAL, H = 40, Ø 80 MM, 0.14 L | 3
25 | KIDNEY DISH, 250X140X40 MM | 4
26 | CONTAINER MS, 60X30X14 CM, HANDLE RED | 1
27 | marTray, 1/1, 532x251x64 mm | 1
28 | COLOR-TAG, RED | 2
29 | Coding label, MS2, w. text, w/o hole | 2
Serial # 34
1 | Patient Transfer Trolley | 4
Serial # 35
| Suction Machine | 4
Serial # 36
| Autoclave Machine | 1
Serial # 37
| Endoscopic Washer | 1
Serial # 38
| Syringe Pump | 10
Syringe pump offers a range of features suited todrug therapy, blood and blood products
Including: large, cleardisplay, intuitive user interface, integrated design and wide rangeof
specialty infusion sets.
TECHNICAL SPECIFICATION:
Infusion Rate Range (With Titration if Enabled)
5ml syringe – 0.1 to 150ml/h
10ml syringe – 0.1 to 300ml/h
20ml syringe – 0.1 to 600ml/h
30ml syringe – 0.1 to 900ml/h
50ml syringe – 0.1 to 1200ml/h
VTBI Range: 0.1 to 1000ml, 1min to 24 hrs.
VI Range: 0.0ml to 9990ml
Volumetric Accuracy: +/- 2% (nominal)
Purge Rate: 100 to 500ml/h (limited by syringe size)
Purge volume range is 0.5ml - 5ml
Purge volume: Range 0.5 to 5ml
Bolus Rate: 5ml syringe - 10 - 150ml/h
10ml syringe - 10 - 300ml/h
20ml syringe - 10 - 600ml/h
30ml syringe - 10 - 900ml/h
50ml syringe - 10 - 1200 ml/h
Bolus Volume: 0.1-25ml (hands-on and hands-free)
Drug Names/Protocols: Configurable for up to 100 drugs per profile
Alarms, Prompts &
Warnings: Alarm beacon with alarm (red) & warning (amber) indicators
Occlusion Settings: 11 Levels from L0-L10.
BackOff: Reduces post occlusion bolus
Event Log: min 1500 events
Key Press Log: 200 Key presses
Battery Type: Rechargeable NiMH
Battery Life: Operates for 6 hrs.from full charge
Recharge Time: 2.5hrs from discharge to 90% charge
Serial # 39
| Blood Warmer | 2
• Free setting of temperature from 37° C to 41° C in 0.5° C increments
• Easy to read display
• Optimized warming circuit
• Economic disposables
• Simultaneous warming of multiple transfusions / infusions
• Splash proof (IPX4)
• High- and low temperature alarm
• Permanent running self-tests
• External service interface for PC
• Mounting on a normed railing and an infusion stand
• Integrated handle
Technical Requirement:-
Device: Blood and infusion warmer
Type designation: Voltage: 230 V, 50–60 Hz Power consumption: 280 W
Current: 1.5 A
Type of protection
against electric shock: I
Degree of protection
against electric shock: B
Degree of protection
Classification: IIb according to rule 9
Type of operation: continuous operation
Type designation: TubeFlow
Dimensions: length 1340 mm
Accessories:-
Extension Set No. 35000, 350 cm, Extension Set No. 46000, 460 cm
Extension Set with bubble trap and
injection port No. 25000, 460 cm
For use with Tube Flow:
Extension Set with injection port,Extension Set with bubble trap and
injection port
Serial # 40
Central Gas System
Design, installation, fucntion Performance, documentation, testing and Commissioning of
Central Medical Gas Pipe Line System for Oxygen,Nitrous Oxide & Air
Medical Gas System Should meet requirement for Supply System, Pipline,
Distribution, Control, Monitoring and Alarm System ensuring Non-Interchangeable between
the Different Gas Sysem
Automatic Mainfold for Oxygen QTY:1
the Mainfold Shall be Fully Automatic Type & Shall Switch From Blank in Use to Reserve bank
with out Flucntion in Delivery Supply Line Pressure & Without the Need for External Power
after the Switch-over the Reserve Bank Becomes the Bank in Use and the bankin use becomes
the reserve ban. The System is abke to maintain the continous supply of electric Power Failure.
Capcity : 2x5
Flow Rate: 70-90m3/hr
Function:
Fully Automatic Self-contained Shuttle-valve with no electric Power Rrequired for Switching.
system should have Electronic Control System and should have following display
Left Source Empty
Right Source Empty
third Soource(Liquied Oxygen)
Malfucntion of Device
Operating Pressure Low
Operating Pressure High
Pressure Sensor for operating Operting Pressure.
Automatic Mainfold for Nitrous Oxide QTY:1
the Mainfold Shall be Fully Automatic Type & Shall Switch From Blank in Use to Reserve bank
with out Flucntion in Delivery Supply Line Pressure & Without the Need for External Power
after the Switch-over the Reserve Bank Becomes the Bank in Use and the bankin use becomes
the reserve ban. The System is abke to maintain the continous supply of electric Power Failure.
Capcity : 2x2
Flow Rate: 70-90m3/hr
Function:
Fully Automatic Self-contained Shuttle-valve with no electric Power Rrequired for Switching.
system should have Electronic Control System and should have following display
Left Source Empty
Right Source Empty
third Soource(Liquied Oxygen)
Malfucntion of Device
Operating Pressure Low
Operating Pressure High
Pressure Sensor for operating Operting Pressure.
Gas outletsBS/DIN System
Gas Outlets complete inbox/Casing Gas Specific, sealed vlave, indexed to eliminate
interchangable/ erroneous tapping of gas,filter and cover plate as per standard design
requirement of the bidder with following color coding with provision of parking stage / position
Oxygen
Nitrous Oxide
Compressed air 4 bar
MEDICAL COMPRESSED AIR STATION; QTY:1
Medical air plant with breathable air . Elimination of toxic gases such as carbon dioxide
nitrogen dioxide nitrogen monoxide carbon monoxide and other pollutants as per
design standrad followed
COMPRESSOR
Duplex compressor unit
oil injected screw / scroll type only
Capacity of air producing approximately minimum of 500L/min each compressor
Mounted on anti vibration base
Dew point monitoring should be availible in the air plant
RESERVOIR
Compatible medical grade
air receiving tank 250 L
vertical type
Galvanized internally with auto drain
Minimum 10 bars output pressure of tank
AIR DRYER
Desiccant type duplex
dew point range between -40 to 30°C
Capacity suitable according to the compressor output
Including oil water separator
FILTRATION SYSTEM
Clean medical grade air supply in accordance with the requirements of HTM/ISO/European
pharmacopeia standards
consisting of pre-filter / humidity oil free and sterile / bacteria filter
Mounted with shut off valves on an assembly panel
Parallel connections of the filters this will make it possible to exchange filter
without interrupting the air supply
Flow/ filtering rate according to the compressed air output
REDUCER PANEL COMPRESSED AIR
DUPLEX
Parallel switched reducer with gauge safety valve pressure switch for high
and low pressure and shut off valve with assembly panel size 4 bar
complete with shut off valve and pressure guage complete assemble
outgoing copper pipe for complete system
CONTROL
Complete with 1 main warning system for compressed air for visual and
acoustic monitoring of alarms conditions for the compressor room.
Test point in the system for air quality
1x switch cabinet for automatic unit incl. all necessary fitting and
installation material for smooth running of the system without any
interruption
POWER : Power supply 3 phases 380 - 400V/50 Hz
Copper Pipping
as per the Site
Three Gasses Valve Box QTY: 3
| SIGNATURE CONTRACTOR NAME | DATE & SEAL

## Specification (4) (hidden)
SPECIFICATIONS
S No. | Description | Qty
ANNEXURE A
Serial No. 3
Neuro Endoscope set with Vertebroscope.
High Definition Camera System
1 | HD module, for use with up to 3 link modules, resolution 1920 x 1080 pixels, with integrated and digital Image | 1
Processing Module, power supply 100 - 120 VAC, 50/60 Hz; 200 - 240 VAC, 50/60 Hz
Consisting of:
Mains Cord, length 300 cm
DVI Signal Cable, length 300 cm
Connecting Cable, length 100 cm
USB Flash Drive,
2 | link module, for use with FULL HD threechip camera heads, power supply 100 - 120 VAC/200 - 240 VAC, 50/60 Hz, for | 1
use with HD
Consisting of:
Mains Cord, length 300 cm
Link Cable, length 20 cm
3 | Three-Chip FULL HD Camera Head, progressive scan, soakable, gas- and plasma-sterilizable, with integrated Parfocal | 1
Zoom Lens, focal length f = 15 - 31 mm (2x), 2 freely programmable camera head buttons
4 | Power LED 175 Light Source, with integrated high-performance LED and one light outlet, | 1
power supply 110 - 240 VAC, 50/60 Hz
Including:
Mains Cord
Connecting Cable
5 | Fiber Optic Light Cable, with straight connector, extremely heat-resistant, enhanced light transmission, diameter 4.8 mm, | 1
length 250 cm
6 | 26" FULL HD Monitor, color systems PAL/NTSC, max. screen resolution 1920 x 1080, image format 16:9, | 1
Interface: RS 232, power supply 85 - 264 VAC, 50/60 Hz, wall mount with VESA 100 adaptor
Including:
External 24 VDC Power Supply Mains Cord
7 | Trolley Powder Coated Rides on 4 Antistics, Dual Wheels, 2 Euipped with Locking (Front) 2 Fixed Shelves, 1 Drawer Unit | 1
with Lock (Locally Supply)

## Specification (5) (hidden)
SPECIFICATIONS
S No. | Description | Qty
ANNEXURE A
Serial No. 10
Endoscopic Spinal System
1 | Puncture Needle, including stylet, diameter 1.8 mm, working length 18 cm, with 1.3 mm opening for guide wire | 1
2 | Guide Wire, unsterile, diameter 1.2 mm, length 31 mm, package of 10 | 1
3 | Dilation Sleeve, OD5,2mm, ID 1.5 mm, graduated, length 23cm, colour code white | 1
4 | Dilation Sleeve, OD8.9 mm, ID 5.3 mm, graduated, length 21 cm, colour code yellow | 1
5 | Dilation Sleeve, graduated, inner diameter 9 mm, outer diameter 12.7 mm, length 19 cm, color code: orange | 1
6 | Dilation Sleeve, OD14.9mm, ID12.9mm, graduated, length 17cm, colour code red | 1
7 | Dilation Sleeve, OD16.9 mm, ID15.1 mm, graduated, length 15cm, colour code green | 1
8 | Dilation Sleeve, OD18.9 mm, ID17.1 mm, graduated, length 14 cm, colour code blue | 1
9 | Dilation Sleeve, OD20.9 mm, ID19 mm, graduated, length 13 cm, colour code black | 1
10 | Trocar, diameter 15 mm, working length 39 mm,for use with Trocar Attachment Telescope Holder and Forward-Oblique | 1
Telescope 25°, 4 mm, 6 cm color code: orange
11 | Trocar, diameter 19 mm, working length 40 mm, for use with Trocar Attachment | 1
12 | Trocar, diameter 23 mm, working length 43 mm, for use with Trocar Attachment Telescope Holder and Forward-Oblique | 1
Telescope 25°, 4 mm, 6 cm color code: black
13 | Trocar Attachment, diameter 15 mm, for use with Trocars | 1
14 | Telescope Holder, diameter 15 mm, for use with Trocar Attachment | 1
15 | Trocar Attachment, diameter 23 mm, for use with Trocars | 1
16 | Telescope Holder, diameter 23 mm, for use with Trocar Attachment | 1
17 | Trocar Attachment, diameter 19 mm, for use with Trocars | 1
18 | Telescope Holder, diameter 19 mm, for use with Trocar Attachment | 1
19 | Forward-Oblique Telescope 25°, eyepiece angled 90°, diameter 4 mm, length 6 cm, autoclavable, fiber optic light transmission | 1
incorporated, for use with 4 cm long Trocars color code: red to use with 4 cm long trocars
20 | Forward-Oblique Telescope 25°, eyepiece angled 90°, diameter 4 mm, length 11 cm, autoclavable, fiber optic light transmission | 1
incorporated, for use with 10 cm long Trocars color code: red to use with 10 cm long trocars
21 | Fiber Optic Light Cable, with straight connector, diameter 3.5 mm, length 230 cm | 1
22 | Bone Punch, dismantling, 90° upbiting, not through-cutting, 2 mm, working length 24 cm | 1
23 | Bone Punch, dismantling, 90° upbiting, not through-cutting, 4 mm, working length 24 cm | 1
24 | Bone Punch, dismantling, 40° upbiting, not through-cutting, 2 mm, working length 24 cm | 1
25 | Bone Punch, dismantling, 40° upbiting, not through-cutting, 4 mm, working length 24 cm | 1
26 | Punch, dismantling, bayonet-shaped, fixed, downbiting 40° forward, 2 mm, working length 17 cm | 1
27 | Punch, dismantling, bayonet-shaped, fixed, upbiting 40° forward, 2 mm, working length 17 cm | 1
28 | Hook Scissors, single action jaws, diameter 2.5 mm, working length 25 cm | 1
29 | Spoon Forceps, dismantling, robust, oval, spoon size 3 x 10 mm, single action jaws, working length 20 cm | 1
30 | Trocar, diameter 19 mm, working length 71 mm, for use with Trocar Attachment Telescope Holder and Forward-Oblique | 1
Telescope 25°, 4 mm, 9 cm color code: green
31 | Dissector n. dead hand, bayonet shaped, 3 mm, curved upwards, with round handle, sharp, working length 16 cm | 1
32 | Dissector, distal end tapered , bayonet-shaped, working length 15 cm | 1
33 | Nerve Hook, distal width 3 mm, bayonet-shaped, working length 16 cm | 1
34 | Nerve Hook, distal width 5 mm, bayonet-shaped, working length 16 cm | 1
35 | scalpel, with telescopic blade, | 1
36 | Suction Tube, with distal nerve retractor, with cut-off hole, LUER-Lock connector, diameter 2.7 mm, working length 15 cm | 1
37 | Nerve Retractor, angled 30°, distal width 5 mm, working length 17 cm | 1
38 | Dissector, sharp, tip angled 15°, with round handle, size 2 mm, length 25 cm | 1
39 | Palpation Hook, straight, distally 10 mm long and angled 90°, with ball, working length 20 cm | 1
40 | Nerve Retractor, hook length 2 mm, diameter 4 mm, angled sheath, working length 20 cm | 1
41 | Bipolar Coagulating Forceps, insulated, bayonet-shaped, tip 0.7 mm, length 23 cm, for use with bipolar high frequency cord | 1
42 | Bipolar Coagulating Forceps, insulated, bayonet-shaped, tip 1.2 mm, length 23 cm, for use with bipolar high frequency cord | 1
43 | Nasal Dressing Forceps, bayonet-shaped, length 20 cm | 1
44 | Plastic Container for sterilizing, suitable for steam, gas, and hydrogen peroxide (Sterrad®) sterilization and storage, | 1
perforated, with lid, external dimensions (w x d x h): 321 x 90 x 45 mm for use with two rigid endoscopes up to max. 20 cm
working length
45 | Rotation Socket to clamp on the operating table with one already mounted butterfly nut 28172 HRS, for use with European | 1
and United States standard rails, with lateral clamping element for height and angle adjustment of the articulated stand
46 | Articulated Stand, reinforced version, only, L-shaped, with one mechanical central clamp for all five joint functions, height 48 | 1
cm, operating range 52 cm, with fastener: KSLock (female)
47 | Bipolar High Frequency Cord, Length 300 cm, | 2
48 | Trocar, diameter 23 mm, working length 94 mm, for use with Trocar Attachment Telescope Holder and Forward-Oblique | 1
Telescope 25°, 4 mm, 11 cm color code: black
Terms & Conditions
Manufacturer | Total FOR
Country of Origin | Rupees Five Million Seven Hundred Thirty Five Thousand One Hundred & Eighty Only.
Delivery Period
Terms Of Payment :
Prices
Validity
Warranty
Karl Storz GmbH & Co.KG.
Germany
12-14 weeks After receipt of order / Advance.
_____________________________ | 100% Advance.
S.Hussain Ali | In Pak Rs (FOR)
National Sales Manager | : 30 Days
Mob:- 0300-826881 | One year warranty against any manufacturing defect

## oRIGINAL (2) (hidden)
SPECIFICATIONS
S No. | Description | Qty
ANNEXURE A
Serial No. 6
High Speed Neurosurgical Drill System.
1 | NEURO motor system with color display, touch screen, two motor outputs, integrated irrigation pump | 1
and integrated SCB module, 100 - 240 VAC, 50/60 Hz
Consisting of:
Motor Unit
Mains Cord
Irrigator Rod
Two-Pedal Footswitch, two-stage, with proportional function
Silicone Tubing Set, for irrigation, sterilizable
Clip Set, for use with Tubing Set
Connecting Cable, length 100 cm
Single Use Tubing Set, sterile, package of 3
2 | High-Speed Micro-Motor, max. speed 60,000 rpm, including connecting cable, | 1
3 | High-Speed Handpiece, long, angled, 60,000 rpm, for use with High-Speed Micro-Motor | 1
4 | Craniotome Handpiece, maximum speed 60,000 rpm, including medium | 1
dura protector for use with High-Speed Micro-Motor as well as 3.17 mm Craniotome Burrs and Dura Protector
5 | High-Speed Craniotome Burr, medium, shaft diameter 3.17 mm, for single use, sterile, package of 5, for use with 60,000 rpm Craniotome | 1
Handpiece and Medium Dura Protector
Terms & Conditions
Manufacturer | Total FOR
Country of Origin | Rupees Five Million Seven Hundred Thirty Five Thousand One Hundred & Eighty Only.
Delivery Period
Terms Of Payment :
Prices
Validity
Warranty
Karl Storz GmbH & Co.KG.
Germany
12-14 weeks After receipt of order / Advance.
_____________________________ | 100% Advance.
S.Hussain Ali | In Pak Rs (FOR)
National Sales Manager | : 30 Days
Mob:- 0300-826881 | One year warranty against any manufacturing defect

## Cardic bypass surgery
CHANDKA MEDICAL COLLEGE HOSPITAL LARKANA
FINANCIAL BID
Supply, Installation, Testing, commissioning & Maintenance Durign maintenance period of Medical Equipment / Instruments under Revenue Component ADP Scheme No 371 Namely Construction of ESTABLISHMENT OF MADAM FARYAL TALPUR CARDIAC SURGERY COMPLEX HAVING FACILITIES OF :i.CORNARY ANGIOGRAPHY CARDIAC CATHETERIZED LAB ii. CARDIAC BYPASS SURGERY iii.CHEST SURGERY AT CMC HOSPITAL LARKANA FOR THE YEAR 2015-2016
1. List of Machinery Equipments (A)
S No. | Description | Qty | Unit Rate on CIF Basis Figure & Words | Unit Rate on FOR Basis Figure & Words | Total Cost (Figure & Words
Serial # 1
1 | SINGLE PLANE CARDIAC ANGIOGRAPHY / Intervational System | 1
SINGLE PLANE CARDIAC ANGIOGRAPHY / INTERVENTIONAL SYSTEM CARDIAC
CATHETERIZATION SYSTEM DEDICATED FOR INTERVENTIONAL CARDIAC PROCEDURES
THE ANGIOGRAPHY SYSTEM SHOULD COMPRISE OF THE FOLLOWING CONFIGURATIONS:
2 | POSITIONING ARM:
GEOMETRY: C-ARM/G-ARM
2.1
2.2 | FLOOR/CEILING MOUNTED STAND REAL TIME DISPLAY OF ROTATIONAL ANGULATIONS.
2.3 | RAO/LAO: ± 120° OR MORE.
2.4 | CRANIAL / CAUDAL: MINIMUM CRA 45° TO CAUD 45° OR MORE.
2.5 | ROTATION SPEED: 30°/SEC OR MORE IN LAO / RAO
2.6 | AUTO POSITIONING:
PROGRAMMABLE AUTO POSITIONING OF SELECTED ANGULATIONS OF UP TO 50
PRE SET POSITIONS.
3 | FLAT PANEL DETECTOR: FIELD TYPE
3.1 | INPUT SIZE OF 20cm X 20cm (8” X 8”)
3.2 | 1024 X 1024 @ 30 FPS AND
512 X 512 @ 60 FPS FOR PEDIATRIC CASES.
3.3 | FPD OPERATION SHOULD REQUIRE NO SPECIAL HUMIDITY OR AIR CONDITIONING.
4 | PATIENT SUPPORT/TABLE
4.1 | CATHETERIZATION TABLE:
FLOOR MOUNTED WITH UP DOWN, LONGITUDINAL, TRANSVERSE AND ROTATIONAL
MOVEMENTS. LONGITUDINAL STROKE: 1000 – 1500 MM, LATERAL STROKE:
100 – 300 MM, CPR: IN ANY TABLE POSITION TABLE TOP SHOULD BE OF SUCH
CONSTRUCTION IN MATERIAL AND DURABILITY TO ACCEPT PATIENT’S WEIGHT
OF NOT LESS THAN 150 KG PLUS 100 KG FOR RESUSCITATION WITH THE TABLE FULLY
EXTENDED. TABLE DIMENSIONS SHOULD BE ABLE TO ACCOMMODATE PATIENT
OF ALL AGES.
TABLETOP SHOULD HAVE LARGE METAL FREE OVER HANG FOR UN-OBSTRUCTED
IMAGING COVERAGE. COMPLETE ACCESSORIES SHOULD BE PROVIDED INCLUDING
ARM HOLDER, HARD GRIP, ARM SUPPORT AND ARM REST. X-RAY GENERATO
SHOULD BE MICROPROCESSOR BASED HIGH FREQUENCY DUAL INVERTER SYSTEM.
DEDICATED X-RAY GENERATOR OF 100 KW
| SERIAL FILMING EXPOSURES WITH SHORTEST EXPOSURE OF L MS WITH AUTOMATIC |  |  |  | .
KV AND MA CONTROL FOR OPTIMUM IMAGE QUALITY.
THE SYSTEM SHOULD HAVE CAPABILITY OF DIGITAL RADIOGRAPHY AND FLUOROSCOPY.
SHOULD HAVE CAPABILITY OF DOING DIGITAL PULSED FLUOROSCOPY 1, 2, 3, 5/6, 7.5/8,
10/12, 15, AND 30 PULSES/SEC. USER SHOULD BE ABLE TO SELECT VARIOUS PULSE
RATES AS AND WHEN REQUIRED INCLUDING SINGLE SHOT FLUOROGRAPHY.
6 | DIGITAL IMAGING AND ACQUISITION / FLUOROSCOPY:
DIGITAL SYSTEM
ACQUISITION, STORAGE AND DISPLAY 1024 X 1024 X 12BITS OR MORE AT 25/30 FPS AND
512 X 512 X 12BITS OR MORE AT 50/60 FPS FOR PAEDIATRIC WORK
PARALLEL PROCESSING CAPABILITY / MULTITASKING FACILITY FOR ACQUIRING AS
WELL AS RECORDING, FILMING, PATIENTS REGISTRATION, POST PROCESSING &
MEASUREMENT, BURNING ON CD SIMULTANEOUSLY. REALTIME FILTERING AND
MAGNETIC DISK CAPACITY FOR STORAGE OF 100,000 FRAMES WITH 1024 X 1024
MATRIX ON ROAD MAP FUNCTION. THE MAGNETIC DISK OF THE MAIN CONSOLE.
DICOM 3.0 WITH STANDARD EXCHANGE MEDIA.
IMAGES TO BE STORED ON AND RETRIEVED FROM ARCHIVAL DISKS FOR POSSIBLE
MANIPULATION AND QUANTIFICATION USING AVAILABLE SOFTWARE PACKAGES.
MULTIPLE (RAID) HARD DISK SYSTEM FOR GREATER SYSTEM RELIABILITY.
MULTIPLE (RAID) HARD DISK SYSTEM FOR GREATER SYSTEM RELIABILITY.
7 | X-RAY TUBE
DUAL FOCUS WITH AT LEAST 1.8/2.4 MHU OR BETTER ANODE HEAT STORAGE
CAPACITY TO ENABLE CONTINUOUS HEAT DISSIPATION DURING SERIAL EXPOSURES.
LIQUID METAL BEARING TUBE TECHNOLOGY FOR LONG LIFE.
DUAL FOCUS ROTATING ANODE.
DOSE MANAGEMENT WITH FLUORO FILTERS WITH AUTO ADJUSTMENT.
8 | MONITORING SYSTEM
2 LCD MONITORS SHOULD BE CEILING-MOUNTED IN THE CATH ROOM OF SIZE 17/18
INCH OR MORE. TWO LCD MONITORS FOR LIVE IMAGES AND ROAD MAPPING IN THE
OPERATING ROOM OF SIZE 17”/18”.
9 | CONTROLS:
ALL CONTROLS OF DIGITAL IMAGING SHALL BE IN THE EXAMINATION AS WELL AS
CONTROL ROOM.
10 | SOFTWARE / HARDWARE PACKAGES:
COMPLETE ANALYSIS PACKAGE FOR THE FOLLOWING CARDIAC APPLICATIONS:
DYNAMIC PRE AND POST PTCA / VALVOTOMY COMPARISON WITH ONE IMAGE LIVE
AND OTHER REFERENCE.
AUTOMATIC POSITIONING OF THE C-ARM CORRESPONDING TO REFERENCE IMAGE.
AUTOMATIC LOOP REPLAY AFTER ACQUISITION OF FLUOROSCOPY.
11 | ACCESSORIES:
(TO BE SUPPLIED AS PER MANUFACTURER’S RECOMMENDATION)
A) SURGICAL SHADOWLESS LIGHT CEILING SUSPENDED.
B) LEAD GLASS
CEILING SUSPENDED TILTABLE LEAD GLASS FOR RADIATION PROTECTION OF
OPERATOR’S HEAD & NECK REGIONS (& LOWER BODY REGION) COLLISION TOLERANT
WITH STAINLESS STEEL OR CAST ALUMIMIUM CEILING MOUNTING.
C) RADIATION PROTECTION DEVICES
LEAD APRONS. 4
LEAD GOGGLES. 4
THYROID PROTECTORS. 4
LEAD GLASS WINDOW OF SIZE 5’ X 2.5 FEET
D) PHYSIOLOGICAL HAEMODYNAMIC MONITORING SYSTEM (01), (AS RECOMMENDED
| BY MANUFACTURER) FOR 2 CHANNEL IBP, CARDIAC OUTPUT WITH THERMO-DILUTION | 1
| METHOD. SIMULTANEOUS 12 LEAD ECG, NIBP AND SPO2 MEASUREMENT. | 1
| THE SYSTEM SHOULD INCLUDE FOLLOWING: | 1
| 2x INVASIVE BLOOD PRESSURE TRANSDUCER | 1
| CABLE FOR 12 LEAD ECG, NIBP AND CARDIAC OUTPUT. | 1
| HOLDER FOR MOUNTING THE IBP TRANSDUCER ALONGSIDE THE PATIENT TABLE. | 1
| E) FULLY PROGRAMMABLE LATEST MODEL CONTRAST MEDIUM INJECTORS: | 1
WITH VOLUME AND FLOW CONTROL AND FACILITY OF CONTRAST MEDIUM HEATING
WITH 100 EXTRA DISPOSABLE SYRINGES.
F) COMPLETE UPS:
FOR THE WHOLE ANGIO SYSTEM OF 120 KVA OR ABOVE.
12 | DICOM 3.0 ACTIVE
12.1 | PRINT, STORE, QUERY, RETRIEVE, WORK LIST TO BE PROVIDED WITH SYSTEM. | 1
| KNOCK OUT CLAUSES. QUOTING FIRMS/SYSTEMS MUST FULFILL THEESE | 1
CONDITIONS OTHERWISE WILL NT BE CONSIDERED FOR PURCHASE :
12.2 | ANGIOGRAPHY SYSTEM MUST HAVE FDA 510K APPROVAL AND CE/MDD . BOTH
ARE REQUIRED
12.3 | ANGIOGRAPHY SYSTEM MUST BE MANUFACTURED IN USA/EUROPE/JAPAN.
12.4 | QUOTING LOCAL FIRM MUST HAVE SOLD AT LEAST 3 OTHER ANGIOGRAPHY SYSTEMS
IN PAKISTAN AND MUST ATTACH STAISFACTORY LETTER FROM THESE USERS
12.5 | SYSTEM MUST CARRY A THREE YEAR WARRNTY FROM THE MANUFACTURER
DIRECTLY ALONG WITH UNDERTAKING OF THE MANUFACTURER THAT THEY WILL
BE RESPONSIBLE FOR THE WARRNTY OF THE SYSTEM
13 | WARRANTY : ONE YEAR COMPREHENSIVE INCLUDING SPARE PARTS & SERVICES
Serial # 2
| Pacernmaker | 1
Serial # 3
1 | Ceilling LED OT Light Triple Dome | 1
Compact design with the significantly smaller & High performance LEDs.
Excellent color retention
Color temperature similar to day light
Extremely long life of the LEDs
Highest hygiene standards.
Maintenance and service costs are completely eliminated
Consisting of:
a | Main Dome : 160,000 lx | 1
b | Satalite Dome I : 160,000 lx | 1
c | Satalite Dome II : 160,000 lx | 1
TECHNICAL DATA
Central illuminance (at a distance of 1 m ) 160,000 lx
Turbulence degree (standard value: 37.5%) 0.27
Power input 45 W
Effective Life of light source 40,000 h
Central illuminance (electric diming from/to) 30-100% /5% in endo mode
Light field diameter (d10, distance of 1 m ) 170-220 mm
Color rendring index (CRI) [Ra (1-8)] 95
Permissible ambient temperature range 5° to + 40° C
during operation
Soft-start option (light does not start YES immediately with full brightness)
Electronic system meets requirements YES
acc. to VDE and IEC
Classification acc. to MPG YES
(Medical Devices Act / MDD)
Protection class acc. to IEC 601 I
Degree of protection: suspension system IP 30
Degree of protection: Light Head IP 42
Conformity CE
Standard Accessories are Included
a | LED LIGHT CAMERA | 1
120x zoom
HD-SDI digital video signal
Full HD video output For teleconsultations
No image interference caused by electro-surgery
Maximum image quality with 2 million pixels in full HD resolution
100% interference-free transmission to external storage media
Compatible with LED Ceiling OT Light
Compatible with DVD Recording System
TECHNICAL DATA OF HD CAMERA
Image Sensor 1/3" CMOS
Zoom 10x optical / 12x digita
Effective pixel number Approx. 2,000,000
Number of pixels/Resolution 1920 x 1080i
Focal length (zoom lens) f = 5.1 mm to 51 mm
White balance AUTOMATIC
Protection class I
Type B
Mark of conformity CE
Wall socket connector standard 3 x BNC socket for video signal
3x multiple DIN socket for control signals
Power requirements: control unit / camera 100-240 V, 50/60 Hz / 24 V AC/DC
b | Recording System | 1
Examinations & Minor procedures can be recorded in a cost-effective manner
Functions
Multi functional image & video recorder
Recording images & videos up yo full HD quality via camera
Enter patient data in combination with the keyboard
Storage of data on an external thumb drive or the internal data hard drive
Provision of data in the network
Video Signal input (DVI, SDI, S-VIDEO and Composite)
Recording Indicator
Recording indicator should disply in the upper-left corner of the monitor
Patient Data
Button on the front panel enables shuld direct access to the patient information screen
Image Options
Should be enable the section of image formates (e.g. JPG or TIFF)
Specification
Video Signal inputes: DVI, HD, SDI, Composite, S-Video, RGB, YPbPr
Video Out put resolution: 1920x1080, 1280x720, 1024x768, 800x600 640x480
Internal Hard drive 320 GB
USB Ports USB 2.0
Network RJ45/Connections as network drive (SMB)
Recording Formats Videos: H264mp4
Images: JPG, TIFF, BMP
Patient Data Saved as .txt file and / or in EXIF format
Consisting of:
19 Monitor (Local)
Key Board /Mouse (Local)
| Serial # 4 | 1
IABP (Intra Aortics Balloon Pump
Serial # 5
Color Doppler Ultrasound with Imaging Modes System Software
and Transducer.
Main Features of System
15” TFT LCD or more Color Monitor
Operational Modes: B Mode, M Mode, Doppler Mode, Color Flow Maping, BFLOW / e flow
Color Flow Mapping in M Mode, HPRF, Duplex/Triplex modes
Power Doppler, Doppler, Directional Power Doppler.
Frame Rate 745 FPS or more
Scanning Dept: 30 cm or more
System Advance Application Software:
Compound precision Imaging /Cross Beam/M-View
System Should Be Capable Of Following Up gradation:
B flow/eflow
Probes
MULTI-FREQUENCY CONVEX PROBE
MULTI-FREQUENCY LINEAR PROBE
MULTI-FREQUENCY ENDOCAVITY PROBE
Other Third Party Accessories UPS
Imported UPS with8-10 minutes backup.
Digital B&W Thermal Printer
| SIGNATURE CONTRACTOR NAME |  |  | DATE & SEAL

## Cardic Bypass Surgery
ESTABISHMENT OF MADAM FARYAL TALPUR CARDIC SURGERY COMPLEX AT CMCH LARKANA, (i) CORONARY ANGIOGRAPHY CARDIC CATHETERIZED LAB (ii) CARDIC BYPASS SURGERY (iii) CHEST SURGERY
2. List of Machinery Equipments (B)
S No. | Description |  |  | Qty | Unit Rate on CIF Basis Figure & Words | Unit Rate on FOR Basis Figure & Words | Total Cost (Figure & Words
Serial # 1
| ICU Ventilator ( Infant to Childran and Adult) |  |  | 6
Serial # 2
| Patient Monitor upto 7 parameters. |  |  | 10
2IBP, ECG, TEMP, NIBP, SPO2, 4WAY, RESP with Bettery backup 8 hours.
TECHNICAL SPECIFICATIONS
Monitor size: Weight: 360mm x273 mm x 122mm
Display Type: 12.1"color LED backlight LCD
Resolution: 800 x 600 pixels
Waveforms: up to 8
Temperature range: 0 to 50 °C (32 to 122 F)
Battery Type: Chargeable Lithium-ion
Run time: 2 hrs. (2600mAh)
INCLUDES: -
12.1" Anti-glare Screen, ECG,NIBP, Temp, SpO2
Oxygen, Ventilation, Renal calculations
Electrode+cable+wires: Adult, 5-lead, Snap, AHA
Reusable SpO2Finger Sensor, Adult.
Reusable NIBP Cuff, Adult.
Reusable Temp probe, Adult, Skin, 2 Pin
1 Li-ion Battery (11.1V, 2600mAh)
Power cord – EU.
2 CHANNEL IBP MODUEL ACCESSORIES
IBP cable assembly 6 Pin x 1.
Disposable transducers x 5.
SIDESTREAM ETCO2 WITH ACCESSORIES:
Including:
Dry-line water trap, Adult, 2 pcs.
Sampling line, Adult, 2.5 m, 2 pcs.
CO2 Nasal sample cannula, Adult, with 7' line, 2 pcs.
Dry-line airway adapter, straight, 2 pcs.
Serial # 3
| Pacemaker |  |  | 5
Serial # 4
| Portable Heavy Duty Suction Machine (34 L/PM) |  |  | 4
Serial # 5
| Low Power Suction Machine |  |  | 4
Serial # 6
| Syringe Pump |  |  | 6
Syringe pump offers a range of features suited todrug therapy, blood and blood products
Including: large, cleardisplay, intuitive user interface, integrated design and wide rangeof
specialty infusion sets.
TECHNICAL SPECIFICATION:
Infusion Rate Range (With Titration if Enabled)
5ml syringe – 0.1 to 150ml/h
10ml syringe – 0.1 to 300ml/h
20ml syringe – 0.1 to 600ml/h
30ml syringe – 0.1 to 900ml/h
50ml syringe – 0.1 to 1200ml/h
VTBI Range: 0.1 to 1000ml, 1min to 24 hrs.
VI Range: 0.0ml to 9990ml
Volumetric Accuracy: +/- 2% (nominal)
Purge Rate: 100 to 500ml/h (limited by syringe size)
Purge volume range is 0.5ml - 5ml
Purge volume: Range 0.5 to 5ml
Bolus Rate: 5ml syringe - 10 - 150ml/h
10ml syringe - 10 - 300ml/h
20ml syringe - 10 - 600ml/h
30ml syringe - 10 - 900ml/h
50ml syringe - 10 - 1200 ml/h
Bolus Volume: 0.1-25ml (hands-on and hands-free)
Drug Names/Protocols: Configurable for up to 100 drugs per profile
Alarms, Prompts &
Warnings: Alarm beacon with alarm (red) & warning (amber) indicators
Occlusion Settings: 11 Levels from L0-L10.
BackOff: Reduces post occlusion bolus
Event Log: min 1500 events
Key Press Log: 200 Key presses
Battery Type: Rechargeable NiMH
Battery Life: Operates for 6 hrs.from full charge
Recharge Time: 2.5hrs from discharge to 90% charge
Serial # 6
| Infusion Pump |  |  | 6
INFUSION PUMP:-
Air-in-line sensor: The air-in-line sensor provides reliable detection of airin the pump fluid
path and therefore reduces the risksassociated with accidental air infusion to the patient.
| Titration | Modes, For seamless infusions, this useful mode allows you tochange the flow rate
without stopping the pump.Free flow protection:The built-in flow stop feature helps reduce
the risksassociated with accidental freeflow. Anti-bolus feature: The built-in bolus function is
designed to reduce thebolus that may occur upon the release of an occlusion,following a
downstream occlusion alarm.
TECHNICAL SPECIFICATION:
| Rate Range: |  | 1 to 999 ml/hr – Standard mode
1 to 99.9 ml/hr – Micro mode
Volumetric Flow
| Rate Accuracy |  | ±5% at 25ml/hr in accordance to IEC60601-2-24
Volume to be Infused
| (VTBI) |  | 1 to 9999 ml – Standard
0.1 to 999 ml – Microl
| Volume, | 0 to 9999 ml – Standard
| Infused (VI) |  | 0.0 to 9999 ml – Micro
| KVO Infusion Rate |  | Configurable: off to 5 ml/hr
| Occlusion Pressure | User selectable – low 250 mmHg,
Normal 350 mmHg, high 500 mmHg
| Battery |  | Rechargeable NiMH
| Battery Life |  | 10h at 25ml/hr when fully charged
| Recharge Time |  | 24h from discharge to 95% charge
| Air-In-Line Detection | Single bubble: 50, 100 (default), 250, 500μl
Bubble accumulation
Bolus Capability
| During Infusion |  | Bolus Flow Rate – 1-999 ml/h
Bolus Volume – 0-99 ml
Serial # 8
| Fase Mask |  |  | 25
| Video Laryngoscope with Monitor |  |  | 2
Latest CMOS technology based scope
Integrated LED Light Source
Fibreless scope with a camera chip on the tip technology
Image quality of 110,000 pixels or more
Controls button for Imaging, White Balance & Video Recording
Deflection up/down: 140°/140°
Direction of view: 0°
Angle of view: 100°
Working length: 65 cm or better
Total length: 93 cm or better
Working channel inner diameter: 1.5 mm or better
Including Accessories:
| Leakage Tester |  |  | 1
| Tube Holder |  |  | 1
| Cleaning Brush |  |  | 1
| Irrigation Adaptor |  |  | 1
| Suction Valve |  |  | 1
| Bronchoscope Insertion Tube, size 4 |  |  | 1
| Bronchoscope Insertion Tube, size 2 |  |  | 1
| Plug, package of 10 |  |  | 1
| Protection Cap |  |  | 1
A | CMOS Technology based Video Laryngoscope
Standard Macintosh blades
Angle of View 80o or better
CMOS Technology with integreted LED light source
Blade tip should be visible on the monitor
Can also be used for direct laryngoscopy in case of power failure
Following blades are required:
| Macintosh Balde Size 4 |  |  | 1
| Macintosh Balde Size 2 |  |  | 1
B | Portable Monitor |  |  | 1
7" or more TFT Wide Angle View Display
Can connect Video Endoscope & Video Laryngoscope with toggle button
Video Imaging in 4:3 format
High image resolution with 110,000 pixels or more
Documentation with SD Storage and backup on USB memory
Connection for use of two video endoscopes
HDMI output for connection to a external monitor
Play back system for images & videos
Operates even during charging
2 hours battery back-up
C | Retromolar Intubation endoscope 5x40 |  |  | 1
For unexpected difficult airway. For use in the case of minimal mouth opening
>1cm possible. Continuous Oxygen flow via tube adaptor. Autoclavable
Outer Diameter: 5mm
For ETT: > 5.5mm
Useable sheath length: 40cm
Distal Blending 40O
Moveable eyepiece
Tube holder for ETT & Oxygen flow
| Tray for cleaning & sterilization of retromolar endoscope |  |  | 1
| Battery Light Source LED for endoscopes rechargeable brightness>110Im, |  |  | 1
lithim ion batteries charging time 60 min max suitable for wipe disnfection
| Charging unit for battery light source |  |  | 1
| One-chip CMOS camera head, resolution 640 x 480, focal length f = 20 mm |  |  | 1
compatible with all endoscopes.
| Imported Stand from the same manufacturer |  |  | 1
| Plastic sterilization container for flexible endoscopes |  |  | 1
Serial # 9
| Breathing Circuit |  |  | 50
Serial # 10
| Ventura Mask |  |  | 34
Serial # 11
| AMBU Bags (Adult & Peads) |  |  | 25
Serial # 12
| Dressing Set |  |  | 4
Serial # 13
| Spot Light ( Mobile OT Light) |  |  | 2
Light system
High-performance LEDs. Computer-calculated optical systems boasting a special
microstructure enable the color shadow-free illumination of small or large surgical sites.
Operation
All operating functions can be controlled via an ergonomic,flush-integrated
membrane keypad.Brightness control and light field adjustment are fully electronic.
setting is indicated via an LED segment display on the control panel. High-precision
positioning of the light head in any surgical situation is guaranteed by integrated handle
elements that are accessible from all sides
TECHNICAL DATA
Central illuminance at a distance of 1 m
Effective service life of light sources
Color rendering index [Ra (1 – 8)]
Red rendering index [R9]
Light head dimensions
Color temperature
Light field diameter (d10, distance 1 m)
Light field diameter (d50, distance 1 m)
Central illuminance, electrically dimmable from / to
Illumination depth (20 % on Ec max. without refocusing)
Illumination depth (60 % on Ec max. without refocusing)
Permissible ambient temperature range during operation
Soft-start option (light does not start immediately with full brightness)
Electronic system meets requirements acc. to VDE and IEC
Battery recharging time
Operating time with battery
Classification acc. to MPG (Medical Devices Act / MDD)
Protection class acc. to IEC 601
Conformity
| Sterilizable Handel |  |  | 1
Local UPS (Best Quality)
Standard Accessories are Included :-
Serial # 13
| Emergency Mobile OT Light |  |  | 2
Light system
High-performance LEDs. Computer-calculated optical systems boasting a special
microstructure enable the color shadow-free illumination of small or large surgical sites.
Operation
All operating functions can be controlled via an ergonomic,flush-integrated
membrane keypad.Brightness control and light field adjustment are fully electronic.
setting is indicated via an LED segment display on the control panel. High-precision
positioning of the light head in any surgical situation is guaranteed by integrated handle
elements that are accessible from all sides
TECHNICAL DATA
Central illuminance at a distance of 1 m
Effective service life of light sources
Color rendering index [Ra (1 – 8)]
Red rendering index [R9]
Light head dimensions
Color temperature
Light field diameter (d10, distance 1 m)
Light field diameter (d50, distance 1 m)
Central illuminance, electrically dimmable from / to
Illumination depth (20 % on Ec max. without refocusing)
Illumination depth (60 % on Ec max. without refocusing)
Permissible ambient temperature range during operation
Soft-start option (light does not start immediately with full brightness)
Electronic system meets requirements acc. to VDE and IEC
Battery recharging time
Operating time with battery
Classification acc. to MPG (Medical Devices Act / MDD)
Protection class acc. to IEC 601
Conformity
| Sterilizable Handel |  |  | 1
Local UPS (Best Quality)
Standard Accessories are Included :-
Serial # 16
| Pharmacy Refrigerator |  |  | 2
Serial # 17
| Electro Hydrulic Operating Table |  |  | 1
Five sections Electro-hydraulic operating table,
- Adjustment possibilities and ranges meets the requirments
demanded by a modern universal operating tables
- The operating table made of stainless, acidproof steel.
Features
Modren Design
Battery powered, with facility for connection to
mains electricity for immediate use.
Unique software providing smooth and precise table movements.
Superior electro- hydraulic system, with technical solutions of the world
leading micro-hydraulics providers.
Controls
Optional override control system, mechanic-hydraulic, used in case of
corded hand control damage, failure of drive or main control system.
Range of table movements controlled via the hand control
- Height adjustment
- Table top tilts: lateral and longitudinal
(Trendelenburg and Rev- Trendelenburg)
- Back rest section and kidney bridge (option)
- Longitudinal slide (option)
Adjustment angles of 40° Trendelenburg/Trendelenburg and and
30° tilt make it ideal solution for minimally invasive procedures.
Range of table movements controlled mechanically or gas spring supported:
- Backrest section and kidney bridge,
- Footrest sections,
- Headrest sections,
- Longitudinal slide
TECHNICAL PARAMETERS
Total length of the table with 5 sections tabletop
Tabletop width
Total tabletop width (with side rails)
Standard height adjustment, with mattress
(electro-hydraulic)
Extended height adjustment with mattress
(electro-hydraulic)
Lateral tilts adjustment (hydraulic)
Trendelenburg / reverse-Trendelenburg
adjustment (electro-hydraulic)
Backrest adjustment (pneumatic)
Optional backrest adjustment (electro-hydraulic)
Footrest adjustment (pneumatic)
Footrest abduction (mechanic)
Headrest adjustment (pneumatic)
Kidney bridge adjustment (mechanic),
optional (hydraulic)
Longitudinal slide(mechanic),Optional(hydraulic
Table mass
Power
Batteries (sealed, maintenance-free)
Battery charging time
Working time between battery charging s
Type of operations
Battery charger
Charger power
Power consumption from mains
Maximum working load
Lifetime
Degree of protection
Class of protection against electric paralysis
Application part type
Following Accessories are Included
| Specialistic head rest with adjustable track |  |  | 1
| arm support /with one-position clamp (One Pair) |  |  | 2
| Hand strap (One Pair) |  |  | 2
| Anesthetic frame with adjustable width |  |  | 1
| Knee support /with multiposition clamp (One Pair) |  |  | 2
| Abdominal Strap |  |  | 1
| Infusion Holder |  |  | 1
| bowl |  |  | 1
| Wireless hand controller |  |  | 1
Serial # 18
| Anesthesia Machine with Ventilator with backup support |  |  | 2
Tidal Volume setting20-1500ml,including: 8.4 inch color screen, ventilator, glass flowmeters
(hypoxic guard),work surface, 1 drawer, 2 sets of flowsensor, 1 battery (90 min), 20pcs of foam
pad, pipeline pressure gauges, user manual, quick user guide, packing material System
software, including: VCV mode, Pressure controlled ventilation mode (PCV) SIMV
mode (SIMV-VC and SIMV-PC) Pressure supported ventilation mode (PSV,
w/apnea backup) TV compensation, pressure monitoring, volume monitoring, PEEP
European (ID) style NIST-2 (European) O2 & N2O & Air - inlet with hoses O2 drive
2 positions of vaporizer (One vaporizer included). PSU Breathing circuit
(with manual ventilation arm)
O2 monitor and accessory kit, including:
O2 sensor, Medicel MOX-2, 1
O2 sensor cable Disposable accessory kit (Adult), including:
Patient circuit (breathing bag 3L, Adapter Wye, Adapter elbow,
Airway filter), adult, 1pcs
Large adult mask, Air-Soft, w/valve, 22 mm, 1pcs Passive AGSS system,
including accessory kit: Tube, 32mm, 1pc
Adaptor, 1pc Auxiliary O2 flowmeter PIN index cylinder yokes
(O2 and N2O) Auxiliary sockets LED lights Power cord (Europe)
PHYSICAL CHARACTERSITICS
Dimenions
| Height |  | 1300mm or more
| Width |  | 700mm or more
| Depth |  | 600mm or more
| Castors: |  | 125 mm with brakes
| Drawers |  | 2 or more drawers with mounting shelf
Modes of Ventilation
VCV, PCV, SIMV-VOLUME, SIMV-PRESSURE, PSV (OPTIONAL).
Ventialtor parameters:
| tidal volume |  | 1500ml
| Pressure: |  | 0 - 60cmH2O
| Rate: |  | 4 - 100bpm
| PEEP |  | Integrated electronic controlled
| FiO2 |  | 0 - 100%
Vaporizer:
2 positions, selectatec, Halothane, EnfluraneSevoflurane, Isoflurane
Pipeline inlet pressure
Gas Type:
O2
NO2
Air
| pipeline input range: |  | 0 to 6 bar
Manometer Gauges
| N2O, Air, O2 |  | 0 – 10 bar
Rotameter:
| O2 tube |  | 1 – 10 L/min
| N2O tube |  | 1 – 10 L/min
| Air |  | 1 – 10 L/min
| Volume of CO2 absorber: |  | 1.5 liters (1500ml)
| Audible and Visual alarm: |  | Low/high airway pressure, Apnea alarm,
O2 alarm, Tidal volume alarm,
upper/lower limit alarm, user and technical alarm,
major leak or patient disconnection.
| Battery |  | Li-ion (built-in)
Power Requirement:
Power input: 100-240 VAC, 50/60 Hz
Serial # 19
| Suction Machine |  |  | 4
Serial # 20
| Diathermy Machine with backup support |  |  | 2
Electrosurgery Unit 400 watts
Functional Test
Progressive power controll
Endo Mode" Time Controll cuting
Bipolar coagulation with foot switch
Connector for Monopolar hand switches
Multifunctional connector for Bipolar instruments
Monopolar cutting & coagulation by foot Switch
Technical Specification of the Unit
Supply voltage: 220-240 V; 50-60 Hz or 100-127 V; 50-60 Hz
Selectable via dummy plug located inside the unit by Technical Service
Power input: with no HF power output: approx.16 VA,
at max. power output: approx. 800 VA
Class of protection : I
Classified acc. to MDD : II b
Leakage currents LF and HF : In acc. with EN IEC 60601, Part 2-2
Type of equipment : CF; defibrillator-proof
Nominal frequency : 450 kHz
Pulse frequency : 30 kHz
HF output power:
Type of current Power Crest factor Voltage
Cutting 1 : max. 400 W at 300 Ω 1.6 at 300 Ω max. 2300 Vpp
Cutting 2 max. 300 W at 300 Ω 1.9 at 300 Ω max. 2500 Vpp
Endo-Mode max. 100 W at 200 Ω 1.6 at 200 Ω max. 2800 Vpp
Contact coagulation max. 250 W at 200 Ω 3.4 at 200 Ω max. 3200 Vpp
Spray coagulation max. 120 W at 300 Ω 5.6 at 300 Ω max. 6000 Vpp
Bipolar coagulation max. 100 W at 100 Ω 2.1 at 100 Ω max. 600 Vpp
Duty type intermittent INT 10 s/ 30 s, equivalent to a duty factor of 25%
220-240 V: T 4 A (slow-blow.)
100-127 V: T 8 A (slow-blow.)
Signal level : HF activation: 55 dB (A) (adjustable between 50 dB and
60 dB by Technical Service)
Weight : 8.6 kg
Interference suppression : Limits in conformity with EN 55011,
Interference immunity in conformity with IEC 801
Dimensions : 405 mm x 135 mm x 380 mm (W x H x D)
CE marking conform with 93/42/EEC
Accessories: (for each set)
| Double-pedal foot switch |  |  | 1
| Rubber neutral electrode for adults single 15X26 cm |  |  | 1
| Non stick Red forceps, Straight, Blunt. 0.6 mm, 20 cm |  |  | 1
| Bipolar Cable |  |  | 1
| Reusable Electrode Handle with two push buttons |  |  | 1
| Knife Electrode, 30mm |  |  | 1
| Trolley (Local - Best Quality) |  |  | 1
Serial 21
| DEFIBRILLATOR |  |  | 2
TECHNICAL SPECIFICATIONS
Display Screen Active LCD screen, with visible viewing area
| Waveform | display should display waveform sweep speed for ECG and SPO2
| Manual | defibrillation mode: Should have energyy select range from 2 to 360 with user
configurable sequence
| External | defibriilator AED with shock advisory system
| Pacing | Available with pacing option, with pacing rate of 40 - 180ppm
| Output | waveform Monophasic output waveform
| Calibration: | Unit should be able to measure impedance of a patient and
calibrate defibrillation waveform and amplitude
| waveform | Type: Biphasic Truncated Exponential
| Modes: | In-service mode, Archieve mode
| ECG n | lead selectio3 wire / 5 wire ECG cable selection option
| Patinet's | Data management and transmission Should be able to record/store patient's data
| Heart | rate range 20-350bpm
| Heart | rate display Digital
| Battery | Rechargeable internal battery
Standard Accessories:
Hard paddles (detachable, Adu/Ped)
3-lead ECG cable
Adult ECG electrodes
Pacing electrodes
Therapy cables
ECG printer papers ( 50 x 30 mm)
Trolley
Serial # 21
| Mobile C-Arm Image Intensifier X-Ray System with dual Monitor |  |  | 1
Image intensifier on “C” arm of wide depth and with orbital/skull-caudal rotation
Generator: High Frequency Monoblock, power rating 5kW, 110kV, 35mA, 125mAs. Continuous
and pulsed fluoroscopy 8mA. Snap shot 10mA.
X-ray tube: Stationary anode, 0.5/1.5mm focus
9” Image Intensifier: Triple field 9” – 6” – 4”,
useful input diameter 215/160/120mm
TV Camera: CCD, high resolution, 625/50 TV lines. Image rotation 360°, automatic return to 0°.
L.I.H. + 330 frames non volatile memory + 2° 19” LCD monitors – L/R image inversion – edge
enhancement Real time Digital Rotation – Grey scale inversion, with keyboard for editing
data patient, Collimator: Motorised with iris diaphragm
Serial # 22
| Ultrasound Therapy |  |  | 1
Serial # 23
| TENS Machine (Transcutaneous Electrical Nerve Stimulation) |  |  | 2
Serial # 24
| IABP (Intra Aortics Balloon Pump) with Backup Support |  |  | 1
Serial # 25
PULSE OXIMERTER: -
| SpO2, Pulse Rate, Built-in Battery (Lead-Acid) |  |  | 2
Suitable for adult, pediatric and neonatal patients.
3.4" high resolution LCD for display of plethysmograph, trended parameter data, alarm events
and menus 0.8” large LED displays
Large, bright LEDs provide optimal visibility, whether across the bed or across the room
Smart and stable design keeps firmly in place when activating buttons, leaving your other hand free
to perform other tasks
ACCESSORIES INCLUDES:-
Main unit
512F Adult Finger Clip SpO2 sensor (Reusable).
6 Pin Extension Cable.
1 lead acid battery
Serial # 26
| Medical Central Oxygen Supply System with Compressor (Air & Suction) |  |  | 1
Serial # 27
| Surgical Operating Instruments and Material Set |  |  | 1
Serial # 28
| HearT Lung Machine with Backup Support |  |  | 2
| Serial # 29 |  |  |  |  | FEATURES:- 2
| Hypothermy Machine |  |  | 2
Serial # 30
| Heavy Duty Autoclave Capacity (300 to 400Liters) |  |  | 1
Serial # 31
| Defibrilator |  |  | 2
TECHNICAL SPECIFICATIONS
Display Screen Active LCD screen, with visible viewing area
| Waveform | display should display waveform sweep speed for ECG and SPO2
| Manual | defibrilation mode: Should have energyy select range from 2 to 360 with user
configurable sequence
| External | defibriilator AED with shock advisory system
| Pacing | Available with pacing option, with pacing rate of 40 - 180ppm
| Output | waveform Monophasic output waveform
| Calibration: | Unit should be able to measure impedance of a patient and
calibrate defibrillation waveform and amplitude
| waveform | Type: Biphasic Truncated Exponential
| Modes: | In-service mode, Archieve mode
| ECG n | lead selectio3 wire / 5 wire ECG cable selection option
| Patinet's | Data management and transmission Should be able to record/store patient's data
| Heart | rate range 20-350bpm
| Heart | rate display Digital
| Battery | Rechargeable internal battery
Standard Accessories:
Hard paddles (detachable, Adu/Ped)
3-lead ECG cable
Adult ECG electrodes
Pacing electrodes
Therapy cables
ECG printer papers ( 50 x 30 mm)
Trolley
| Serial # 32 |  |  | 1
ABG Machine (Blood Gas Analyzar)
| Serial # 33 |  |  | 2
ICU Slush
Serial # 34
| ECG Machine-12 Channel:- |  |  | 2
FEATURES:-
8-inch high-resolution color display touchscreen. User friendly workflow buttons. Modern designed soft touch hotkeys. Standard keyboard layout. Patient demographics can be retrieved from the work-list. Anatomically designed cable to minimize tangling and lead reversal. Acquire, analyze, displays, and record 12-lead ECG information. Provide ECG algorithm to automatically analyze the acquired ECG waveforms; output measurements and diagnosis. Support auto measurement, manual measurement, and rhythm measurement. Print ECG reports through either an internal thermal recorder or an external printer. Store, preview, and review ECG reports. Support entering patient information through the keyboard or a barcode reader. Present messages in case of lead off, interference, low battery, or other abnormity.
Technical Specfications:-
Display: 8 inch color LCD with LED backlight
Resolution: 800 × 480 pixels
Measurement mode: Auto, manual, rhythm
Lead type: 12-lead
PACE pulse markers:
Amplitude: ±2 mV - ± 250 mV
Width: 0.1ms - 2ms
Rise time: < 100μs
Amplitude: ≥0.2 mV RTI
Build-in thermal recorder
Number of waveform channels Max. 12
Paper speed: 5 mm/s, 12.5 mm/s, 25mm/s, 50 mm/s
Accuracy: ±5%
Resolution: Vertical resolution: ≥8 dots/mm
Standard Accessories:-
Patient Cable, Limb Electrode, Rechargeable Battery, Power Cord.
Serial # 35
| ETT Machine with Heavy Duty Trade Mill |  |  | 1
EXERCISE TOLERANCE TEST (ETT) MACHINE: ‐
Exercise ECG Test System downloads information to the Workstation Software,
which provides paperless efficiency as well as connectivity to many leading EMR Systems.
This software puts patient records from a variety of diagnostic instruments at your fingertips,
giving you the ability to store, access, review, compare and edit diagnostic test data from any
computer on your network.
Multiple clinical features will meet your testing requirements.
Performs all commonly used stress protocols, and has the ability to customize additional
protocols for personal preference.
Displays leads in 3‐, 6‐ or 12‐lead screen formats.
ST measurements reference point can be modified at any point before, during, or after the test.
oCalculates average ECG complexes at regular intervals and
superimposes them on a reference ECG complex, which
highlights ST changes in detail. o Displays ST level and slope measurements both in a trend graph
and in a table during the test.
Detection of leads‐off, arrhythmias, and pacemakers.
Calculates Duke Score and supports Borg Scale (patient effort) entry.
MEANS ECG INTERPRETATION SOFTWARE: ‐
By employing a new, world‐class version of the MEANS interpretation program for full
morphology and rhythm analysis, the PC‐based Resting ECG is able to provide extremely
accurate interpretive statements. The MEANS interpretive software is based on 25 years of
research from the renowned University of Rotterdam. A pediatric version of the interpretation
software is also available.
12‐lead Exercise ECG Test System.
Resting ECG Functionality.
Digital ECG Patient Cable with USB Connections.
Stress Belt, Stress Electrodes, Manual.
Desktop Computer and Laser Printer (Locally Provided).
Heavy Duty Treadmill
Frame: 10 guage steel side rails / 12 guage steel motor pan
Powder coat
Speed Range: 0.5‐12 mph
2 hp motor; 220 volt
Elevation Range: 0‐25% Standard.
Running Surface 22” x 63”
Antislip strips
Ultra quiet, long running surface
User capacity: 500 lbs.
Floor surface requirements: Level 33” x 88” surface
Serial # 36
| Endoscopic Saphenous Vein Harvesting System |  |  | 2
1 | Endoscopic Vein Retractor, model, for harvesting the great saphenous vein, enhanced
telescope guiding channel, distal width 25 mm, working length 27 cm, with integrated
guide in handle for fiber optic light cable, autoclavable, incl. Cleaning Adaptor,
| for use with Telescope |  |  | 1
2 | Telescope 45°, diameter 5 mm, length 29 cm, autoclavable, fiber optic light transmission
| incorporated. |  |  | 1
3 | Vein-Dissector, blunt, distal, angled to left, size 3 mm, working length 41 cm, autoclavable |  |  | 1
4 | Vein Dissector, blunt, distal angled to right, size 3 mm, working length 41 cm, autoclavable |  |  | 1
5 | Forceps, rotating, dismantling, with connector pin for bipolar coagulation, width of the jaw
2 mm, 45° angled, to grasp and coagulate vessels, double-action jaws, size 5 mm, length 43 cm.
| Plastic Handle, Metal Outer Sheath, insulated, Forceps Insert. |  |  | 1
6 | Scissors, rotating, size 5 mm, length 43 cm, with serrated jaws, curved spoon blades, length
of blades 17mm, double action jaws.
| Plastic Handle, without ratchet , Outer Tube, insulated, Scissors Insert |  |  | 1
7 | Hook Scissors, rotating, with connector pin for unipolar coagulation, size 5 mm, length
43 cm, single action jaws.
| Plastic Handle, withour ratchet, Outer tube, insulated, Scissors Insert. |  |  | 1
8 | Dissecting and Grasping Forceps, rotating, with connector pin for unipolar coagulation,
size 5 mm, length 43 cm, long, double action jaws.
| Plastic Handle, without ratchet, Outer Tube, insulated, Forceps Insert Insulated Metal. |  |  | 1
9 | Dissecting Forceps, rotating, with connector pin for unipolar coagulation, size 5 mm, length
43 cm, double actionjaws,
| Plastic Handle, without ratchet, Outer Tube, insulated, Forceps Insert. |  |  | 1
10 | Clip Applicator for ligating clips , (medium),dismantling, rotating, with ratchet to lock the
jaw part holding the clip, size 10 mm, length 36 cm,
| Metal Handle, with ratchet, Metal Outer Tube, Insert. |  |  | 1
11 | Titanium-Clips, medium, box with 10 sterile cartridges, 25 clips each, for use with applicators |  |  | 1
12 | Endo-Loop Ligature, with knot, for bleeding stumps, with absorbable synthetic thread, sterile,
| package of 12, USP 0 |  |  | 1
13 | Bipolar High Frequency Cord, length 300 cm, for AUTOCON® II 400 SCB system (112, 114, 116,
| 122, 125), AUTOCON® II 200, AUTOCON® II 80 and Valleylab coagulators |  |  | 1
14 | Unipolar High Frequency Cord, with 4 mm plug, length 300 cm, for models , Erbe type T,
older models and Ellman.
Serial # 37
| ECG Machine 3 Channel |  |  | 2
Serial # 38
| 300MA & 100KV Mobile X Ray Sytem |  |  | 1
Technical Characteristics of 300mA at 100kV X-Ray System
Micro processer control, Three Phase, 32 KW High Frequency X-Ray
Automatic Line Compensation
Over Load Tube Protection
Radiographic kVP 40 – 150kV
Radiographic mA 50-500mA, 11 steps
(2) mAs range: 0.5 – 800mAs
(3) Exposure time: 2.5ms – 57s
(4) Floor Mounted Tube Stand
Column rotation (+- 180 degree)
(5) Collimator
Manual collimator with 4 pairs of leafs.
Haleogen Lamp
(6) Wall Bucky Stand
(7) X-Ray Tube
(300kHU, 0.6/1.2, 15/45kW, 12°)
Small focus 0.6, large focus 1.2
Anode Heat Capacity 300KHu
Anode rotation, speed 8.5min
(8) X-Ray High Voltage Generator
X-ray high-voltage generator (32kW, Single-phase, 400VAC class)
(9) Vertical Bucky Radiographic unit
(stationary, 60lines, 10:1, f=180cm)
(10) Bucky Radiographic Table with Fixed Tabletop
(stationary, 60lines, 8:1, f=100cm)
Other Third Party Accessories
a. | Lead Approns
2 Lead aprons (Medium/Large Male/Female)
b. | UPS
UPS for whole X-Ray system:
Imported UPS with AVR capable to operate the whole system for 8-10 minutes.
Serial # 38
500MA Static X-Ray System
a | Technical Characteristics of 500mA at 100kV X-Ray System |  |  | 1
Micro processer control, Three Phase, 50 KW High Frequency X-Ray
Automatic Line Compensation
Over Load Tube Protection
Radiographic kVP 40 – 125kV
Radiographic mA 10-800mA, 20 steps
(2) mAs range: 0.5 – 800mAs
(3) Exposure time: 1 ms – 8 ms
(4) Floor Mounted Tube Stand
Column rotation (+- 180 degree)
(5) Collimator
Manual collimator with 4 pairs of leafs.
Haleogen Lamp
(6) Wall Bucky Stand
(7) X-Ray Tube
(300kHU, 0.6/1.2, 40/100kW, 12°)
Small focus 0.6, large focus 1.2
Anode Heat Capacity 300KHu
Anode rotation, speed 8.5min
(8) X-Ray High Voltage Generator
X-ray high-voltage generator (50kW, Three-phase, 400VAC class)
(9) Vertical Bucky Radiographic unit
(stationary, 60lines, 10:1, f=180cm)
(10) Bucky Radiographic Table with Fixed Tabletop
(stationary, 60lines, 8:1, f=100cm)
Other Third Party Accessories
a. | Lead Approns
2 Lead aprons (Medium/Large Male/Female)
b. | UPS
UPS for whole X-Ray system:
Imported UPS with AVR capable to operate the whole system for 8-10 minutes.
a. | Training:
03 days training on site training.
| SIGNATURE |  | CONTRACTOR NAME | DATE & SEAL

## Lab Items
CHANDKA MEDICAL COLLEGE HOSPITAL LARKANA
FINANCIAL BID
Supply, Installation, Testing, commissioning & Maintenance Durign maintenance period of Medical Equipment / Instruments under Revenue Component ADP Scheme No 373 Namely Upgration of SURGERY DEPARMENT LIST OF MACHINERY / EQUIPMENTS Modernization of Laboratories, Blood Bank, CMC HOSPITAL LARKANA FOR THE YEAR 2015-2016
S No. | Description |  |  | Qty | Unit Rate on CIF Basis Figure & Words | Unit Rate on FOR Basis Figure & Words | Total Cost (Figure & Words
Serial # 1
| Plasma Extractor |  |  | 3
Manual plasma extractor EASY TO USE MANUAL SYSTEM THAT
ACCEPTS ALL KINDS OF PLASMA BAGS. TRANSPARENT PLATFORM FOR
VISUAL CONTROL OF THE PLASMA. STAINLESS STEEL AISI 304 CABINET.
Serial # 2
| Deep Freezer |  |  | 1
Serial # 3
| Platlet Agitator |  |  | 1
Horizontal shakers for platelet bags
HORIZONTAL MOVEMENT AT A FIXED SPEED OF 66 R.P.M., LATERAL SHUFFLE OF 40 MM.
CAPACITY FOR 54 PLATELET BAGS.
FEATURES Epoxy covered case with treated aluminium platform with AISI
304 stainless steel racks for bags. Stirring automatic restart in 30 seconds
and pause device for bags handling. Acoustic alarm in case of power failure.
LCD display informing of the events log: it counts the number of power
failures and their duration. When power is restored, the device
automatically restarts the stirring.
CONTROL PANEL Single button to adjust:
-Date and time programmation. -Moving to review the number of
incidences. -Shaking stop button for 30 seconds. -Shaker stop time
programmation. -Alarms reset.
Serial # 4
| Cryo Centrifuge |  |  | 1
Refrigerated high capacity microprocessor controlled centrifuge suitable blood bags
CONTROLLABLE TEMPERATURES FROM 0 TO 35 °C. CAPACITY
ml AND ESPECIALLY FOR TRIPLE AND QUAD BLOOD BAGS FROM 500-800 ml.
MAXIMUM CAPACITY 12 BAGS.
Maximum volume: . . . . .9600 ml
Maximum bucket volume: . . . . . . . .1600 ml
Speed: . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . from 300 to 4070 r.p.m.
Accuracy: . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .± 5 r.p.m.
Resolution r.p.m.: . . . . . . . . . . . . . . . . . . . . . . . . . .10 r.p.m.
R.C.F.: . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .from 29 to 5370
Imbalance: . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .± 5 g
Accuracy: . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . ..4 g
Temperature range: . . . . . . . . . . . . . . . . . . . . . . . . from -10 °C to 35 °C
Temperature stability: . . . . . . . . . . . . . . . . . . . . . . . ±4 °C
Resolution: . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .1 °C
Temperature range: . . . . . . . . . . . . . . . . . . . . . . . . 99 min. 59 sec.
Centrifuge radius: . . . . . . . . . . . . . . . . . . . . . . . . . .290 mm
Acceleration profile: . . . . . . . . . . . . . . . . . . . . . . . . 5 Brake
profile: . . . . . . . . . . . . . . . . . . . . . . . . . . . . . 9 linear 6 quadratic
Voltage: . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .230/400 three phase.
Power: . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .4500 V
| weight: |  | 420 kg
Serial # 5
| Scrum Electrolytes Analyer |  |  | 2
Serial # 6
| Hematology |  |  | 2
Hematology analyzer 20 parameter with WB and PD mode.
The display analysis result with 3 histograms on a colour LCD Screen.
Specifications
| Parameters | Whol blood mode, Pre-diluted mode
WBC, RBC, HGB, HCT, MCV, MCH, MCHC, PLT
LYM%, MXD%, NEUT%, LYM#, MXD#, NEUT#, RDW-SD, RDW-CV, PDW*,
MPV, PCT, P-LCR*
| Analysis Principle | WBC : DC Detection method
RBC/PLT : DC Detection method
HGB: Non-cyanide hemoglobin analysis method
| Throughput | Aprrox. 60 samples per hour
| Sample Volume | WB mode Approx. 50uL
PD mode Approx. 20uL
| Sample ID | Up to 15 characters
| Quality | Control Material: Eight check-3WP (Low, Normal, High)
| Calibrator | Material: SCS-1000
| Dimension | Main unit: 420 x 355 x 450 mm
| Printer | Built-in Thermal Printer
Serial # 7
| T B System |  |  | 2
Serial # 8
| Blood Bank Refregerator |  |  | 2
Technology in our blood.
A wide range of blood bank refrigerators to store up to 700 blood bags and meet small, medium
hospitals andand large blood centers, Analytical laboratories needs. Each product line ensures
the most secure environment for your valuable blood products. MINI and BBR are landmarks
in Blood Bank Refrigerators for reliability and lasting.
●Certified according to medical devices directive 93/42/EEC.
●Wide range of capacity (up to 700 blood bags)
●Temperature graph recorder (7 days)
●Forced air circulation
●High stability and uniformity of the internal temperature
●Stainless steel sliding drawers complete with stainless steel sliding guide which prevents rust.
●Fast temperature recovery after door opening
●Stainless steel internal chamber which grant maximum hygiene
●Triple vacuum-sealed glass door, with anti-moisture resistors
●Tropicalized execution (up to 40°C)
●Superior insulating panel
●Internal light (manual or automatic switch)
●Safety key lock
The microprocessor control system E2003
●Alarm section independent from regulation section
●Remote audio/visual alarm control
●Rechargeable buffer battery (48h in case of power outage)
●Advanced security
●User friendly control display
●Continuous monitoring
Serial # 9
| Safety Cabinet |  |  | 2
Serial # 10
| Air Condition 1.5 Ton with Stabilizer |  |  | 8
Serial # 11
| Computer P 4 with all accessiries |  |  | 2
Serial # 12
Automated Bio Chemistry Analyzer
Photometric with following features and close system for:
Clinical Chemistry Testing • Specific Protein Analysis •
Drugs of Abuse Analysis • Therapeutic Drug Monitoring •
Immunosuppressant Drug Testing & Cancer testing like MTX & Viral testing like
VEROCANAZOLE assyas and others.
●includes combined sample and reagent disk for max 6 racks.
●6-positions reagent rack use
●sample disk for 54 samples : 9-pos. sample racks, max 6 racks.
●Reagent disk positions for 42 reagents.
●9-positions sample rack use
●Cuvette loader for 360 cells. Low-volume reaction cuvette uses less reagent
and water for lower environmental impact.
●Temperature controlled incubator for 90 cuvette cells. Cuvettes are
disposable, so there is no need for extra rinsing to prevent carry-over.
●The analyzer is self-contained with no need for external water or drainage
connections, which makes the analyzer quick and easy to insta
●Intuitive user interface, supported either by a tradiotional or by a touch
screen, helps to manage the daily workload fluently.
Capacity
On-Board Sample Capacit
On-Board Reagent Capacity
Walk-away time
Water Consumption
Dimensions & Weight
Continuous access to samples, reagents and cuvettes without interrupting test processing.
Fully traceable results with archiving function.
Serial # 13
| Simple Bio Chemistry Analyer |  |  | 2
Serial # 14
| Cell Counter |  |  | 2
Serial # 15
| Plasma Seprator Manual |  |  | 2
Manual plasma extractor EASY TO USE MANUAL SYSTEM THAT
ACCEPTS ALL KINDS OF PLASMA BAGS. TRANSPARENT PLATFORM FOR
VISUAL CONTROL OF THE PLASMA. STAINLESS STEEL AISI 304 CABINET.
Serial # 16
| Multi Head Microscope |  |  | 2
Serial # 17
| Binocular Microscope |  |  | 3
Serial # 18
| Oven (Large Size) |  |  | 2
Serial # 19
| Pathology Incubator |  |  | 2
Serial # 20
| Water Bath |  |  | 2
Water bath with inner stirring
with digital control and display of temperature and time
ADJUSTABLE TEMPERATURE RANGE FROM AMBIENT +5ºC UP TO 80ºC. STABILITY
±0.1°C. SET TEMPERATURE ACCURACY ±1 °C, WORKING TEMPERATURE PRECISION 0.1 °C.
FEATURES Double skinned tank, inner case made of AISI 304 stainless steel. External
case of AISI 304 steel. Submerged heating elements in INCOLOY special alloy. It comes
complete with drain tap
STIRRING SYSTEM These baths have an aspiration and outflow internal circulation
pump which allow a constant flow stirring. Flow: 8 litres per minute.
All baths with heating shelf and drainage. Faucet included.
Toatlcacpacity Must be 20 litres for water bath
Serial # 21
| Auto Clave Machine |  |  | 1
SELECTABLE TEMPERATURE BETWEEN 40 ºC & 121ºC.
CAPACITY: 20, 80 & 150 L.
. WATER SUPPLY NEEDED WITH LOW-CAL CONTENT AND DRAINAGE. (WATER PRESURE:
3 BAR MINIMUM). CONFORMITY EN 61010-1 SAFETY FOR EQUIPMENTS FOR LABORATORY
USE . EN 61010-2-040 SAFETY FOR STERILIZER EQUIPMENTS. EN 61326
ELECTROMAGNETIC COMPATIBILITY. EN/97/23/EC SAFETY FOR PRESSURE
EQUIPMENTS.
FEATURES External case, top, tank and lid made of stainless steel. Tank fill
by external water inlet. Temperature control probe inside the sample.
Counter pressure system to prevent the cans breakage. Temperature, time
and pressure registered in USB.
SAFETY Device that avoids opening the lid when the chamber is under pressure. Safety
valve: Prevents overpressure over the maximum limit. Safety pressure switch:
Disconnects heater when there’s overpressure. Thermally insulated lid. Door properly closed d
MicroProcessor includes:
Parameters LCD display. 2.
Microprocessor single button. 3.
Operation Led. 4.Start/stop
push button. 5. USB adapter. 6.
Open lid knob. 7. USB function
light indicator. C
MICROPROCESSOR FEATURES Microprocessor controls sterilization and cooling
process: Any error message of functions or execution is displayed in the panel and an
acoustic alarm sounds, automatically cutting the steam generator out. BACK VIEW 1.
Mains input.
Serial # 22
| Weighting Balance |  |  | 3
Serial # 23
| Biological Safety cabinet |  |  | 2
Serial # 24
Centrifuge Machine
| Cetnrifuge Bench top Laboratory Centrifuge |  |  | 2
APPLICATIONS General laboratory use and teaching.
FEATURES External case made from coated ABS plastic. Rubber
suspended motor, reduces any motor vibrations. Fixed angle 6
place rotor with identification code. Transparent polycarbonate
lid that allows the user to see inside the centrifuge. Speed
selection
CONTROL PANEL Digital display of speed and time.
Push button Start-Stop. Push button selector for
speed and time. Push button increase value. Push
button decrease value.
Supply complete with 6 x 15 ml tubes with adapters and 6 adapters for 15 ml Falcon tubes
Serial # 25
| Incubator |  |  | 2
Bacteriological incubators
TEMPERATURE THERMOSTAT CONTROL WITH ANALOGUE THERMOMETER.
ADJUSTABLE TEMPERATURES FROM AMBIENT +5 °C UP TO 80 °C.
STABILITY: ±0.1 °C, UP TO 37 °C. HOMOGENEITY: ±0.5 °C, UP TO 37 °C
INTERNAL GLASS Door.
SAFETY: OVER TEMPERATURE CUT OUT INCORPORATED ACCORDING TO THE
EN.61010 STANDARD. ADJUSTABLE SAFETY THERMOSTAT DIN 12880. FITTED.
internal shelves 3-5
capacity in litres Minimum 19 litres
Serial # 26
| Co2 Incubator |  |  | 2
Serial # 27
| Vertex Mixer |  |  | 4
| SIGNATURE |  | CONTRACTOR NAME |  | DATED & SEAL

## Pathological
ESTABLISHMENT OF PATHOLOGICAL LABORATORY, CMCH LARKANA
3. List of Machinery Equipment (Pathology Laboratory) (C)
S No. | Description |  |  | Qty | Unit Rate on CIF Basis Figure & Words | Unit Rate on FOR Basis Figure & Words | Total Cost (Figure & Words
1
Serial # 1
Automated Bio Chemistry Analyzer
Photometric with following features and close system for:
Clinical Chemistry Testing • Specific Protein Analysis •
Drugs of Abuse Analysis • Therapeutic Drug Monitoring •
Immunosuppressant Drug Testing & Cancer testing like MTX & Viral testing like
VEROCANAZOLE assyas and others.
●includes combined sample and reagent disk for max 6 racks.
●6-positions reagent rack use
●sample disk for 54 samples : 9-pos. sample racks, max 6 racks.
●Reagent disk positions for 42 reagents.
●9-positions sample rack use
●Cuvette loader for 360 cells. Low-volume reaction cuvette uses less reagent
and water for lower environmental impact.
●Temperature controlled incubator for 90 cuvette cells. Cuvettes are
disposable, so there is no need for extra rinsing to prevent carry-over.
●The analyzer is self-contained with no need for external water or drainage
connections, which makes the analyzer quick and easy to insta
●Intuitive user interface, supported either by a tradiotional or by a touch
screen, helps to manage the daily workload fluently.
Capacity
On-Board Sample Capacit
On-Board Reagent Capacity
Walk-away time
Water Consumption
Dimensions & Weight
Continuous access to samples, reagents and cuvettes without interrupting test
processing. Fully traceable results with archiving function.
| Serial # 2 |  |  | 1
Easylyte For Electrolyte
Serial # 3
| Hematological Analyzer for Computerized CBC |  |  | 2
Hematology analyzer 20 parameter with WB and PD mode.
The display analysis result with 3 histograms on a colour LCD Screen.
Specifications
| Parameters | Whol blood mode, Pre-diluted mode
WBC, RBC, HGB, HCT, MCV, MCH, MCHC, PLT
LYM%, MXD%, NEUT%, LYM#, MXD#, NEUT#, RDW-SD, RDW-CV, PDW*,
MPV, PCT, P-LCR*
| Analysis Principle | WBC : DC Detection method
RBC/PLT : DC Detection method
HGB: Non-cyanide hemoglobin analysis method
| Throughput | Aprrox. 60 samples per hour
| Sample Volume | WB mode Approx. 50uL
PD mode Approx. 20uL
| Sample ID | Up to 15 characters
| Quality | Control Material: Eight check-3WP (Low, Normal, High)
| Calibrator | Material: SCS-1000
| Dimension | Main unit: 420 x 355 x 450 mm
| Printer | Built-in Thermal Printer
Serial # 4
| Microscope (Binocular) |  |  | 1
| Serial # 5 |  |  | 1
Selectra-E (Fully Automatic System
| Serial # 6 |  |  | 1
Centrifuge Machine
APPLICATIONS General laboratory use and teaching.
FEATURES External case made from coated ABS plastic. Rubber
suspended motor, reduces any motor vibrations. Fixed angle 6
place rotor with identification code. Transparent polycarbonate
lid that allows the user to see inside the centrifuge. Speed
selection
CONTROL PANEL Digital display of speed and time.
Push button Start-Stop. Push button selector for
speed and time. Push button increase value. Push
button decrease value.
Supply complete with 6 x 15 ml tubes with adapters and 6 adapters for 15 ml Falcon tubes
| Serial # 7 |  |  | 1
Distiller
Serial # 8
| Water Bath |  |  | 1
Water bath with inner stirring
with digital control and display of temperature and time
ADJUSTABLE TEMPERATURE RANGE FROM AMBIENT +5ºC UP TO 80ºC. STABILITY
±0.1°C. SET TEMPERATURE ACCURACY ±1 °C, WORKING TEMPERATURE PRECISION 0.1 °C.
FEATURES Double skinned tank, inner case made of AISI 304 stainless steel. External
case of AISI 304 steel. Submerged heating elements in INCOLOY special alloy. It comes
complete with drain tap
STIRRING SYSTEM These baths have an aspiration and outflow internal circulation
pump which allow a constant flow stirring. Flow: 8 litres per minute.
All baths with heating shelf and drainage. Faucet included.
Toatlcacpacity Must be 20 litres for water bath
Serial # 9
| Incubator |  |  | 1
Bacteriological incubators
TEMPERATURE THERMOSTAT CONTROL WITH ANALOGUE THERMOMETER.
ADJUSTABLE TEMPERATURES FROM AMBIENT +5 °C UP TO 80 °C.
STABILITY: ±0.1 °C, UP TO 37 °C. HOMOGENEITY: ±0.5 °C, UP TO 37 °C
INTERNAL GLASS Door.
SAFETY: OVER TEMPERATURE CUT OUT INCORPORATED ACCORDING TO THE
EN.61010 STANDARD. ADJUSTABLE SAFETY THERMOSTAT DIN 12880. FITTED.
internal shelves 3-5
capacity in litres Minimum 53 litres
Serial # 10
| Lab Oven |  |  | 1
Serial # 11
| Lab Sterilizer |  |  | 1
SELECTABLE TEMPERATURE BETWEEN 40 ºC & 121ºC.
CAPACITY: 20, 80 & 150 L.
WATER SUPPLY NEEDED WITH LOW-CAL CONTENT AND DRAINAGE. (WATER PRESURE
3 BAR MINIMUM). CONFORMITY EN 61010-1 SAFETY FOR EQUIPMENTS FOR LABORATORY
USE . EN 61010-2-040 SAFETY FOR STERILIZER EQUIPMENTS. EN 61326
ELECTROMAGNETIC COMPATIBILITY. EN/97/23/EC SAFETY FOR PRESSURE
EQUIPMENTS.
FEATURES External case, top, tank and lid made of stainless steel. Tank fill
by external water inlet. Temperature control probe inside the sample.
Counter pressure system to prevent the cans breakage. Temperature, time
and pressure registered in USB.
SAFETY Device that avoids opening the lid when the chamber is under pressure. Safety
valve: Prevents overpressure over the maximum limit. Safety pressure switch:
Disconnects heater when there’s overpressure. Thermally insulated lid. Door properly closed d
MicroProcessor includes:
Parameters LCD display. 2.
Microprocessor single button. 3.
Operation Led. 4.Start/stop
push button. 5. USB adapter. 6.
Open lid knob. 7. USB function
light indicator. C
MICROPROCESSOR FEATURES Microprocessor controls sterilization and cooling
process: Any error message of functions or execution is displayed in the panel and an
acoustic alarm sounds, automatically cutting the steam generator out. BACK VIEW 1.
Mains input.
| Serial # 12 |  |  | 3
Juster Pippete
● user friendly
● easy to adjust
● Must use universal tips
● Autoclavalble
●Maintenance free
● standard sizes, like 1-0,5-50,10-100 & 100-1000 Microlitres
| Serial # 13 |  |  | 1
Blood Mixer
Shaking monitor for blood collection FOR BAGS BLOOD COLLECTION PROCESSES.
SWINGING PLATFORM.
FEATURES HEDO2 is a digital blood collection monitor,
created to control and monitoring blood collection
processes. It has been designed to perform simple,
precise and accurate operations. It is composed of:
-Easily removable tray. -Wired or wireless inte
Automatic clamp closing at the end of the donation.
-Automatic tare calculation.
-Real time monitoring of the donation showed on the digital display
-Stand alone mode (operating with rechargeable battery). 8 hours life
-Tube sensor presence.
Power supply: 110-220V / 50-60 Hz
Must include the followin accesories:
ACCESSORIES Wireless data transfer
Wireless connection module Barcode
reader Alarm lamp pool
Must supplied with transport case
Serial # 14
| Resuscitation Trolley |  |  | 2
Serial # 15
| Laminar Down Flow System |  |  | 2
●Bearing structure in steel epoxy-polyester powder painted
●Work Surface AISI 304L "2B" finishing Tempered glass slide doors
●Lighting of the work surface by means of 2 fluorescent electronic lamps fitted in
non-contaminated areas
●Ventilation system by means of N1 or 2 Ip55 protection factor motor blowers
●Recirculating and air filtration achieved by non-returnable high retention capacity pre-filters
and HEPA H14 filters.
●Microprocessor control system
●Work Surface 120cm
●Work Surface180cm
Serial # 16
| Surround Light System |  |  | 2
Serial # 17
| Media Bridge 3-Sided |  |  | 2
Serial # 18
| Modular Surgeon Panel with X-Ray Viewing Box |  |  | 2
Serial # 19
| Scrub Station (3 Person) |  |  | 1
Serial # 20
| Angio Simulator |  |  | 1
Serial # 21
| Blood Bank Refregerator |  |  | 1
Technology in our blood.
A wide range of blood bank refrigerators to store up to 700 blood bags and meet small, medium
hospitals andand large blood centers, Analytical laboratories needs. Each product line ensures
the most secure environment for your valuable blood products. MINI and BBR are landmarks
in Blood Bank Refrigerators for reliability and lasting.
●Certified according to medical devices directive 93/42/EEC.
●Wide range of capacity (up to 700 blood bags)
●Temperature graph recorder (7 days)
●Forced air circulation
●High stability and uniformity of the internal temperature
●Stainless steel sliding drawers complete with stainless steel sliding guide which prevents rust.
●Fast temperature recovery after door opening
●Stainless steel internal chamber which grant maximum hygiene
●Triple vacuum-sealed glass door, with anti-moisture resistors
●Tropicalized execution (up to 40°C)
●Superior insulating panel
●Internal light (manual or automatic switch)
●Safety key lock
The microprocessor control system E2003
●Alarm section independent from regulation section
●Remote audio/visual alarm control
●Rechargeable buffer battery (48h in case of power outage)
●Advanced security
●User friendly control display
●Continuous monitoring
|  | SIGNATURE | CONTRACTOR NAM |  | DATE & SEAL