SPECIFICATION OF MAJOR EQUIPMENTS FOR

SINDH INSTITUTE OF ORAL HEALTH SCIENCES

JINNAHSINDHMEDICALUNIVERSITY, KARACHI

SPECIFICATION OF MAJOR EQUIPMENTS

| Item # | Description of Store | Quantity Required | unit Price (FOR,KAR) | UNIT PRICE (C&F, kAR) |
| --- | --- | --- | --- | --- |
| 01. | 3D DIGITAL DENTAL PANORAMIC AND CEPHLOMETRIC X-RAY UNIT (CBCT) AND VIRTUAL IMPLANT PLANNING SOFTWARE & ENT Module (complete system) LATEST VERSION MODEL: MAKE: COUNTRY OF ORIGIN: FDA/ CE ETC CERTIFICATION. SPECIFICATIONS: - Digital Sensor for Panoramic, Cephalometric & 3D (CBCT) imaging. - Programs: - Panoramic Program - Segmented Panoramic Program - Maxillary Sinus Program - LA TMJX2 Program - LA TMXJ4 Program - Lateral Ceph Program - Lateral Ceph Program with PA or AP View - Oblique Ceph Program - Submentovertex Program - Carpus View Program - Cephalometric Image Size Program- Option of 30x30, 24x30, 24x24, 18x24 & 18x18 cm Images - 3D Program for Dental & ENT FOV- 5x5, 8x8, 10x5, 10x10, 17x6, 17x11, 17x13.5 TECHNICAL DATA: - 3D Scanning mode- Continuous & Pulse - 3D Scanning time- 15 to 30 sec. - Voxel size: 95 to 550µM - Pan exposure Time: 04 to 16 sec. - Ceph Exposure Time: 0.1 sec. to 3.2 sec. - One Shot Cephalometric Sensor to reduce patinet blur - Patinet position- Standing/ Sitting/ Wheel chair - Tube Voltage: 60 – 90KV - Tube Current: 2 – 15mA - Frequency: 130 – 150KHz - Tube Focal Spot: 0.7mm - Input Voltage: 220/230/240V, 50/60Hz SOFTWARE: - DICOM Compatible Imaging software - Image Acquisition software with UNLIMITED LICENSES & INSTALLATIONS on unlimited computers - Implant Planning Software with installations on UNLIMITED COMPUTERS - Auto Ceph Analysis Software with Tweed, McNamara, Stiener & Riketts Analysis without Annual Subscription Cost ACCESSORIES: - Compatible Branded PC with manufacturer’s recommended specs (Documentary proof will be required) & HD LED monitor with Keyboard & Mouse. - Pure Sine Wave UPS to take load of X-Ray, PC & Printer. - Compatible DICOM Printer (Agfa/Fuji/Kodak). - 08 x Computers with LCD. - 01 x Sever Core i7 Hard drive 1TB+ 9X - Client Core i3 Hard Drive 500GB - 18.5” LCD Monitors with wireless Keyboard & Mouse. - With Complete Wireless networking. - Digital Printer Dry Laser Printer Compatible with Digital OPG Machine; - Sizes- OPG, Cephalometric & CBCT - X-Ray Sheet & transparent sheets. - Power: 220V 50/60hz - All necessary software with networking for printer shall be provided by supplier. computer with lcd (Details) (08 qunitity) - 1x Server Core i7 Hard Drive 1TB+ 9x. - Client Core i3 Hard Drive 500Gb. - 18.5" LCD Monitors with wireless Keyboard & mouse - With Complete wireless networking. Digital Printer Dry Laser Printer Compatible with Digital OPG Machine. - Supreme quality printing, High speed, Remarkable volume, Easy usage, Easy maintenance, Excellent performance - Sizes: OPG and Cephlometry, CBCT - X-Ray Sheets / Transparent Sheets. - Power : 220V 50/60Htz - All necessary software with networking for printer provided by the supplier. - Five year full warranty for all CBCT related Equipments. International Training from the country of Origin must be Provided by the Supplier.(Three Person) Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | _________ _________ _________ _________ 01 NO. 08 nOS. 01 no. |  |  |
| 02. | CENTRAL STERILE SUPPLIES DEPARTMENT (CSSD), COMPLETE PACKAGE INCLUDING CLEANING & WASHING UNIT, WASHING/ DISINFECTING DRYING UNIT, AUTOCLAVE, RO SYSTEM, OIL FREE COMPRESSOR, ULTRASONIC CLEANER, TRANSPORT AND DISTRIBUTION TROLLEY, PAPER DISPENSING TROLLEY, PAPER SEALING MACHINE, TAPE DISPENSORS, FREE STANDING BASKETS, PACKING TABLE, GLASS BEAT STERILIZER AND DRY HEAT STERILIZER.. CLEANING & WASHING UNIT 01 NO. USA/ EU/ Japan/ Equivalent/ Indigenous. - Cleaning & Washing Unit approx. 2000 x 700 x 900mm (WxDxH) consisting of Stainless Steel square meter, Table Top with 50mm backsplash made of Ni-chrome steel. - Two (02) Sinks each 500/500/250mm made of stainless steel with the provisions of counter top; equipped with drain valve and stand pipe and with siphon trap. - Substructure cabinets with shelve. Cabinet in ni-chrome steel. - Two (02)integrated Spray Guns for dematerialized water / air with one set of 08 cleaning nozzles hose table lead through (one behind each sink) - One (01) mixing tab 3/4" in middle of the both sinks WASHING/DISINFECTION / DRYING UNIT 01 NO. USA/ EU/ Japan/ Equivalent - High capacity washer disinfector with capacity up to 10-DIN net baskets. - Double door pass-through machine with automatic sliding glass doors. - Two big capacity pre-heating tanks directly connected through large valves, into the washing chamber, grant effective results in reducing cycle time. - Water connection: cold, warm, demineralised - Following EN ISO 15883-1/-2 rules - Washing chamber entirely made in stainless steel AISI 316L, corrosion proofing - Construction system with total removal of angle and breeding ground dirty self cleaning tank with rounded edge. - Two pre heating water tanks. - Water filtering system in three levels for an excellent Washing result - Automatic motorized double doors with vertical down sliding and safety device. - Door made in HST tempered glass - Equipped with 02 dosing pumps of chemical products, under a complete micro-processor control, with the possibility to modify the required products quantity in each cycle. - Flow meters for checking and validation of chemical products quantity. - check of chemical products quantity, with minimum level alarm. ACCESSORIES - A.N Wash Cart 01 no. - 4/5 level washing cart 01 no. - Instrument trays 10 nos. - Wash cart for dental instruments 01 no. - Transfer trolleys for cart 02 nos. HEAVY DUTY STEAM AUTOCLAVE AUTOMATIC DOUBLE SIDED (2 DOORS VERSION) 01 NO. USA/ EU/ Japan/ Equivalent - The sterilizer should be designed according to DIN EN 285 and bears the CE making in compliance with MDD. - Steam heated electrically operated with microprocessor control system of latest version, Chambers Capacity 200 lit. - Two door model Built in electrical steam generator. - TECHNICAL SPECS: - A) Programs: Different programs for time, temperature, pressure, cycles etc. - Program and display menu shall provide all important cycle information and message of the safety check system in case of deviations from preset values. - The service menu comprises the display menu, control menu and intervention menu for the service engineer. - B) Safety Devices: According to DIN EN 285, - C) Operating Control: Integrated operating Panel, attached to the front of the sterilizer. Panel with colored touch - screen, signal lamps, switch "Open Door / Close Door", Program-stop-button. The signal lamps are represented on the touch-screen. - Door System: Double Door System, Motorized Pass Through Type. - Sterilization Chamber: Double-Walled; interior walls and steam jacket made of Ni-chrome-molybdenum-titanium steel 316. - Highly efficient heat insulation with solid insulation. - Vacuum Device: Powerful, low - noise water ring vacuum pump with counter-flow heat exchanger with water saving system controlled by microprocessor. - Built in steam generator entirely made of high quality stainless steel AISI 316L - Chassis made of high quality steel. - Frame and panels: AISI 304 - Mechanized door gasket groove for a longer gasket life and easier maintenance. - Three phase voltage 400 V/50Hz - Printer - Programs: - 134°C Universal-program, total process time for 20 kg load: 20 minutes (5.5 minutes killing time) plus 20 minutes drying time (abort by user) - 134°C Quick B –program, total process time for 5 kg wrapped load: 17 minutes (3.5 minutes killing time) plus 8 minutes drying (abort by user) - 121°C Gentle-program, total process time for 10 kg porous load: 40 minutes (20.5 minutes killing time) plus 20 minutes drying time. - 134°C Prion-program, total process time for 20 kg load: 40 minutes (20.5 minutes killing time), plus 20 minutes drying time - Vacuum-test program. - Bowie & Dick-test program - Air leakage detection program REVERSE OSMOSIS SYSTEM (TO FULFIL WATER REQUIREMENTS OF CSSD AND STEAM GENERATOR) 01 NO. USA/ EU/ Japan/ Equivalent/ Indigenous. - Water softening RO unit for steam Sterilizers washer disinfector and cleaning / washing units. - The RO system including Filter housing, Hardness stabilizing unit, Reverse-osmosis-system suitable to the offered equipment with Storage tank, Level control for tank, Pressure pump, Conductivity meter, Distribution manifold, Hose set and installation material imported component with local fabrication. OIL FREE AIR COMPRESSOR 01 NO. USA/ EU/ Japan/ Equivalent/ Indigenous. - Compressor oil free, Noiseless, single stage, Dual Head cylinders with electric motor 1.2 KW – 2.2 KW, pressure 6-8 bar, capacity 140 liter/min @ 5bar with 50 liter tank. Pressure Gauge, Filter With Regulator, Safety Valve, Drain Valve, ULTRASONIC CLEANER 02 NOS. USA/ EU/ Japan/ Equivalent - Ultrasonic Cleaner integrated type, cleaning chamber constructed of corrosive resistant AISI 304. - Heating thermostatically controlled. - Digital display of time and temperature. - Tank Capacity 28 Ltr. - System complete with basket, lid and other standard accessories. - Drain Valve for Quick drain the waste water. - Operation 220V, 50 Hz. TRANSPORT AND DISTRIBUTION TROLLEY 05 NOS. USA/ EU/ Japan/ Equivalent/ Indigenous. - Distribution trolley made of 3 shelves stainless steel, Dim. 700 x 500x 895mm PAPER DISPENSING TROLLEY 02 NOS. USA/ EU/ Japan/ Equivalent/ Indigenous. - With floor stand, movable 1410 x 725 x 995 mm - Movable paper dispenser trolley made of stainless steel for storing of sterilization paper sheets at packing tables. 04 frames for paper sheet of 1200mm and 1 bottom shelve. Mobile with two lockable wheels. PAPER SEALING MACHINE 02 NOS. USA/ EU/ Japan/ Equivalent - Complete stainless steel body - Sealing speed 10 m / min - Sealing temperature 0 - 200 C - Sealing edge Sealed seam width: 12 mm - Main connection: 230 V/50Hz - Microprocessor controlled automatic heat sealer for sterilization bags and pouches. - Stainless steel body with printing mechanism. CUTTING DEVICE 02 NOS. USA/ EU/ Japan/ Equivalent - For storage and preparation of paper/Plastic bags in rolls. The cutting knife is made of tempered stainless steel and is self grinding. Size approx. 873x125x420mm. TAPE DISPENSER 10 NOS. USA/ EU/ Japan/ Equivalent/ Indigenous. - Tape dispenser for autoclave tape. Main parts of the dispenser are made of metal. Design for two roll of tape. FREE STANDING BASKETS RACK USA/ EU/ Japan/ Equivalent/ Indigenous. PACKING TABLE 02 NOS. USA/ EU/ Japan/ Equivalent/ Indigenous. - Table top of Stainless Steel with one drawer and power sockets (692/376/198). Stand of stainless steel including two top mounted shelves and Illuminated Magnifier*. Dimensions approx. 1600 x 800 x 800mm. GLASS BEADS STERILIZER 05 NOS. USA/ EU/ Japan/ Equivalent - Automatic thermostat, safety thermic fuse, double thermic fibre glass insulation, careful electric insulation, feeding cable with safety plug, insulating material externally protected by stainless steel plate, armored resistance, air circulation on 5 sides, switch and warning lights with safety marks. One removable nickel copper cup of easy replacement with holding edge to avoid leakages of quartz bal. Three gauges (network, wait, work) to control the condition of the equipment. Sterilizing times: Files, probes: 3 seconds, Mirrors, surgical instruments, curettes, forceps, milling cutters, hand pieces in 5-10 seconds. Technical details: Sizes: 11x10x15cm, Basket Diameter 5.5x7.5cm Max absorbed Power: 60 Watt. Power: 220V/50 Hz HOT AIR STERILIZER 06 NOS. USA/ EU/ Japan/ Equivalent. - Inside Chamber is Double walled made of Anodized Aluminum/ stainless steel. Outer chamber made of mild steel & powder coated. Chamber Volume 27 Liter with 2 trays. Thermostatically controlled form 50°C to 250°C (± 1°C). Beaded heating elements are placed in the ribs at the bottom and sides. Complete with L-Shape thermometer, 02 indicators, thermostat and on/off switch. To operate in 220/230 volts. International Training from the Manufacturer must be provided by the Supplier to the Three Persons of SIOHS. Complete Job: Supply, Installation, Testing, Commissioning, Maintenance & Operational/ Application training to the designated staff of SIOHS during maintenance Period of 06 months or as appropriate. Five (05) Years full Warranty for all equipment. UK/USA/JAPAN/GERMANY OR EQUIVALENT | 01 no. 01 no. 01 nO. 01 no. 01 no. 02 no. 05 no. 02 no. 02 no. 02 no. 10 no. 02 no. 02 no. |  |  |
| 03. | CAD CAM SYSTEM FOR DIGITAL DENTAL LABORATORY COMPLETE WITH SCANNING DEVICE, MILLING MACHINE, SINTERING FURNACE & NECESSARY HARDWARE SOFTWARE. MODEL: MAKE: COUNTRY OF ORIGIN: FDA/ CE ETC CERTIFICATION. SPECIFICATIONS/ TECHNICAL DATA 3D DENTAL SCANNER: - Latest Generation 3D Scanner with 10-micron accuracy. - Fast fully automatic scanner. - White-light scans providing fully open STL scans. - Fully automatic 3D calibration. - Capable of transferring Jaw joint position into virtual articulator. - Scan Fixators for Articulators. - Large measurement field. - Suitable for longer directly screwed implant bars. - Output in open STL file format. - Fast 2 or 5 axis system capable of providing accurate cranio-facial information from physical articulators. - System shall be capable of transferring the scan models, jaw joint position and articulator setting from the physical articulator to the CAD software. - Accuracy: 6µM - Implant Bar Accuracy: 8µM - Measurement Field (XxYxZ): 80x60x85mm - Complete Jaw measurement time: 90 sec. - Single Tooth measurement time: 55 sec. - Measurement time 3-unit bridge: 90 sec. - USB Interface. - Power 100-240VAC, 50/60Hz - Standard Accessories; - Connector cable set, - Software CD with installation program, - Operational manual. - Dell Workstation. - Lifetime license for CAD - Free of cost all software updates SOFTWARE: - Capable of constructing up to 32 crowns and bridges. - Off-set, reduced-anatomic of fully-anatomic module. - Simultaneous construction of upper and lower jaws. - Fully automatic recognition of preparation lines. - Results in open STL format. SPECIFICATIONS/ TECHNICAL DATA CAD CAM: - Five (05) simultaneously working axis. - Automatic axis compensation for exact results. - Capable of Dry machining & Wet machining. - Plastics, Wax, Zirconium oxide, Composites, CoCr, Glass ceramics. - A-Axis Rotary axis for processing work pieces on their full circumference of 360°. Harmonic drive free from backlash. - B-Axis Rotary axis rotation range 30° for undercuts. Harmonic drive free from backlash. - Powerful & highly precise synchronous spindle and an automatic tool changer for 16 tools. - Automatic safety interlock. - Under pressure sensor for monitoring the air extraction. - Power 100-240VAC, 50/60Hz - Multiple block loading in milling machine - Free of cost all software updates SPACIFICATIONS/ TECHNICAL DATA OF SINTERING FURNACE: - Fully automatic high performance furnace for final sintering of distortion free zirconia framework. - Four (04) sintering program locations, for different restoration sizes. - High temperature furnace with constant temperature control. - Even temperature distribution in the firing chamber. - Notification of termination of the sintering program due to power failure etc. - Frictionless sintering process. - Large firing chamber with Two (02) stackable sintering bowls for maximum utilization of the furnace. - Power: 3.5kW, 220-240V, 50/60Hz Five (05) Years full Warranty for all equipment. International Training from the Manufacturer must be provided by the Supplier to the Three Persons of SIOHS. Complete Job: Supply, Installation, Testing, Commissioning, Maintenance & Operational/ Application training to the designated staff of SIOHS during maintenance Period and production along with the materials of 06 months or as appropriate. Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 01.NO |  |  |
| 04. | oil free compressor with dryer - For 04 Executive clinics. - One Motors Single phase 1700rpm. (UK, USA, JAPAN, GERMANY) - Output power: 3~ 50Hz-2.2 kW 12A - Dust Free, Moisture Free, - 3 Hp Motor - Pumping capacity 1500 l/min - Mounting on Tank - Tank Capacity : 100 L - Noise level: 81dB(A) - Automatic Dryers with Drain system - Automatic Thermal Protection Switch - Automatic ON/OFF Switch - Air Gauge, Air Filter with Regulator. - Automatic safety release valve - Digital Display for Power supply. - Automatic Circuit Breaker. All the plumbing and pre installation work done by the supplier Training must be Provided by The Supplier wherever it required Complete with standard accessories # UK/USA/JAPAN/GERMANY OR EQUIVALENT | 01.NO |  |  |
| 05. | SUCTION Machine for 04 Surgeries - High power Motorized suction System Suitable for 04 Dental Units. - Individual separator for 04 unit - Mechanical and Electrical Air bleed Opening System, Controlled with P.L.C (programmable Logic Controller.) - single phase motors type (UK, USA, JAPAN, GERMANY) - Output power: 2.2 / 3.5 kW-12A 50/60Hz 3600/rpm each motor - Control Panel board with VSD Inverter, safety Switch - Digital Display for Power supply. - Automatic Drainage system - Max. vacuum level: 240mbar (continuous service) - Max. flow: 1000 l/min - Noise level: 77dB(A) - Electronic /Electric Panel With Electronic Level Sensors and Automatic Washing System. All the plumbing and pre installation work done by the supplier Training must be Provided by The Supplier wherever it required Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 01 No. |  |  |
| 06. | HIGH SPEED AIR TURBINE-(02 hole) - Ultra push Chuck, Ceramic Ball Bearing - Maximum Speed 300,000 rpm - 400,000 rpm - Autoclave-able 135C - Clean Head System - Four Hole Tubing System Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 25 Nos. |  |  |
| 07. | SLOW SPEED AIR MOTOR Complete set (02 hole) - Slow Speed Hand Piece Standard Set for Four Hole Tubing System. - Maximum Speed 25000 rpm with Straight Hand pieces 1:1 Direct Drive. Max Speed 20000 - 40000 rpm For HP burs (Ø2.35) - Contra Angle Hand piece 1:1 Direct Drive, For CA burs (Ø2.35) Push button chuck. - Max Speed 30000 – 40000 rpm - External water spray Having facility to rotate the bur clockwise and anti-clockwise) - Autoclave-able 135C Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 20 Nos. |  |  |
| 08. | light curing MACHINE - Rechargeable pen type curing light - Easy to use with one touch button - General curing mode for 10sec, slow rise mode for 30 sec and Pulse mode. - Non light guide tip - Li-ion battery without memory effect - Output intensity 1200-1500 Mw/cm2 - Wavelength 450-480 nm - Rechargeable battery. - Disinfect able Detachable Curing Probe Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 05 Nos. |  |  |
| 09. | Dental scaler (PEizon technology) - Table Top Scalar - Feedback system for built in circuit module keeps constant strong power. - Regulator controls output power moderately in different quantity of tartar. - Circuit module with digital system works more accurately. Standard Accessories - 01 Detachable Hand Piece (Autoclave able). - 05 Tips with individual tip holder.(Different Shape) - 01 Torque wrench (Autoclave able). - 01 Autoclave able box. Technical Specification - Working frequency 25-35kHz - Power supply AC 220V, 50/60Hz - Water supply pressure1-5bar - Adapter 24VDC~#30VDC - Foot Switch - Digital Buttons for Power Control Shown by LED - Water Valve for Water Pressure Control with Built-in Filter. Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 10 Nos. |  |  |
| 10. | AMALGAMATOR - Mechanical Dental Amalgamator - Automatic amalgamation with timer is a compact and functionally designed mixer which assures efficient and reliable triturating with a simple operation - Excellent durability and mixing power with 4 capsules - Digital Display with time setting. - Safety casing with limit switch. - Power 220V 50/60 Hz Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 05 Nos. |  |  |
| 11. | PORTABLE DENTAL UNIT: - Mobile trolley with briefcase. - Complete Standard set. - Two Pcs Four Hole Tubing for Hand Pieces - Triple Syringe - Built-In Type LED Type Scalar with Deatechabl Scalar Hand piece and Five Autoclavable Scalar Tips - Air operated Suction with Waste Container - Air Filter with regulator - Built-In Pressure Bottle - Built-in Compressor 550 watts - 220 Volts 50/60 Htz Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 03 Nos. |  |  |
| 12. | PIEZO ULTRASONIC & IMPLANT MOTOR Includes 20:1 Hand piece + Piezo Hand piece with 10 Tip's ultrasonic and rotary technologies to be used in diverse clinical areas, especially for dental implantology and Piezo Surgery with dual functions. The ultrasonic hand piece is active only on hard tissues, and it is non-invasive on soft tissues, resulting in improved safety, efficiency and less pain during surgery. In addition, when the dentist takes the hand piece off from the patient, the unit moves to a standby mode, which provides durability on the unit and hand piece as well as increasing the tip’s life span. For implant motors, it provides all necessary and efficient functions for dental implant surgery, such as highly qualified torques for drilling and insertion of dental implant fixture, and visibility with LED in hand piece. Power Supply 220 V Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 01 No. |  |  |
| 13. | PROPHYLAXIS SYSTEM Mode of Operation: Continuous operation with intermittent loading: 10 minutes ON / 20 minutes OFF Electrical Voltage 220V Gen 122J 115V Gen 122 Current 1.1/1.1 amps @ 100v/115v Frequency 50/60 Hertz Water Pressure 25–60 psig (172–414 kPa) Air Pressure 65–100 psig (448–690 kPa) Flow Rate Minimum setting (CCW) 0–15 ml/min Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 05 Nos. |  |  |
| 14. | GUTTA PERCHA CUTTER ELECTRICALLY DRIVEN - Press rapid heating release the button a few seconds for cooling - Using wireless charging technology stylish fashionable. - 2400mAh large capacity nickel cadmium rechargeable, battery avoid frequently charging trouble. - Fever needle can be sterilized, it can be repeatedly used. - Battery: 1.2v 2400mAh - Charger input voltage: 220-240V, 50-60Hz. - Power: 2.5w Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 05 Nos. |  |  |
| 15. | WIRELESS OBTURATION DEVICE: - Cordless and compact design - Fast and continuous flow of filling tightly seals the canals - Variable temperature setting (100°C ~200°C) - Quick time for heating - Ergonomic design 360° swill needle offers - Thinner needle can easily the narrow root canal Accessories: - Obturation Gun - Charger input 220 - Safety cup - Multi purpose Wrench - Scubbing Brush - Battery - ObturationGuttapercha medium type 100 each Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 02 Nos. |  |  |
| 16. | PORTABLE dental x-rays - Convenience of cordless: Use the charging battery-pack of loading/unloading type - Incredible reduction of X-ray exposure: Up to 1/7 of existing X-Ray systems - More compact size and weight than existing portable dental X-ray (2Kg) - Adopting high-frequency circuit of inverter type - Simultaneous using of digital sensor and film is possible - User control time set: Time set-up of exposure on the location of teeth - More than 300 radiography is possible with one time charging of battery - Image checking in real-time by built-in TFT LCD - Acquisition of X-ray images of high quality using built-in diverse image processing technology X-Rays Generator : High frequency Tube voltage : 60 KV (Fixed) Tube Current : 2 mA (Fixed) Tube Type : Stationery Tube Focal Spot : 0.8mm Target angle : 20◦ Weight : 1.8 Kg Approx Detector : Film / Digital Sensor Exposure Time : 0.05 ~ 1.25 Sec – Digital Sensor 0.20 ~ 1.35 Sec – Film Exposure Power Requirement : DC 14.2 ~ 16.8 V Power Supply : Rechargeable Battery Must Provide one Extra Battery for each Unit Training must be Provided by The Supplier wherever it required Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 02 Nos. |  |  |
| 17. | DIGITAL IMAGING PSP SENSOR Phophorus Plate: - working upto 08 surgeries with software installation - Lightning fast: image readout completed in seconds. - Repeatable image quality automatically, without special training. - Advanced image processing technology. - Minimized need to adjust images. - Virtually no under - or overexposures. - No need to change exposure settings for each patient. - One step handling, fully daylight operation. - True cordless best for multi surgeries. - Uniquely hygienic for improved patient and user safety. - Touch less operation, protective covers, liquid proof bags. - Patient comfort with new latex free, food grade hygiene bags. - New generation of durable, cordless, thin and flexible imaging plates. - Re-usable plates, sizes equal to intra oral films (0, 1, 2, & 4C) Readable. - Imaging plates are as easy to position into patient’s mouth as film. - Small footprint with many installation options. - Compatible with all intra oral x-ray unit. Technical Data: - Technology Imaging plate PSP = Photo Stimulable Phosphorus plate - Digital LCD for quick View Imaging at the front - Readout time 5-15 seconds - Standard network connection(Wifi / Ethernet RJ45)/ USB - Operating Voltage 100-240VAC, 50/60Hz - Operating SystemWindows Vista or Window 7 Imaging plates set - Size#0x1, Size#1 x5, Size#2x2, Size#3x1, Size#4Cx1set Protective Covers & Hygienic Bags for all sizes. User’s software licenses software for multi Surgeries. computer with lcd (01 qunitity) - 1x Server Core i7 Hard Drive 1TB+ 9x. - Client Core i7 Hard Drive 500Gb. - 18.5" LCD Monitors with Wireless Keyboard &Wireless Mouse - With Complete wireless networking Training must be Provided by The Supplier wherever it required Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 01 no |  |  |
| 18. | AUTO CLAVE - Auto Clave Class-B 22-23 LITTER - Five Stainless Steel Trays. - Complete Standard Set Special Features - Provides full traceability of every single sterilization cycle. - Sterilization cycle for Solid, Hollow, Pours Instruments, Cloths, Liquid also. - Three Pre Vacuum and One Post vacuum cycle. - The Anti dusts filter traps impurities in the cooling air and protects the sterilizer interior and steam condensation system. - Steam generator and double head vacuum pump. - The Electro polish stainless steel sterilization chamber. - Double Door locking system (electrical and Pneumatic) - Door locking mechanism which ensures the load can be recovered in the event of a power failure. - The quality of the distilled water being used is guaranteed by the conductivity sensor incorporated in the tank and an efficient warning message/alarm system. If the water is of unsuitable quality, the sterilizer informs the user with appropriate messages prevents the use to avoid damage risks. - Leakage test, Cleaning Cycle and B&D Test - The quick coupling system for water filling and draining on the front section of the sterilizer simplifies and speeds up to relevant tasks. And that’s not all: it is also possible to equip the sterilizer with an automatic filling system that withdraws water from an external tank. - Thermal Printer for check the step by step process and recording for sterilization cycle. Technical Data - Chamber capacity20-25 Liter - loading capacity05 kg - electro – polished stain less steel chamber (316 ti) - platinum temperature sensorpt 1000 - internal storageof last40 cycle - Power Supply220/240V - Display LCD Display - Main Frequency50/60Hz - Rated Power2300W - Tank Capacity4.5liters - Waste Water Tank Approx 3.5 Litres - Fully compliance according to EN 13060 Training must be Provided by The Supplier wherever it required Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 04 Nos. |  |  |
| 19. | AUTOMATIC WASHER & DISINFECTION SYSTEM WITH FORCE DRYING - Forced drying system, with warm air generator. - Eco Slim Steam Condenser, 2back supply connection for detergent& neutralizer loading. - Detergent/Neutralizer dosing system with peristaltic pumps - 15 to 20 washing/disinfection programs upto 95C, - 6 custom programs - Integrated water softener with automatic volumetric re-generation 60F - Automatic electro-mechanical door locking system. - Washer chamber dimension (WxHXD) 380x590x480 - 3300W. - Digital LCD Display - 220 V - Capacity 22-30 liters Training must be Provided by The Supplier wherever it required Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 01 no. |  |  |
| 20. | POUCH SEALER - Electronic sealing control. Digital display with soft touch keypad With Temperature Controller. - Automatic bag preparation, with pre-setting of bags quantity & length with Cutting system - 12mm sealing band. - Allow the sealing of pouches up to 300mm - Automatic sealing of pouches up to 200mm - Integrated roll-holder. Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 04 Nos |  |  |
| 21. | Water distiller - Independent water distiller, complete with container - Fully automatic process to produce 4Ltrs within 6Hrs - Digital Display - Safety Thermal Protector. - Auto Off when Process Completed. - Automatic Builton Cooling Fan. - Water Filter with Carbon Filter. - Stainless Steel Chamber, Rust proof. Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 04 nos. |  |  |
| 22. | ULTRASONIC CLEANER - Heating Power: 400W, Capacity 5.7Ltrs, Ultrasonic frequency 38KHz - With Basket & Glasses holder complete with 2PVC glasses - Digital display - One Big Grill for Instrument and one for Small Grill for Tips and Burrs. - Drain Valve for Cleaning. - Automatic Heat Temperature Detector Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 04 nos. |  |  |
| 23. | - DRY HEAT STERIZER - Capacity: 22 liter - Voltage : 230 V 50/60 Htz - Absorbed Thermostat 50°C ~ 200°C - Sterilization Temperature 150/180°C - Adapted Timer 0~120 Min Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 03 nos. |  |  |
| 24. | PH METER - Resolution : 0.01 Ph - TemperatureRange : 0.0 to 100ºC (Manual Compensation) - Display : 31/2 Digit LED Display - Power Supply : 230VAC ±10%, 50 Hz - Calibration check facility & Calibration Error indicationfor 7.00 & 4.00 pH - Calibration check facility & Calibration Error indicationfor 7.00 & 4.00 pH. Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 01 no. |  |  |
| 25. | INFRARED THERMOMETER General Specifications Display 3½ digit backlit LCD display with function indicators Display rate 1 second approx. Over range indication Audible tone sounds and ‘OVER’ appears on the LCD when the measurement exceeds the measurement range spec. Operating Temperature 0°C to 50°C (32°F to 122°F) Operating Humidity Max. 80% RH. Power Supply 9V battery Automatic Power Off Meter shuts off automatically after 7 seconds Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 05 nos. |  |  |
| 26. | MEDICAL OXYGEN CYLINDER THERAPY UNIT PORTABLE AND TROLLEY MOUNTED Technical Data: Flow Rate: (0~3) LPM Oxygen Purity : 65% @ 3 L/min; ≥90% @ 2 L/min Outlet Pressure : 0.04 - 0.06 MPa Noise : <45dB (A) Applied Voltage : AC 220V / 50Hz ( or customized) Power Consumption : 150 W Size : 440x340x290mm Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 02 nos. |  |  |
| 27. | FIRST AID BOX Technical Data: Minimum quantities for low risk establishments and activities may be considered as a general guidance leaflet on first aid. 20 individually wrapped sterile adhesive dressings (assorted sizes) appropriate for the activity (detectable dressings (colored blue) should be available if catering is to be undertaken) 2 sterile eye pads 4 individually wrapped triangular bandages (preferably sterile) 6 safety pins 6 medium sized individually wrapped sterile unmediated wound dressings (approx. 12cm × 12cm) 2 large sterile individually wrapped unmediated wound dressings (approx. 18cm × 18cm) 1 pair of disposable gloves. Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 05 nos. |  |  |
| 28. | WHEEL CHAIR Technical Data: Intended User adults without postural support needs Drive train storable lever drive train and push rims Gear Ratio variable mechanical advantage from 1:1 to 3:1 Frame rigid mild-steel frame with three wheels Axles adjustable rear axle position Rear Wheels 26 in. (66 cm) pneumatic bicycle tires Front Wheel 8 in. (20.5 cm) solid rubber front castor Footrest adjustable footrest with straps Armrests fixed armrest with mudguard Support comfort cushion, seatbelt, chest and foot straps Bearings bicycle ball bearings on all bearing surfaces Reparability repairable in a local bicycle shop Features Technical Specifications Seat Depth 14 in. (35.5 cm) Length 44 in. (112 cm) Sizes Small Standard Large Width 24 in. (61 cm) 26 in. (66 cm) 29 in. (74 cm) Seat Width 13 in. (33 cm) 15 in. (38 cm) 18 in. (46 cm) Weight 49 lbs. (22 kg) 51 lbs. (23 kg) 53 lbs. (24 kg) Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 05 nos. |  |  |
| 29. | - INTRA ORAL CAMERA with VGA arm and LCD - Power 5V ± 1V, 1500MA - Image Sensor ¼ Sony HAD CCD - Focus Range – 3 -40 mm - Out VGA - Image resolution – 1024 x 768 - Handle LED 200 mm - Illumination 6 LED - Adaptor 5V DC - VGA Cord - 18.5 inches LCD/LED - VGA arm for Holding the LCD in Dental Units Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 05 nos |  |  |
| 30. | - TRANSPORT STRETCHER Technical data Maximum weight limit: 700 lbs (317.5 kg) Overall length: 83" (210.8 cm) Overall width: 32" (81.3 cm) or 36" (91.4 cm) Siderail length: 47" (119.4 cm) Siderail height above deck: 14.5" (36.8 cm) Mattress size: 26" (66 cm) x 75" (190.5 cm) or 30" (76.2 cm) x 75" (190.5 cm) Maximum head elevation: 90° Maximum Trendeleburg/Reverse Trendelenburg: 18° Floor to base clearance: 3.5" (8.9 cm) (except under the hi-lo cylinders) Caster size: 8" (20.3 cm) Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 05 nos. |  |  |
| 31. | PEDIATRIC DENTAL UNIT - Three standard peado dental unit with Four hole Tubing system. - All peadsEquipment and Instruments. - Multimedia Projector with Portable Screen. - Complete Computer System. - Audio Video Sound system with Gaming Zone. - Matching color furniture with department for pediatric. - Fully Computerized control motorized, which gives smooth and non-jerky start and stop. - Lowest height range for children seating with safety belts. - Comfortable arm rests. Right arm swings for patient use - Hand piece console easy to wipe and clean. - Soft touch Control Panel for chair positions, - Interruptible time controlled automatic cup filler (adjustable). - Interruptible time controlled automatic Bowl flush.(adjustable) - Dental light low and high intensity, halogen lamp 12 V 50 Watts. - Spare Holder for built-in curing light. - All Internal tubing is durable & long lasting. - Rectangular air pressure gauge’s - Water flow control. - Individual air pressure controls. - One 3 way syringe for doctor with quick disconnect tip for autoclave. - 3-Aseptic easy clean, tubing’s and fitting for the hand piece straight 04 hole tubing. - Multi function foot operated switch and Chip Air Switch. - Built-In Bottle for Hand Piece Exhaust Drain - Double articulating Head rest adjustable for neck positioning & pediatric use. - Emergency stoppage by single touch of switch. - Foot control for chair Up/Down and Back/forward movement. - Two operator stool. - Smooth seamless upholstery in a large selection of colors. - All the outlet &inlet for the services to the chair should be concealed in the box to be at the foot area of the chair or within the unit, as an infection control measure. - Should be with minimum 5 delivery ports for various hand pieces, 3-waysyringe and extra - One piece spittoon with the Cartoon face. - Strainer Basket for Bowl. - High volume evacuator (air operated). - Saliva Ejector (Air operated). - Air Filter & Regulator, Water Filter, Air & Water supply lines. - Safe electrical compartment with Transformer, fuse, main Switch etc. to give output. 12/24 VAC for safety. - Main 220/240V. - They come standard with detachable lugs, two closure plugs and ground screws, making installation fast and easy. STANDARD LAYOUT: • Dinosaur chair 1pc • 4holes high-speed hand piece tubing 2pcs • 4holes low speed hand piece tubing 1pc • Operation light 1pc • Fish instrument tray 1pc • Three way syringe 1pc • Automatic heating water supply 1set • Centre suction system 1set • Cuspidor 1pc • Big X-film viewer 1set • Foot control 1pc INSTRUMENTS FOR PEDO DEPERTMENT: • Multi color examination set (probe, mirror, twezzer) 10 sets • Filling instruments set for pedo department 03 sets • Elevators sets consists of 5 elevators 03 sets • Forceps sets consist of 5 different forceps 03 sets MULTIMEDIA PROJECTOR High resulation multimedia projector with portable Screen 12’x10’ • Computer i-3 with 80 GB hard drive • DVD Rom • Wireless Mouse • Wireless Key board • Built-in games, cartoons, videos, Tooth brushing Techniques and other knowledge able information for Childs. • Multimedia projector VGA cable • 12’x 10’portable screen • Multimedia speakers with woofer FURNITURE FOR PEDO DEPERTMENT: FURNITURE FOR PEDO DEPERTMENT (1) SET - Table top cabinet with the matching color of pedo unit standard sizes dryer’s size as per department. - One toys shelves for pedo department with 20 small toys. - Color matching for walls of Pedo Department size as per department. Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 03 Nos. |  |  |
| 32. | DIGITAL SLR CAMERA WITH RING FLASH - Most convenient, versatile, easy to use camera for dental applications is the digital single lens reflex (DSLR) camera. 18-75 mm Lens - The body Additional lens 75-300 macro lens The flash TTL ring flashes. - Resolution 24 megapixel CMOS - Scene intelligent Auto - 4.3” LCD screen - focal length of 85-105mm - 9 AF point Flash - ISO: 6400 - Aperture: f/8 for full face - f/32 for intraoral - Shutter Speed: 3 Frame Per Seconds - Big memory cards, i.e. 16GB, - Big hard drives, both internal and external, - Compatible Ring Flash with Charging Device - Photography Digic 4software - image previews, or the equivalent of thumbnails, when you are viewing large numbers of pictures, instead of the full image, to avoid bogging down the computer, and photography software allows easy viewing of RAW files,just as with JPEGs. Tripod with ball head. | 02 Nos. |  |  |
| 33. | BIO-MEDICAL ENGINEERING TOOL KIT: Sensitive and expensive medical and laboratory equipment demands a tool selection capable of maintaining both electronic and small mechanical systems. That is why we built the biomedical repair tool kit series. Standard tools Kit Bag Kit Highlights: • 5 pc. Needle File Set - Great for reshaping and smoothing. • 7 pc. Nutdriver Blades Set - From 3/16" to 3/8" for small and big jobs. • Ignition Wrenches to 7/16" - Compact wrenches for tight spots. • Two Pair Cutters - One for delicate work and the other for heavier cutting. • Soldering Iron and Supplies- Everything you need to repair solder joints. Electric Sucker • A complete set of 9 screwdrivers - To handle any size fastener. • Tubing Scissors - Easily trim and reattach plastic tubing. • Long Nose and Chain Nose Pliers - For tight spots, forming and bending. - L-Keys Set mm and Inches - Monkey Plier - Grip Plair - Digital Multi meter UNI-T UT201 - Watch set - File set - Adjustable wrench 8” - Lock Opener (Positive / Negative) - 52 Pieces Socket Wrench set - Hammer Small Size • Plus many more commonly used basic hand tools. Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 02 no. |  |  |
| 34. | LIGHT THERAPY & TEETH WHITENING MACHINE Light source cool blue led power 36000mw, light intensity 2400mw upgradeable for adding functions in future two timer, can operate with remotewith wheel stand SPECIFICATIONS: - Wavelength: 455nm-475nm - Power: 36,000 mW or 36W - Light source: cool, blue LED - Light intensity: up to 2,400 mW/cm² - Voltage: 12V - Input voltage: 100V ~ 240V AC, 50/60Hz - External power supply: 36W Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 05 nos. |  |  |
| 35. | B.P APPARATUS Technical Data: Sphygmomanometer - Sphygmomanometer -Mercury Type Mercury Type Mercury Type 1 Should be Portable mercurial type. 2 Should have ISI mark. 3 Should have ON and OFF provision for mercury reservoir. 4 Should have a measuring range from 0 to 300 mmHg. 5 Should be provided with adult arm cuffs of size medium & large and pediatric cuff. 6 The control valve should have a knurled thumb control device. The leak rate should not exceed 10 mm of mercury per minute. 7 The manometer scale markings and graduations should be permanent and clearly visible and filled with pigments. 8 The internal diameter of the manometer glass tube should be 4.1 ± 0.1 mm and the thickness not less than 2 mm. 9 All plastic parts, if any used should not crack, flake, peel or disintegrate in normal use. 10 The inflating rubber bag should be capable of withstanding an internal pressure of 450 mmHg without leaking. 11 The inflating bulb should be soft and should not have any joints or ridges. 12 The mercury used should be clean, double distilled and of 99.9% purity. 13 The fastening arrangements of the cuff should be of hook and loop type (Velcro). 14 The threading and fastening arrangement of the cuff should show no sign of slip or failure when subjected to the maximum test conditions. 15 The rubber tubes used should have an internal diameter of 3 ± 0.5 mm and the external diameter should not be less than 8mm. 16 The tubes should be fitted with male and female leur connectors. 17 The housing case should be of robust design. It should have press to release lock. It should have metal hinges. The tube should be secured with metal screws and clamps. It should have mechanism to hold the lid in right angles and should prevent accidental dropping. All parts should be replaceable in case of breakage. 18 A cleaning brush to clean the manometer tube and a set of spare washers may be provided with each unit. Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 05 nos. |  |  |
| 36. | DIGITAL GLUCOMETER WITH STRIPS (BLOOD GLUCOSE METER) TECHNICAL SPECIFICATION 1. Should be a hand held meter 2. Should require no routine maintenance 3. Should have reading range/linearity from 20 to 600 mg/dl 4. Should have a maximum reading time of less than 10 seconds 5. Should use electrochemical technology 6. Should use a minimum blood sample less than 1.5µl 7. Should have a LCD display 8. Should have measuring unit in mg/dl. 9. Should have wide operating temperature 10.Should have a minimum memory of 50 11.Should have life time replacement offer 12.Should have easy code entry technique 13.Battery should be replaceable without using any tools. 14.Should have facility to ensure accuracy of measurements. 15.Should be supplied with three types of control solutions of each at least 20 ml 16.Should have safety certificate from a competent authority CE / FDA (US) / STQC CB certificate / STQC S certificate or valid detailed electrical and functional safety test GLUCOSE STRIPS GLUCOSE STRIPS 1. Should be able to use capillary blood samples. 2. Should have a minimum 4 months shelf life after opening the strip vial. 3. All strips should have at least one year expiry date from the date of supply. 4. 50 strips should be supplied along with the equipment. 5. Strips should be available in the local market. 6. Connection to PC via cable, SW application Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 05 nos. |  |  |
| 37. | Induction casting Machine: technical data: Induction casting machines capable of casting all alloys, except titanium. has been designed for the laboratories where the room is a problem or the freight costs should be minimized. Casting of precious and non-precious metals. Self-diagnostics of circuits and safety devices. Manual and fully automatic casting cycle. Equipped with motor acceleration regulator to control the metal flow to the ring. Casting power setting feature. Continues-reading control and stabilization of the temperature in the range from 800 to 2000°C Powered movement of the casting coil. Manual or automatic centrifugation control. Built-in cooling system. Use of crucibles in ceramic, coated ceramic, graphite or sintered graphite. Standard accessories: Ring holders, Rings of stainless steel, Universal rubber sprue former, crucible per type, Ceramic stirring rod, Pair of safety goggles. Max Casting Capacity 5CC precious 15 CC non precious Casting Time: 30 Gr Cr-Co 60”Power Supply 220V Voltage: 220V/ 110V Power: 2500W Centrifugal rotational speed: 500rpm Maximum melting amount: 50g (CO-CR alloy) Melting duration: 30g ≤ 185s Casting. Training must be Provided by The Supplier wherever it required Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 01 No. |  |  |
| 38. | POLYMERIZING UNITS - The polymerizing units ACRYDIG are completelybuilt in stainless steel and painted. - No electric component may come in contact,even accidentally, with water or steam, inconformity with the European Directives. - The special design of the machine closing coveravoids any scattering of vapors or condensatedripping, thus insuring the polymerization cycle isrun in steam-saturated air. - The wide opening and the proper dimensions of thepolymerization tank permit uniform arrangementof the flasks. - Due to its programmability, the ACRYDIG will runa polymerization cycle and automatically switch off at the end of the cycle itself, the flasks coolnaturally, thus avoiding the resin micro-fractures, usually caused by an excessively aggressive cooling. Instruments - The digital controller built in the polymerization units, makes programming extremelysimple (3 keys only). - The program is indefinitely stored. - The cycle is started or stopped through the RUN key. - Every cycle step is continuously displayed. Working cycle - X1 Raising to the pre-polymerizing temperature A1. - A1 Pre-polymerizing temperature. - T1 Pre-polymerization period. - X2 Raising to the polymerizing temperature A2. - A2 Polymerizing temperature. - T2 Polymerization period. - STOP automatic switching off . - Power supply Cod. Art. ACRYDIG 4 - 220 / 240 V - 50 / 60 Hz - AD044 - 110 V - 50 / 60 Hz - AD041 Technical data - Max. consumption : 1400 W - Inside Width : 26 cm - Inside Depth : 31 cm - Inside Working height : 26 cm - Capacity : 10-18 - Outside Width : 30 cm - Outside Depth : 42 cm - Outside height : 38 cm - Net weight / Gross weight : 10 kg / 11 kg Training must be Provided by The Supplier wherever it required Complete with standardaccessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 01 no. |  |  |
| 39. | SURVEYOR (Manual - Surveyor with one (1) swivel arm andcentral articulated joint, ideal forprosthesis framework design. Kit of accessories - Model-holder FIXIT. - 1 wax chisel. - 1 scanning needle probe. - 3 undercut gauges. - 1 carbon marker holder. - 1 Ø 3 mm Collet. - 1 Ø 2.5 mm Collet. Technical Data. - Width : 19 cm - Depth : 20 cm - Height : 31 cm - Net weight /Gross weight: 4 kg / 4,5 kg Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 05 nos. |  |  |
| 40. | MICRoMOTOR (bench top) - Heavy Duty Laboratory micromotor 35,000 rmp.Recommendedfor bridges crowns and resins. - Optimal quality, minimal vibration with handpice holder. - Digital Display with Speed Controller Knobe. - Power On-Off - Foot Control - Reverse and Forward Option TECHNICAL FEATURES - Power absorbed: 70 WATT - Speed: 1.000-50.000 rmp - Feeding: 230V-Hz50/60 Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 06 nos. |  |  |
| 41. | PORCELAIN PRESSING FURNACE Automatically monitors the flow of the glass and senses the movement of the ceramic into the ring in real time and dramatically increases the likelihood of a complete press, even with cracked molds. Enables successful pressing of more units per ring Presses faster than conventional pressing furnaces Minimizes or eliminates the reaction layer up to 300 different Programs Facilitates successful pressing with cracked molds Heavy Duty Oil Free Vacuum Pump 220 Volts TECHNICAL DATA: Power supply 220V 50/60Htz Power consumption; 1200 watts Muffle; Quartz Tubing Temperature: 1200 C Training must be Provided by The Supplier wherever it required Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 01 no. |  |  |
| 42. | SAND BLASTER TECHNICAL DATA: Power supply: 220v/110v 50/60 Hz. Power consumption :< 8w. Working pressure: 4kg/cm. Headpiece: twin pencil. Nozzle: replaceable. Nozzle materials: carbide/ceramic. Nozzle diameter:1.2mm/0.8mm. Particles size: 120m/50m. Optional accessories: dust extractor (reference,). Size 31cm x 42cm x42cm (h w d). Package size: 38cm x49cm x 51cm (h w d). Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 02 nos. |  |  |
| 43. | DUPLICATING MACHINE WITH DUPLICATING FLASK TECHNICAL DATA: - model duplications in little more than 6½ minutes - bubble-free surface allowing for repeated duplications unlike manual mixing machines - The adjustable disk allows you to align the flow, and facilitates insertion of the mixing tip - The aluminum support helps position the flask correctly on the work surface - The quick insertion system and flow shut-off valves prevent the formation of air bubbles during liquid delivery - Adjustable bracket design is compatible with all of duplication silicone pack sizes - The start button is specially positioned to facilitate filling the flask - Dosing pump with 1:1 mixing ratio - Adjustable plate for holding mixing tip - Anodized aluminum support with flask alignment - Adjustable container support brackets - Quick container insertion with flow shut-off valve - Provision for direct connection to all Elite Double pack sizes (1kg, 5kg, 10kg) - Sintered bayonet filters for air venting from tanks - Three set of Duplicating flask must be provide by Suppliers Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 01 no. |  |  |
| 44. | VACUME MIXTURE FOR ALGINATE AND PLASTER TECHNICAL DATA: Vacuum mixing all types of gypsums, plasters and investments. enter up to 30 separate programs using several parameters to custom mix a material. Utilizing multiple stages. It can be programmed for time, speed and paddle direction. Pre-mix stage blends powder and liquid using alternating paddle direction, eliminating the need for hand mixing Twin chamber, oil-less vacuum pump offers a powerful vacuum quickly, providing a bubble-free mix Vortex inducing paddle design quickly and efficiently incorporates mix Inlet filter prevents material from entering vacuum pump Display programmable in 5 languages (English, German, French, Italian and Spanish) Compact, space saving design Standard Container 500ml container x 2 cups Mixing Speed 0, 100~500rpm; 100rmp Increment Rotation Speed 0, 50~1500rpm; 100rmp Increment Vacuum 27.5 In Hg (930 mbar) Time Setting 0 ~ 300 seconds (5 Min) Program Memory 30 slots Multi-step Mixing 1-5 steps in a slot Vacuum Pump Capacity Exhaust velocity: 195 L/min(60Hz), 165 L/min(50Hz) Vacuum Setting Motor voltage: 220V, 1 Phase Machine Power Supply Revolution: 1000W; Rotation: 400W Safety Features Emergency Button Vibration Sensor Top Locking Sensor Door Locking Sensor Peripheral Cover Sensor Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 01 no. |  |  |
| 45. | AUTOMATED DIVESTING UNIT Use the variable speeds to ensure proper depth of cut for any investment type every time. Press to Zirconia, Press able Ceramics, , Press to Captek, Press to metal, Metal Castings Monitor the ring cutting through the clear shield Place the hot ring directly into the machine, saving an hour or more of cooling time. Select between the two most common size rings (100 and 200 gram) Touch up rings in less than one minute Set up in just five minutes with three simple connections: electricity, vacuum and compressed air (1.0scfm, 40 psi, 1.5" diameter vacuum hose connections) It is compatible with virtually any vacuum system and is small enough to fit on a bench top Power Supply 220V 50/60 Htz Ring Capacity: 100g, 200g. Delay start time: 100 g delay = 6.25 min, 200 g delay = 9.6 min Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 01 no. |  |  |
| 46. | WATER AND SPECIAL LIQUID DISPENSER: The ensures proper water powder ratio’s by automatically calculating the powder weight and then accurately dispensing the correct amount of water. Automating the dispensing of water to powder takes the human error out of the process creating more predictable working times, expansion, set times and hardness. The result is less re-work and increased profits. Programmable for up to 15 various gypsums, the user will have the flexibility needed to standardize their lab processes for optimal product performance and results. The unit dispenses both water and special liquid for your investing needs. This innovative unit takes the guess work out of measuring the powder and the water/liquid ratios by weighing the bowl with the powder on a calibrated scale and calculating the proper liquid for the volume. By dispensing the exact amount of water/liquid for the investment, inconsistencies and resulting remakes are virtually eliminated. - Produces accurate casting sizes with less finishing time - Eliminates remakes due to human error and variability in the measuring process - Calculates the powder weight and accurately dispenses water and/or special liquid, resulting in a consistent mix every time. - Programmable memory for up to 15 investments - For the Model Department: Dispense liquids such as Gypsum Hardener or other enhancing liquids in place of the Special Liquids for Investments Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 01 no. |  |  |
| 47. | Digital WAX POT - The unit HOT WAX ,firmly maintain the wax’s temperature, in order to leave un - alterated its physical charactertics - complete container of N◦4 square basins 30x30.9 , programmable regulation-from 36 ◦ C TO 120◦C. TECHNICAL FEATURES - Feeding: 220V AC – 50Hz - Temperature: 40-120◦C - Digital Display with touch Panel. - Led light for power and another for reaching the required temperature. Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 02 nos. |  |  |
| 48. | FULLY ADJUSTABLE ARTICULATOR WITH MANDIBULAR RECORDING DEVICE (3D Printer with complete set) The fully adjustable articulator has independently adjustable fossae walls that can be easily adjusted to duplicate a patient's condylar path of movement as recorded by the pantograph. Fully adjustable. Arcon construction simulates true anatomical structure. Adjustments: protrusive angle 0-60 degrees; immediate side shift 0-4mm; progressive side shift 0-30 degrees rear wall 30 degrees backward; top wall 30 degrees up, 30 degrees down; inter condylar distance 90-150mm. Medial and Superior wall inserts: removable for custom grinding; choice of curvatures. Lingual visibility allows for easy viewing and access to models from the rear of the instrument. Adjustable incisal table (T3), custom step incisal table (T2) and long centric incisal pin (P2) are standard. Comes complete with: 10 disposable mounting plates, wrench pack, carrying case and instruction manual. FACEBOW / EARBOW Face Bow Compatible with Fully Adjustable Articulator, features: Dual purpose Facebow can be located to the arbitrary or precise hinge axis or to the external auditory meatus. Independent side arms for accurate alignment so its assembled dimensions are determined by the size of the face. Ear bow mounts behind the condyle and the face bow mounts by means of mounting studs in the condyle. MANDIBULAR RECORDING DEVICE: The Mandibular Recording Device System is designed for registration and display of hinge axis movements of the human mandible. The data collected is used to program the articulator providing more predictable restorative results. - Patient set-up and recordings are accomplished in less than 10 minutes - Personalizes documentation, communication and articulator settings for each patient - Settings can be calculated for 15 different models of articulators Features Include: - Real time recordings to visually located the reference position before recording - Increased resolution - Gamma Dental Software provides many options for data analysis and patient presentation / Consultation. - Computer core i7 system with all accessories with 18.5” LCD display - Compatible with Vista, Windows XP and Windows 2000 - Optimized for recording biomechanical movements of the TMJ through new magnetic technology - Comparison of tracings - Verify accuracy - Space saving dimensions: 1.5"w x 6.5"h x 6.5"d - Portable, Plugs directly into computer USB port All Accessories including Complete with Compact unit, upper and lower face bow, clutch, Tray Clutch, Functional Clutch, Central Bearing Clutch, Deprogrammers foot switch, 2 sensor flags, 2 styli instruction manual and recording software. 3D Printing Details: - Prints in multiple build layers up to 7cm in total height - Print layer thickness is adjustable - Designed to print with different resins - Built-in LED Beamer of 405nm for curing of materials - Resolution of 34μm - Large touch screen for ease of programming and monitoring of program - Economical option for getting into 3D printing - Your choice for Ortho and all crown and bridge applications - Largest build plate - Ideal for printing patterns, drill guides, models, implant trays temporaries and flexible gingiva masks - Built-in UV Beamer of 385nm for curing of materials - Print layer thickness is adjustable from 25μm to 100μm - Resolution of 94μm - Fast and easy to change tray for different materials - complete curing of resins post printing - UV-LED with long lifetime—no bulb changes - Optimal material properties - Homogeneous irradiation distribution - Status monitoring - Intuitive menu display - To be used with Nitrogen Gas design—ideal size for desktop printing. - Irradiation source : UV 385nm - Maximum build envolpe: (WxDxH) 75x125x100mm - Layer size 25μm, 50μm, 100μm - Resolution 94μm - Machine Dimensions (W × D × H) 520 × 313 × 510mm Complete Job: Supply, Installation, Testing, Commissioning, Maintenance & Operational/ Application training to the designated staff of SIOHS during maintenance Period and production along with the materials of 06 months or as appropriate. International Training from the Manufacturer must be provided by the Supplier to the Three Persons of SIOHS. Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 01 no. |  |  |
| 49. | SEMI ADJUSTABLE ARTICULATOR WITH FACEBOW: Semi-adjustable articulator with fixed progressive side shift and adjustable condylar inclination. Features include: Progressive side shift fixed at 15 degrees. Condylar inclination adjustable 0 - 60 degrees. Retaining springs hold upper and lower members together. Ability to tilt the instrument back at a 45 degree angle. All finger screws, no tools required. Durable black anodized finish. Positive Centric Latch 110 mm Condylar Distance Comes complete with: custom (flat) incisal table, short round incisal pin, 12 disposable mounting plates, carrying case and instruction manual. FACEBOW: Face Bow Compatible with Semi Adjustable Articulator, quick and accurate method of securing an axis location without the use of expensive equipment and time-consuming techniques. Employing as the three reference points, the two external auditory meatus of the ears and the bridge of the nose, no measurements or markings on the patient are needed. The self-centering caliper design is easily adjusted, and the nasion relator utilizes the bridge of the nose to obtain an accurate and reproducible anterior reference Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 06 nos. |  |  |
| 50. | HIGH SPEED ALLOY GRINDER: - Machine polish all kind of metals in prosthetic laboratory. - Aluminum Alloy reliable and corrosion resistant. - High Speed and Low noise - Rotation speed upto 14,000 rpm - High quality Spindle with easy quick fit attachments. Technical Data: Voltage :220V 50/60 Htz Power 400 Watts Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 01 no. |  |  |
| 51. | ELECTROLYTIC POLISHER: The Electrolytic Polisher is used in restorative dentistry, it can be to polish objects that are polished by sandblasters such as cobalt chrome alloys or stainless steel, dentures frames, crown and bridges and clasps. Two cylindrical electrolytic baths can work at the same time to improve its efficiency and can polish object from all directions. - Cathode Electrolyte - Anode Electrolyte - Timer controls - Current control during the polish. - Must provide the liquid for polisher TECHNICAL DATA: - Voltage: 220V - Power : 160 Watts - Working Current: 6~7 Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 02 nos. |  |  |
| 52. | FINE BLASTING UNIT - Using to polish the surface of porcelain crowns, - Large clear Windows, with white light for clear vision. - Twin Pen using different hardness material - Two transparent Bottle Container for different micron pen blast - Air filter regulator with Pressure guage. - Ceramic Nozzle - Chromium Nozzle - Vacuum device with dust collector. - Gloves for holding the models Technical Data: - Voltage : 220V/50 Htz - Power : 150 watt - Pressure : 4 kg/cm Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 02 nos. |  |  |
| 53. | - DENTAL SURGICAL BINOCULAR LOUPES MAGNIFIER GLASSES HB + LED HEADLIGHT ## Magnification: 2.5X, 3.0X, 3.5X, 4.0X, 4.5X, 5.0X (01 Each) ## Depth of field:80mm ## Visual field:90mm ## Working distance: 300mm~420mm ## Viewing Angle: Adjustable ## Exact to adjust, turn freely ## Meet the requirement of different interpuillary distance ## Extremely High resolution optics, anti-reflection coating ## Lightweight design ## The frame is steady and reliable ## High intensity :--------------------> 15000—30000Lux ## Continuous run-time :-------------> 5 hours ## The power of the LED bulb: --------1 W ## The lifetime of the bulb: ---------10000 hour ## Very small volume and weight ## Battery pack weight: -------------100g ## Chargevoltage: ----------------AC110V/240V/50-60Hz ## Accessories: ## LED head light ## Adaptor ## Battery Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 05 nos. |  |  |
| 54. | HANAU ALCOHOL TORCH: - Volume Capacity 90 mL (3 ounces) - Adjustable Flame Intensity Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 10 nos. |  |  |
| 55. | DENTAL LASER 5TH GENERATION FOR HARD & SOFT TISSUES 3 IN 1 SYSTEM. USA/ EU/ Japan/ Equivalent. - 3 Wavelength dental laser with built-in Er:YAG (2940nm) and two Nd:YAG (1064 & 1320nm) laser generating system. - Greater efficiency and maximum output of 8 watts (15 – 40mJ, 5 – 30Hz) allows to select and perform patient specific treatment. - Quick connect swivel hand piece, detachable, autoclavable. - Comprehensive array of tips including a chisel tip. - LCD Touch panel, easy access to setup keys, convenient storage of user parameters. - Forced water cooling system. Laser; Er:YAG (2940nm); Nd:YAG (1064nm); Nd:YAG (1320nm). Max. Ave. Power; 8W; 10W; 8W. Pulse Energy (mJ/pulse); 15 - 400; 50-600; 50-500. Pulse duration (µS); H-mode: 150±50 S-mode: 500±100; H-mode: 150±50 S-mode: 500±100; H-mode: 150±50 S-mode: 500±100. Pulse Rate (Hz); 5-30; 5-50; 5-30. Hand piece type; Contact & non-contact, water and spray gun.; Contact & non-contact; Contact & non-contact. - Power: 220-230VAC, 50/60Hz Complete Job: Supply, Installation, Testing, Commissioning, Maintenance & Operational/ Application training to the designated staff of SIOHS during maintenance Period of 06 months or as appropriate. Five (05) Years full Warranty for all equipment. International Training from the Manufacturer must be provided by the Supplier to the Three Persons of SIOHS. Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 01 no. |  |  |
| 56. | DIE CUTTING SAW: - Die cutting for cutting dental models. - High speed - Low noise - Tungsten Carbide Saw blade - Excellent cutting result - Angle-adjustable magnetic working platform. TECHNICAL DATA: Voltage : 220V Rotational speed : 2800 rpm Power : 1.5A Watt : 100 W Saw blade thickness : 0.1~0.3 mm Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 01 no. |  |  |
| 57. | LASER PINDEX HOLE MACHINE: - Fabricating accurate models and dies for use in fabrication of dental restorations. - High rotational speed - Accurate drill parallel hole on under side of the model - Dept of the drill should be adjustable - Power On the motor when Work table pressing and stop after the work table is being released TECHNICAL DATA: Voltage : 220V Rotational speed : 2800 rpm Power : 1.5A Watt : 100 W Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 01 no. |  |  |
| 58. | VACUUM FORMER FOR RETAINER: Forming various kinds of plastic films, night guards, Ortho retainers with built-in high vacuum pump. Must provide 100 sheets for Standard model TECHNICAL DATA: Voltage : 220V Rotational speed : 2000 rpm Watt : 1000 W Grinding head : Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 02 nos. |  |  |
| 59. | BURNOUT FURNACE: - Multi-stage burnout furnaces large sizes and will really perform . Designed for maximum versatility, - can be used for a 1, 2, 3 or 4-stage burnout procedure and have a memory for 1-30 "open" easy to use programs. - can heat up, cool down and have a "Ready-Set-Go" smart overnight timer. Technical Data: Power Supply 115V ~ 220 V 50/60 Htz Watts : 1300 W Capacity: 10 rings approx Number of program: 30 Night time Delay : 0-99 hours Temperature : 150ºF - 2012ºF (66ºC - 1100ºC) Hold Time 0~4 Hours Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 01 no. |  |  |
| 60. | BENCH PRESS MANUAL (HEAVY DUTY) - Heavy Duty for Flasking. - Rust Proof - Powder Coating color - Easy Movement for Pressing. - Fising with two Rowal Bolts in Marble Tops Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 10 nos. |  |  |
| 61. | DENTAL POLISHING LATHE cut - Polishing Lathe ,For Continuous Duty, 2 Speeds 1400-2800 Rpm,500 Watt with guard with two different Spindles (Conical and Straight). Technical Info - ON-Off Switch - Low High Speed Regulator - Speed: 1.400/2.800 R.P.M. - Power: 500 Watt - Sizes Mm 520x120x235 Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 05 nos. |  |  |
| 62. | VIBRATOR (Heavy Duty) - They have a removable rubber surface with counter –sinkedges for an easy cleaning. Hugh noiseless and steadinessby means of the perfect vertical movement. Technical Features - Overall size: Table dimensions 10" x 6" Useful work plane: Table dimensions 10" x 6" - Absorption: 150 watt - Feeding: 230v/50hz - Regulator for speed control - On-Off Switch with two speed selection switch with Led Showing Power. Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 05 nos. |  |  |
| 63. | Digital ELECTRIC WAX Carving KNIFE (Dual) - A duo-option modeling station which enables two different handpieces to be readysimultaneously. Each handpiece can have three individually programmable temperatures which can be easily switched over. The exchange mechanism of the modeling tips is a safe and simple process. Suitable for all types of wax-up work and modeling, the device is best combined with the Waxjet when making full dentures. The soft and stress-free treatment of the waxes is improved. - The instrument alloy used for the probes has superior conductivity, being 20 times more than that of stainless steel. The copper tips are gold plates for enhanced heat productivity. - Must provide different tips complete kit (Offset explorer, Discoid, Mini-Beavertail, Large Beavertail, Hollenbeck - Previously temperature recorded in memory. - Silicon cool grip hand piece. - Carving pencil holder. - Touch Control Temperature. TECHNICAL FEATURES - Feeding: 220V AC – 50Hz - Temperature: 45-200◦C Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 01 no. |  |  |
| 64. | WEIGHING BALANCE - Analytical balance range, 0.0001-120g Qty 01. - Analytical balance range, 1g-5kg Qty 01. Digital Display Accurate Weight. Stainless steel platform Zero tare option. Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 02 Nos. |  |  |
| 65. | HYDRAULIC DENTAL PRESS - The hydraulic press is used by the dental technician to mould mobile prosthesis, through the compression of the resin inside the muffle. - The hydraulic press “GD” is able to clamp up to 3 muffles and reach a maximum pressure of 400 bars, exerting on the muffles a force of 15000 Kg (150.000 n). - This is an excessive pressure and in order not to damage themuffles, the working pressure should be limited to 200 bar. - The press features a manual pump operated directly by the user through a lever. - The strength operated on the lever is proportioned to the inner pressure; therefore with an increasing pressure the strength used by the operator mustincrease as well. This is to avoid reaching too high pressures which coulddamage the machine. - The oil-pressure circuit is fully airtight and doesn’t need any maintenance. Technical features - Working pressure: 200 bar - Max pressure : 400 bar - Max force : 15.000 Kg - Base material : hydraulic cast iron - Weight :20 Kg Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 03 nos. |  |  |
| 66. | PLASTER DISPENSING UNIT - Multifunctional turn/ push button for simple and quick programming and programme selection - Easy operation – error-free working for everyone - Reproducible mixing ratio at the push of a button, accurate to within a gram - Integrated scales automatically measure the dosage or for manual weighing - Enables clean and dust free operation﻿ - Refine the plaster - Strong Vibrator for pouring the powder. - Inside rubber container to protect humidity and pollution. Technical Data: - Power Supply : 220 Volts 50/60 Htz - Mains supply: //220 V, 50/60 Hz - Output: 95 W - Capacity Powder: appr. 30 kg - Time to alot 20ml: 100g appr. 30s - Easily Installede in Wallmount - A strong vibrator is activated while the dispenser is open to pour the powder. Training must be Provided by The Supplier wherever it required Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 01 no. |  |  |
| 67. | VACUUM MIXER WITH VIBRATOR - The mixer for coating and plasters can perform the entire cycle (paddling,pouring of the coating, vibration)in vacuumconditions.The incorporated vacuum pump is a “dry pump” without oil and at is entirely maintenance free . The speed of paddle is varying from 100 to 500rpm.The machine is supplied with a digital electronic trimmer that can be used to set the paddling time. As the pre-set time elapses(shown on the display), the padding stop and the pouring and the vibration can start. During the pouring of the coating , the vibrator can be turned on and the intensity regulated. All the quantities are visualized by a big LCD Display backlit. This vacuum paddle and pouring processing assures excellent result, with compact coating air bubbles and blowholes free. Technical Data - Power Supply 230 V-50Hz - Power 300W - Work plane 15x23cm - Dimension(WxDxH) 15x23x9.5cm - Three different sizes Container provide with Machine. Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 01 no. |  |  |
| 68. | HANGING ENGINE motor - Complete with gear type connection, 180W, 25,000rpm. - Complete with electronic rheostat flexible shaft - handpiece with four collets: 0/1/2,35/3 mm - Foot Control - Cable attachement Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 10 nos. |  |  |
| 69. | PRESS DOUBLE FLASK: - flat bench press for 2 alloy flasks (upper+lower). unit weight is 2.175kg - The castings of the flasks shall be clean and free from surface and casting flaws. The finished castings shall have a reasonably good finish and shall require no further machining except at the mating lugs. All sharp edges shall be rounded off. The two halves of the flask shall engage without any jerk and shall not leave any gap at the mating -edge when observed against light. Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 10 nos. |  |  |
| 70. | FLASK - Dimensions and requirements of dental flasks for lower and upper dentures. - Material- Brass/ Aluminum alloy, Designation A-6M of IS: 617-l959* or brass, Grade 3 of IS: 292-196lt. - The castings of the flasks shall be clean and free from surface and casting flaws. The finished castings shall have a reasonably good finish and shall require no further machining except at the mating lugs. All sharp edges shall be rounded off. The two halves of the flask shall engage without any jerk and shall not leave any gap at the mating -edge when observed against light. Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 10 nos. |  |  |
| 71. | MICROTOMES: Technical Data: Features include digital section counter, 60 mm vertical stroke length, section thickness adjustable between 0.5 and 60 microns. Low maintenance and backlash-free micrometer feed system has pre-lubricated cross roller bearings for consistent, accurate sectioning. Unit comes fully restored, serviced and calibrated to factory specifications. Accessories Included: Low or high profile knife holder. Can be outfitted with optional screw down specimen clamp and/or non disposable steel knife holder. Dimensions: 17′ wide X 19′ length X 10′ height Training must be Provided by The Supplier wherever it required Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 01 no. |  |  |
| 72. | BINOCULAR(COMPOUND MICROSCOPE) Binocular polarizing Microscope 40x-600x Technical Data: Eyepieces: PLAN10X (22mm field of view) flat-field eyepieces with eye-guard and PL10X22mm with reticle Objectives: infinity Plan distressing objectives: 4XS, 10XS, 40XS, 60XS Head: 30° inclined, binocular head Midst Head: build-in analyzer and Bertrand lens, rotating range of analyzer: 90°, a dose value of vernier 6', being able to shift out the Optical way. With Bertrand lens diverter and focusing handwheel, centering adjustment Compensator: λ plate (f18mm, first-order red, light path aberration: 551nm); λ/4 plate (f18mm, light path aberration: 147.3nm); Quartzose chock (12×28mm, I - IV) Nosepiece: quadruple distance should be adjustable. Upper limit stopper. Stage: rotatable circular, 360° rotatable circular worktable, a dose value of mark 1, adose value of vernier 6' Condenser: NA1.2/0.22 swing type achromatic condenser Polarizer: 360° (0°/90°/180°/270° four type graduation) rotatable vibration, apparatus with locking, being able to shift out the optical way Illumination: Koehler illuminator system, White LED Illuminator. With adjustable brightness, CE certified Power Supply: 90V~240V fluctuate of voltage Total Overall Height of Microscope: 485mm Accessories TWO Group Observation Systems Each for 2 persons (1 teaching head). Face-to-face or in side-by-side configuration of teaching head. Pointer with bright LED. Training must be Provided by The Supplier wherever it required Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 25 nos. |  |  |
| 73. | DENTAL MODELS - Model of human canine tooth showing enamel-dentine-pulp that should be anatomically designed with sensible hardness diff. between enamel, dentine red pulp should be larger than anatomical size. - Model of human lower molar tooth should be anatomically designed with enamel, dentine, .pulp layers with sensible hardness diff. between enamel, dentine red pulp. Should be larger than anatomical size. - Model of human maxillary molar with anatomically designed enamel, dentine and pulp layers with sensible hardness diff. between enamel, dentine red pulp. Should be larger than anatomical size. - Model of human upper/lower incisor tooth with anatomically designed enamel, dentine and pulp layers with sensible hardness diff. between enamel, dentine red pulp. Should be more than larger than anatomical size. - Model of human lower premolar tooth should be larger than anatomical size. - Model of human maxillary/ mandibular arch showing various types of crown/briges and partical denture. - Model showing concept of post and core treatment of human tooth with removable crown and post. - Model showing human molar teeth with inlay and gold crown restorations. - Individual tooth models of permanent 32 human teeth. Made of plastic or resin. Individual tooth Model size should be 2.5 times or larger than the anatomical tooth reproducing all anatomical details. - Individual tooth models of Primary20 human teeth. Made of plastic or resin. Individual tooth Models should be of anatomical size or larger with all anatomical details. - Model tooth with caries no. 21 - Model tooth with caries no. 24 - Model tooth with caries no. 26 - Model tooth with caries no. 46 - Model tooth no. 17 with fissure caries - Model tooth no. 21 for adhesive restoration model teeth for Adhesive restoration exercises with etching gel. - Partially dentate upper jaw (Kennedys classes) - Partially dentate LOWER jaw(Kennedys classes) - Model tooth for pulpotomy exercises filled with wax. - Pediatric anesthesia model with holder HA incl. syringe. - Spare battery for anesthesia models. - Syringe for training model - Endodontic tooth No. 21 model teeth with root canals, fitting models, execution in transparent plastic material with red pulpal walls and ivory crowns in hard thermosetting material for preclinical training in root canal treatment and cavity preparation. - Endodontic tooth No. 24 model teeth with root canals, fitting models, execution in transparent plastic material with red pulpal walls and ivory crowns in hard thermosetting material for preclinical training in root canal treatment and cavity preparation. - Endodontic tooth No. 26 model teeth with root canals, fitting models, execution in transparent plastic material with red pulpal-walls and ivory crowns in hard thermosetting material for pre-clinical training in root canal treatment and cavity preparation. - Endodontic tooth No. 46 model teeth with root canals, fitting models, execution in transparent plastic material with red pulpal walls and ivory crowns in hard thermosetting material for pre-clinical training in root canal treatment and cavity preparation. - Permanent Tooth brushing demonstration model larger than life size with large tooth brush and large Inter dental brush. - Deciduous Tooth brushing demonstration model larger than life size with tooth brush. - Abnormal teeth alignment tooth brushing demonstration model. - Periodontal disease model. # 4x size sealant study model # 3 stages of caries progression model (C1 to C3) - Pathological mandible natural size w.hinged segment representing the nerve system and the roots supp. pulp infection and an impacted wisdom tooth caries, calculus and a sagittal section through crown and roots of tooth 44. - Pathology. molar socket model scale 5:1 lateral exposure repr. caries, calculus, abscess, abrasion. - Caries demonstration model scale 5:1 with 3 removable sectioned teeth representing caries progression in 3 stages and pulp affection. - Sectioned molar demonstration model scale 2.5 :1 representing tilenerve system and the roots a crown with inlay and an amalgam filling an abscess and an impacted wisdom tooth with signs of dissociation by inflammation. - Demonstration model with individual contact points. - Strip crown forms, glass-clear. Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 01 Set Each. |  |  |
| 74. | DENTAL MOULDS - Individual silicon rubber tooth moulds of 28 human permanent teeth Made of transparent flexible rubber. - Upper and lower jaw Silicone rubber mould maxillary and mandibular deciduous dentition. - Upper and lower jaw Silicone rubber mould maxillary and mandibular permanent dentition. - Rubber moulds for model partial denture (upper &Lower Jaws) - Rubber moulds for model complete denture(upper &Lower Jaws) - Rubber moulds for Edentulous arch (Upper & Lower) Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 06 sets each |  |  |
| 75. | KENNEDY’S DENTAL MOULDS - Kennedy’s Class I (Upper & Lower) (01 set) - Kennedy’s Class II (Upper & Lower) (01 set) - Kennedy’s Class III with all modifications (Upper & Lower)(01 set) - Kennedy’s Class IV (Upper & Lower)(01 set) Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 01 sets each. |  |  |
| 76. | PHANTOM HEAD WITH TABLE MOUNT ASSEMBLY COMPLETE REMOVAL TEETH AND JAWS AND GUMS FOR WORK STATIONS: - Phantom head with removable Jaws, removable teeth and Gums - Brackets set for clamping with Work bench. - Rod and all accessories. - Half mask / full mask for Phantom head. All the plumbing and pre installation work done by the supplier Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 15 nos. |  |  |
| 77. | CENTRIFUGE MACHINE FOR ORAL PATHOLOGY LAB Floor centrifuges, refrigerated/non refrigerated, high/low speed. Centrifuges for collection of microorganisms, cells and cellular debris, and precipitates. These high speed floor units operate at max. speeds of 10,000-25,000 rpm, and are also capable of sedimenting viruses and cellular organelles. Low-speed refrigerated units or nonrefrigerated ones which operate at room temperature are also available. Bench Top high Speed Lab Centrifuge with Digital Speedometer, Digital Timer in the range of 0-59 Minutes and Speed Regulator. Maximum Speed 20000 r.p.m.(without Load). Supply completed with dust cover, instruction manual, cord & plug to work on 220 volts 50 Hz A.C. Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 01 no. |  |  |
| 78. | COLORIMETER FOR ORAL PATHOLOGY LAB Wavelength range 400 to 710nm Wavelength selection 8 in-built gelatin filters: 430, 470, 490, 520, 540, 580, 600 and 710nm Measurement ranges 0-100%T, 0-1.50Abs, 0.1 to 1000 Concentration Resolution 1%T, 0.01Abs, 0.1 to 1 Concentration Bandwidth Typically 40nm Blank drift 2%/hr after 15 minute warm-up Photometric linearity 1%T Light source Tungsten filament lamp Output Analogue, 10mV per digit Power 230V, 50Hz Size (l x w x d)300 x 355 x 120mm Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 01 no. |  |  |
| 79. | INCUBATOR FOR ORAL PATHOLOGY LAB 1. LCD display, with 12V to 24V DC power for ease of portability 2. PeratingTemperatureRange: From 5°C above ambient to 80°C 3. TimeRange: 1min ~ 99h59min 4. Temp.Control Accuracy: ≤±0.5°C 5. Display Accuracy: 0.1°C 6. Heating Time: <6min(from 20°C to 57°C) 7. Temp. Control Accuracy @57°C: ≤±0.35°C Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 01 no. |  |  |
| 80. | DENTAL MILLING MACHINE Milling surveyor with double-articulated arm for the molding and positioning of attachments and supporting any kind of turbine engine or micro motor, with vertical micrometric adjustment. Equipped with axial milling device. Accessories store with low-voltage halogen lamp and air gun for Air supply. Kit of accessories: • Model-holder FIXIT. • 1 wax chisel. • 1 scanning needle probe. • 3 undercut gauges. • 1 carbon marker holder. • 1 Ø 3 mm Collet. • 1 Ø 2.5 mm Collet. • 1 Ø 2.5 mm Collet. Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT |  |  |  |
| 81. | BUNSEN BURNER - Quick On/Off Valve - Flame Adjustable. - High Low option - Brass material used in Burner Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 50 nos. |  |  |
| 82. | X Ray Manual Developing Unit - Three separator Container for Developer, Fixer and Water - Fixed in wall mount - Container capacity 01 liters - Electrically motor driven - Film Collector - Screen for X-Rays Viewing - 220 Volts operated Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT |  |  |  |
| 83. | PLATELET RICH FIBRIN (PRF & CGF MACHINE) PRF Machine Dual Function with PRF and CGF centrifuge should has a «switch» button that affords you to select Various settings mode for the PRF PRP and CGF. The values of these settings are preset and locked except for PRP / PRF / CGF (we can set Other values). Simply turn the knob on the desired mode of operation, thecorresponding LED will light on. The machine is ready. Simply press thebutton to start the centrifugation. Centrifuge Machine with Complete Accessories Including Instruments: PRF Scissors PRF Forceps PRF Mini Tray PRF Forceps giraffe PRF Bowl Pad PRF Compactor Small Compactor Big Double Spoon Tube-holder Tourniquet Sticker Tubes and Blood Collectors | 01 no. |  |  |
| 84. | ELECTRIC CAUTERY MACHINE Bipolar: bipolar coagulation Of the 6 Working Modes 300W Electrosurgical Unit: 1. The electrosurgical generator is used to control bleeding and surgical cutting in all general and open surgical procedures 2. Touch control panel, three-way power output, with different audible and visual indication during the process of outputting. 3. Patient return electrode monitoring system and Power Peak System that provides surgeons with improved performance at Lower power settings, minimizing the risk of tissue damage. 4. Applicable for Prostate Vaporizing Surgical Operation(TUR), perfectly used under water. The electrosurgical generator can be cooperated with argon plasma coagulator, endoscope and laparoscope. 5. Pencil Remoting Output. The output power can be controlled by the remoting control pencil Technical Specifications of the electric cautery machine, cautery machine, Diathermy Cautery Skin surgery : Power: 220V±22V, 50Hz±1Hz(110V±11V, 60Hz) Operating frequency: 512KHz Power rating: 880VA±10% Six working modes: Monopolar Cut a) Pure cut: 0W~300W( Load 800Ω) b) Blend1: 0W~250W(Load 800Ω) c) Blend2: 0W~150W(Load 800Ω) MonopolarCoag d) Contact Coag: 0W~80W( Load 800Ω) e) Forced Coag: 0W~120W( Load 800Ω) Bipolar f) Bipolar Coagulation: 0W~50W( Load 100Ω) Power Consumption: ≤880VA Electrosurgical Pencil: 5pcs Power Control Pencil: 1pc Electrosurgical Pad: 10pcs Electrosurgical pad cable: 1pc Footswitch: 1set Bipolar Forceps and cable: 1set Complete with standard accessories UK/USA/JAPAN/GERMANY OR EQUIVALENT | 01 no. |  |  |
| 85 | Reagent Refrigerator Capacity 200 Liters Cooling Type static Temperature Range 1-10 O C 1 Door Glass Three Baskets / Shelves Dimensions W * D * H (553*553*1575) Made in Germany/uk/Japan or equivalent | 02 Nos |  |  |
| 86 | Water bath Working temperature range 20 - 99.9 °C Temperature control PID Temperature stability ±0.2 °C Display LED Display resolution 0.1 °C Integrated programmer not available Heater capacity 2000 W Bath opening / bath depth (W x L / D) 50 x 30 / 18 cm Filling volume 8 ... 26 Liters Dimensions (W x L x H) 56 x 35 x 32 cm Weight 20 kg Ambient temperature 5 - 40 °C Classification according to DIN 12876-1 1 .0 Made in Germany/Uk/Japan or equivalent | 01No |  |  |
| 87 | Staining Rack Holds 25 Slides Length (Metric) Slide 22mm Material Stainless Steel Product Type Cover slip Staining Rack Width (Metric) Slide 22mm Made in Germany/Uk/Japan OR EQUIVALENT | 02 |  |  |
| 88 | Vertical Staining Jar with Cover Wide top jar especially suitable for staining slides inscribed on one end. Holds 8 single 3 x 1" (75 x 25mm) slides vertically or 16 slides back-to-back. Made of soda lime glass. Approximate inside dimensions: 90 x 50 x 30mm (3.5" x 2" x 1.18"). Holds about 50ml of liquid. Made in Germany/Uk/Japan OR EQUIVALENT | 04 NOS |  |  |
| 89 | Wax Dispenser Weight3.7kg Dimensions (O.D. x H) 21.5 x 41cm Electrical Voltages Frequency Power 220–240V (~) 50/60 Hz 250 VA Maximum supply voltage fluctuation must not exceed 10% of nominal Fuses Mains plug Mains input 5A 240V (where applicable) 230V (~) units F1.6A H250V 115V (~) units F3.15A H250V Environmental Requirements General Indoor use only Temperature (operating limits) +5°C to +40°C (+41°F to +104°F) Temperature (Transport & Storage) -25°C to +55°C (-13°F to 131°F) +70°C (158°F) for short exposure Relative Humidity Max. 80% RH up to 31°C Decreasing linearly to 50% RH at 40°C Altitude Up to 2000 m (6,500 ft) Pollution Degree 2 | 01 |  |  |
| 90 | Heamoglobinometer Haemoglobinometer square tube type The kit Specifications Comparator Holder: -1 Nos. It is black in colour. The back of the holder is fitted with milky colour screen made of acrylic. In front of the holder there are three windows. The two side windows are fitted with color comparison. The sizes of all the three windows are Height 3 cms. (Minimum) Width 0.5 cms Square H.B. Tube: -2 Nos. Square H.B. Tube is graduated on both the side for the measurement of haemoglobin in respect of percentage and gram. H.B. Pipette -1 Nos. 20 microlite H.B. Pipette with latex tube and mouth piece Amber coloured glass bottle -1 Nos. Glass Stirrer -1 Nos. Dropper with teat -1 Nos. Cleaning Brush -1 Nos. Spare latex tubing one meter -1 Nos. All the above are fitted in thermoform tray. Further packed in plastic box. | 02 |  |  |
| 91 | ESR rods Fast: reading time 30 min and 60 min equivalent to the 1 h and 2 h of the Westergren method High quality vacuum glass tubes used for blood collection and running tests: ensure more precise sedimentation rate than plastic tubes Cost of tests equals cost of tubes only! No blood handling: conceived to simplify ESR analysis as much as possible, avoiding sample manipulation and removing the risk of operator infection Positive detection: optional barcode reader for ESR tubes avoids identification errors Traceability: built-in thermal printer with automatic print-out of sedimentation rate and kinetic curve of sedimentation Simplicity: stand-alone instrument with mono-directional LIS interface Throughput: 60 samples per hour | 10 |  |  |