Nuclear Institute of Medicine and Radiotherapy (NIMRA) Jamshoro

Tender for Purchase of Linear Accelerator, MRI System and C.T. Simulator to be installed at Nuclear Institute of Medicine And Radiotherapy (NIMRA) Jamshoro 2015-16

Table of Contents

Section I.  Invitation for Bids

Section II.  Instructions to Bidders

Table of Clauses

Section III.  Bid Data Sheet

Section IV.  General Conditions of Contract

Table of Clauses

Section V.  Special Conditions of Contract

Table of Clauses

Section VI. Schedule of Requirements

Section VII. Technical Specification

Section VIII.  Sample Forms and Procedures

1. Bid Forms and Price Schedules

2. Form of Bid Security / Earnest Money (Bank  Guarantee)

3. Form of Contract Agreement

4. Performance Security Forms

5. Form of Completion Certificate

6. Form of Operational Acceptance Certificate

7. Integrity Pact

Acronyms

BDS	Bid Data Sheet

CIF	Cost, Insurance and Freight

CIP	Carriage and Insurance Paid to (place)

CPM	Critical Path Method

EDI	Electronic Data Interchange

ENAA	Engineering Advancement Association of Japan

EXW	Ex factory, ex works or ex warehouse

FCA	Free Carrier

FIDIC	Fédération Internationale des Ingénieurs Conseils (International Federation of Consulting Engineers)

FOB	Free on Board

GCC	General Conditions of Contract

IBRD	International Bank for Reconstruction and Development

ICC	International Chamber of Commerce

IDA	International Development Association

IFB	Invitation for Bids

ITB	Instructions to Bidders

SBD	Standard Bidding Document

SCC	Special Conditions of Contract

TS	Technical Specifications and Drawings

UNCITRAL	United Nations Commission on International Trade Law

LC		Letter of Credit

# Section I.  Invitation for Bids

Islamic Republic of Pakistan

Name of Project: 	Provision of Purchase of Linear Accelerator, MRI System and C.T. Simulator, to be installed at Nuclear Institute Of Medicine And Radiotherapy (NIMRA) Jamshoro

- This invitation for bids follows the general procurement notice for this project that appeared in newspapers and on Sindh Government website.

- The Department of Health now invites sealed bids from eligible bidders for Provision of Linear Accelerator, MRI System and C.T. Simulator with all accessories and fittings.

- Bidding will be conducted under SPP Rule 2010 (Amended 2013) clause 46(2). Bidding procedures specified in this document is open to all eligible bidders.

- Interested eligible bidders may obtain further information from the Office of Director General Health Services, Health Department, Government of Sindh, Hyderabad.

- A complete set of bidding documents may be purchased by interested bidders on the submission of a written application to the address below and upon cash payment of a non-refundable fees of Rs. 2000/-.

- Bids must be delivered to the address below up to 12:00 noon.  All bids must be accompanied by a bid security / earnest Money @ 2 ½  % of the Total bid cost in local currency or an equivalent amount in a freely convertible currency. Late bids will be rejected.  Bids will be opened in the presence of the bidder’s authorized representatives who choose to attend at the address below at 01:00 p.m.

Address:

I.I. Depot, Near JPMC,

Rafique Shaheed Road, Karachi

# Section II.  Instructions to Bidders

## Table of Clauses

A.  Introduction

1.	Source of Fund

2.	Eligible Bidders

3.	Eligible Plant, Equipment, and Services

4.	Cost of Bidding

B.  The Bidding Documents

5.	Content of Bidding Documents

6.	Clarification of Bidding Documents;

7.	Amendment of Bidding Documents

C.  Preparation of Bids

8.	Language of Bid

9.	Documents Comprising the Bid

10.	Bid Form and Price Schedules

11.	Bid Prices

i)	Bid Currencies

13.	Bid Security

14.	Period of Validity of Bid

15.	Format and Signing of Bid

D.  Submission of Bids

16.	Sealing and Marking of Bids

17.	Deadline for Submission of Bids

18.	Late Bids

19.	Modification and Withdrawal of Bids

E.  Bid Opening and Evaluation

20.	Opening of Bids by Purchaser

21.	Clarification of Bids

22.	Preliminary Examination of Bids

23.	Conversion to Single Currency

24.	Technical Evaluation

25.	Financial Evaluation

26.	Contacting the Purchaser

F.  Award of Contract

27.	Post-qualification

28.	Award Criteria

29.	Purchaser’s Right to Accept Any Bid and to Reject Any or All Bids

30.	Notification of Award

31.	Signing the Contract Agreement

32.	Performance Security

33.	Corrupt or Fraudulent Practices

Instructions to Bidders

A.  Introduction

| 1. Source of Fund | 1.1 The Government of Sindh allocated funds for Provision of Purchase of Linear Accelerator, MRI System and C.T. Simulator to be installed at Nuclear Institute Of Medicine And Radiotherapy (NIMRA) Jamshoro |
| --- | --- |
| 2. Eligible Bidders | 2.1 This Invitation for Bids is open to all original manufacturer / sole distributor in Pakistan for supply & installation of Provision of Linear Accelerator, MRI System and C.T. Simulator. 2.2 Bidders shall not be under a declaration of ineligibility for corrupt and fraudulent practices |
| 3. Eligible Plant, Equipment, and Services | 3.1 For the purposes of these bidding documents, the words “facilities”, “plant and equipment”, “installation services”, etc., shall be construed in accordance with the respective definitions given to them in the General Conditions of Contract. 3.2 The Bidder shall sign the integrity pact attached as annexure. |
| 4. Cost of Bidding | 4.1 The Bidder shall bear all costs associated with the preparation and submission of its bid, and the Purchaser will in no case be responsible or liable for these costs, regardless of the conduct or outcome of the bidding process. |

B.  The Bidding Documents

| 5. Content of Bidding Documents | 5.1 The facilities required, bidding procedures, contract terms and technical requirements are prescribed in the bidding documents. The bidding documents include the following sections: Invitation for Bids (IFB) Instructions to Bidders (ITB) Bid Data Sheet (BDS) General Conditions of Contract (GCC) Special Conditions of Contract (SCC) Schedule of Requirements (SR) Technical Specifications (TS) Forms and Procedures (FP) 1. Bid Form and Price Schedules 2. Bid Security Form 3. Form of Contract Agreement 4. Performance Security Forms 5. Form of Completion Certificate 6. Form of Operational Acceptance Certificate 5.2 The Bidder is expected to examine all instructions, forms, terms, specifications, and other information in the bidding documents. Failure to furnish all information required by the bidding documents or submission of a bid not substantially responsive to the bidding documents in every respect will be at the Bidder’s risk and may result in rejection of its bid. |
| --- | --- |
| 6. Clarification of Bidding Documents; | 6.1 A prospective Bidder requiring any clarification of the bidding documents may notify the Purchaser at the Purchasers address indicated in the Invitation of Bids. At least ten days prior to deadline for submission of bids. 6.2 The Bidder is advised to visit and examine the site where the facilities are to be installed and its surroundings and obtain for itself on its own responsibility all information that may be necessary for preparing the bid and entering into a contract for supply and installation of the facilities. The costs of visiting the site shall be at the Bidder’s own expense. |
| 7. Amendment of Bidding Documents | 7.1 At any time prior to the deadline for submission of bids, the Purchaser may, for any reason, whether at its own initiative, or in response to a clarification requested by a prospective Bidder, amend the bidding documents. 7.2 The amendment will be notified in writing to all prospective bidders that have purchased the bidding documents and will be binding on them. Bidders are required to immediately acknowledge receipt of any such amendment, and it will be assumed that the information contained therein will have been taken into account by the Bidder in its bid. 7.3 In order to afford prospective Bidders reasonable time in which to take the amendment into account in preparing their bid, the Purchaser may, at its discretion, extend the deadline for the submission of bids, in which case, the Purchaser will notify all bidders in writing of the extended deadline. |

C.  Preparation of Bids

| 8. Language of Bid | 8.1 The bid prepared by the Bidder and all correspondence and documents related to the bid exchanged by the Bidder and the Purchaser shall be written in English. Supporting documents and printed literature furnished by the Bidder may be written in another language, as long as such literature is accompanied by a translation of its pertinent passages in the language specified in the Bid Data Sheet, in which case, for purposes of interpretation of the bid, the translation shall govern. |
| --- | --- |
| 9. Documents Comprising the Bid | 9.1 The bid prepared by the Bidder shall comprise the following documents: (a) Bid Form duly completed and signed by the Bidder, together with all Attachments identified in ITB Sub-Clause 9.2 below. (b) Price Schedules duly completed by the Bidder. 9.2 Each Bidder shall submit with its bid the following attachments: (a) Attachment 1: Bid Security A bid security furnished in accordance with ITB Clause 13. (b) Attachment 2: Power of Attorney A power of attorney, duly notarized, indicating that the person(s) signing the bid has(ve) the authority to sign the bid and thus that the bid is binding upon the Bidder during the full period of its validity in accordance with ITB Clause 14. (c) Attachment 3: Bidder’s Eligibility and Qualifications The documentary evidence of the Bidder’s qualifications to perform the contract, if its bid accepted, shall establish to the Purchaser’s satisfaction that the Bidder has the financial, technical, production, procurement, shipping, installation and other capabilities necessary to perform the contract, and in particular, meets the experience and other criteria outlined in the Bid Data Sheet. - Attachment 4: Eligibility and Conformity of the Facilities - The bidder shall furnish the all-original equipment manufacture (OEM) certificate of equipment supplied/ offered and a certificate that the machine / equipment being used in the country of origin. - The documentary evidence of the conformity of the facilities to the bidding documents may be in the form of literature, drawings and data, and shall furnish (i) a detailed description of the essential technical and performance characteristics of the facilities; (ii) a list giving full particulars, including available sources, of all spare parts, special tools, etc., necessary for the proper and continuing functioning of the facilities for the period named in the Bid Data Sheet, following completion of facilities in accordance with provisions of contract. (e) Attachment 5: Subcontractors Proposed by the Bidder The Bidder shall include in its bid details of all major items of supply or services that it proposes to purchase or sublet, and shall give details of the name and nationality of the proposed Subcontractor, including vendors, for each of those items. Bidders are free to list more than one Subcontractor against civil works of the facilities. Quoted rates and prices will be deemed to apply to whichever Subcontractor is appointed, and no adjustment of the rates and prices will be permitted. The Purchaser reserves the right to delete any proposed Subcontractor from the list prior to award of contract, and after discussion between the Purchaser and the Contractor. - Attachment 6: Tender Purchase Receipt - The original tender purchase receipt issued by the purchaser should be attached along with Technical bid. - Bidders shall purchase separate tender document and furnish purchase receipts for alternate offer, in case they want to submit alternate offer. Any item’s bid with alternate offers without separate purchase receipt (Original) is supposed to be rejected. |
| 10. Bid Form and Price Schedules | 10.1 The Bidder shall complete the Bid Form and the appropriate Price Schedules furnished in the bidding documents as indicated therein and in the Subsection “Bid Form and Price Schedules” of the bidding documents, following the requirements of ITB Clauses 11 and 12. |
| 11. Bid Prices | 11.1 Unless otherwise specified in the Technical Specifications, bidders shall quote for the entire facilities basis such that the total bid price covers all the Contractor’s obligations mentioned in or to be reasonably inferred from the bidding documents in respect of the procurement, delivery, installation and commissioning of the machines. This includes all requirements under the Contractor’s responsibilities for Inspection, Training & commissioning of the facilities and where so required by the bidding documents, the acquisition of all permits, approvals and licenses, etc.; the operation, maintenance and training services and such other items and services as may be specified in the bidding documents, all in accordance with the requirements of the General Conditions of Contract. Items against which no price is entered by the Bidder will not be paid for by the Purchaser when executed and shall be deemed to be covered by the prices for other items. 11.2 Bidders shall give the prices in the following manner and detail called for in the Price Schedules. Separate numbered Schedules shall be used for each of the following elements. The total amount from each Schedule (1 to 4) shall be summarized in a Grand Summary (Schedule 5) giving the total bid price(s) to be entered in the Bid Form. Schedule No. 1 Plant and Equipment Supplied from Abroad Schedule No. 2 Plant and Equipment Supplied from within the Purchaser’s Country Schedule No. 3 Local Transportation Schedule No. 4 Installation Services (Electrical works and other services) Schedule No. 5 Grand Summary 11.3 The prices quoted by the Bidder shall be fixed during the Bidder’s performance of the contract and not subject to variation on any account. A bid submitted with an adjustable price quotation will be treated as non-responsive and rejected. |
| 12.Bid Currencies | - Prices shall be quoted in Pak currency and US dollars |
| 13. Bid Security | 13.1 The Bidder shall furnish, as part of its bid, a bid security in the amount stipulated in the Bid Data Sheet in the currency of Pakistan, or in the equivalent amount in a freely convertible currency. 13.2 The bid security shall, at the Bidder’s option, be in the form of a pay order, bank draft /or bank guarantee from a reputable entity located in the country. The format of the bank guarantee shall be in accordance with the form of bid security included in the bidding documents; Bid security shall remain valid for a period of thirty (30) days beyond the original bid validity period, and beyond any extension subsequently requested under ITB Sub-Clause 14.2. 13.3 Any bid not accompanied by an acceptable bid security / short bid security shall be rejected as being non-responsive. 13.4 The bid securities of unsuccessful bidders will be returned as promptly as possible, but not later than thirty (30) days after the expiration of the bid validity period. 13.5 The bid security of the successful Bidder will be returned when the Bidder has signed the Contract Agreement, and has furnished the required performance security, 13.6 The bid security may be forfeited (a) if the Bidder withdraws its bid during the period of bid validity specified by the Bidder in the Bid Form (b) in the case of a successful Bidder, if the Bidder fails to supply within the specified time limit - to sign the Contract Agreement, in accordance with ITB Clause 31, or (ii) to furnish the required performance security, in accordance with ITB Clause 32. |
| 14. Period of Validity of Bid | 14.1 Bids shall remain valid for the period of Ninety (90) days named in the Bid Data Sheet after the closing date prescribed by the Purchaser for the receipt of bids. A bid valid for a shorter period shall be rejected by the Purchaser as being non-responsive. 14.2 In exceptional circumstance, the Purchaser may solicit the Bidder’s consent to an extension of the bid validity period. The request and responses thereto shall be made in writing. If a Bidder accepts to prolong the period of validity, the bid security shall also be suitably extended. A Bidder may refuse the request without forfeiting its bid security. A Bidder granting the request will not be required or permitted to modify its bid. |
| 15. Format and Signing of Bid | 15.1 The Bidder shall prepare bid comprise one single envelope containing separately financial proposal and technical proposal in original. The Envelop shall be marked as “Financial Proposal” and “Technical Proposal” in bold and legible letter to avoid confusion. 15.2 The financial and technical bids, each consisting of the documents listed in ITB Sub-Clause 9.1, shall be typed or written in indelible ink and shall be signed by the Bidder or a person or persons duly authorized to bind the Bidder to the contract. The letter of authorization shall be indicated by written power of attorney, accompanying the bid and submitted as Attachment 2 to the Bid under ITB Sub-Clause 9.2. All pages of the bid, except for un- amended printed literature, shall be initialed by the person or persons signing the bid. 15.3 The bid shall contain no alterations, omissions, or additions, unless such corrections are initialed by the person or persons signing the bid. |

D.  Submission of Bids

| 16. Sealing and Marking of Bids | 16.1 The Bidder shall seal the technical and financial bids in separate envelopes, duly marking the envelopes as “Financial Proposal” and “Technical Proposal”. The envelopes shall then be sealed in an outer envelope. 16.2 The inner and outer envelopes shall (a) be addressed to the Purchaser at the address given in the Bid Data Sheet, and (b) bears the contract for name indicated in the Bid Data Sheet, the Invitation for Bids title and number indicated in the Bid Data Sheet, and the statement. 16.3 The inner envelopes shall also indicate the name and address of the Bidder so that the bid can be returned unopened in case it is declared “late.” 16.4 If the outer envelope is not sealed and marked as required by ITB Sub-Clause 16.2 above, the Purchaser due to any reason will assume no responsibility for the bid’s misplacement or premature opening. 16.5 In Technical Proposal the bidder must provide the original Data sheet and Technical brochure, Income Tax, Sales Tax, Sole Agency certificate and other documents, otherwise the bid will be ignored - In Financial Proposal, the bidder should provide bid security / earnest money at the rate of 2.5% of the quoted value. A photocopy of the same should be attached with Technical Proposal but its value must not be disclosed and this can be achieved by placing a piece of paper on the place where amount is mentioned. |
| --- | --- |
| 17. Deadline for Submission of Bids | 17.1 Bids must be received by the Purchaser at the address specified under ITB Sub-Clause 16.2 no later than the time and date stated in the invitation of bids. 17.2 The Purchaser may, at its discretion, extend this deadline for submission of bids. |
| 18. Late Bids | 18.1 Any bid received by the Purchaser after the bid submission deadline prescribed by the Purchaser, will be rejected and returned unopened to the Bidder. |
| 19. Modification and Withdrawal of Bids | 19.1 The Bidder may withdraw its bid after submission, prior to the deadline prescribed for bid submission. 19.2 No bid may be withdrawn in the interval between the bid submission deadline and the expiration of the bid validity period specified. Withdrawal of a bid during this interval may result in the Bidder’s forfeiture of its bid security. |

E.  Bid Opening and Evaluation

| 20. Opening of Bids by Purchaser | 20.1 The Purchaser will open all bids, in the presence of authorized bidders’ representatives who choose to attend, at the time, on the date and the place specified in the Invitation of Bids The bidders’ representatives who are present shall sign the Attendance Sheet evidencing their attendance. 20.2 The bidders’ names, and the presence or absence of bid security, and any such other details as the Purchaser at its discretion, may consider appropriate will be announced at the opening. No bid shall be rejected at bid opening, except for late bids, which shall be returned unopened to the Bidder. 20.3 The Purchaser will prepare minutes of the bid opening. 20.4 The bids found having without Bid security (earnest money) will also be returned the bidders. However, prior to return to the bidder, the chairman of the purchase committee will record statement /reason on such bids. Bids that are not opened and read out / announced at bid opening shall not be considered further for evaluation, irrespective of the circumstances. |
| --- | --- |
| 21. Clarification of Bids | 21.1 During bid evaluation, the Purchaser may, at its discretion, ask the Bidder for a clarification of its bid. The request for clarification and the response shall be in writing. |
| 22. Preliminary Examination of Bids | 22.1 The Purchaser will examine the bids to determine whether they are complete, whether any computational errors have been made, whether required securities have been furnished, whether the documents have been properly signed, and whether the bids are generally in order. 22.2 Bids shall be submitted either by the manufacturer or its authorized contractor, if submitted by the manufacturer itself then bid of authorized contractor will be rejected. 22.3 Arithmetical errors will be rectified on the following basis. If there is a discrepancy between the unit price and the total price, which is obtained by multiplying the unit price and quantity, or between subtotals and the total price, the unit or subtotal price shall prevail, and the total price shall be corrected. If there is a discrepancy between words and figures, the amount in words will prevail. If the Bidder does not accept the correction of errors, its bid will be rejected. 22.4 Prior to the detailed evaluation, the Purchaser will determine whether each bid is of acceptable quality, is complete and is substantially responsive to the bidding documents. For purposes of this determination, a substantially responsive bid is one that conforms to all the terms, conditions and specifications of the bidding documents without material deviations, objections, conditions or reservations. A material deviation, objection, conditionality or reservation is one (i) that affects in any substantial way the scope, quality or performance of the contract; (ii) that limits in any substantial way, inconsistent with the bidding documents, the Purchaser’s rights or the successful Bidder’s obligations under the contract; or (iii) whose rectification would unfairly affect the competitive position of other bidders who are presenting substantially responsive bids. 22.5 Conditional bid(s) will not be entertained. |
| 23. Conversion to Single Currency | 23.1 To facilitate evaluation and comparison, the Purchaser will convert all bid prices expressed in the amounts in to - Currency of Pakistan at the selling exchange rate established for similar transactions by the central bank or by a commercial bank in Pakistan. |
| 24. Technical Evaluation | 24.1 The Purchaser will carry out a detailed evaluation of the bids in order to determine whether the technical aspects are in accordance with the requirements set forth in the bidding documents on the basis of the information supplied by the bidders, taking into account the following factors: (a) Overall completeness and compliance with the Technical Specifications. (b) Type, quantity and long-term availability of mandatory and recommended spare parts and maintenance services (c) Any other relevant factors, if any, listed in the Bid Data Sheet, or that the Purchaser deems necessary or prudent to take into consideration. |
| 25. Financial Evaluation | 25.1 Pursuant to ITB Sub-Clause, the following evaluation methods will be followed: (a) Time schedule (program of performance) The plant and equipment covered by this bidding are required to be shipped and installed, and the facilities shall have the pre-commissioning completed within the period named in the Bid Data Sheet after the effective date specified in the Form of Contract Agreement. Bids offering a completion date beyond the maximum designated period shall be rejected. |
| 26. Contacting the Purchaser | 26.1 From the time of bid opening to the time of contract award, if any Bidder wishes to contact the Purchaser on any matter related to its bid, it should do so in writing. 26.2 Any effort by a Bidder to influence the Purchaser in the Purchaser’s bid evaluation, bid comparison or contract award decisions may result in rejection of the Bidder’s bid. |

F.  Award of Contract

| 27. Post-qualification | 27.1 In the absence of pre-qualification, the Purchaser will determine to its satisfaction whether the Bidder selected as having submitted the lowest evaluated responsive bid is qualified to satisfactorily perform the contract. 27.2 The determination will take into account the Bidder’s financial, technical and production capabilities, in particular Bidder’s contract work in hand, future commitments and current litigation. It will be based upon an examination of the documentary evidence of the Bidder’s qualifications submitted by the Bidder in Attachment 3 to the bid, as well as such other information as the Purchaser deems necessary and appropriate. 27.3 An affirmative determination will be a prerequisite for award of the contract to the Bidder. A negative determination will result in rejection of the Bidder’s bid, in which event the Purchaser will proceed to the next lowest evaluated bid to make a similar determination of that Bidder’s capabilities to perform satisfactorily. 27.4 The Purchaser, at any stage of the procurement proceedings, having credible reasons for or prima facie evidence of any defect in capacities, may require the bidder to provide information concerning their professional, technical, financial, legal or managerial competence whether already pre-qualified or not. 27.5 The Purchaser, will disqualify a bidder if it finds, at any time that the information submitted by him concerning his qualification was false and materially inaccurate or incomplete. |
| --- | --- |
| 28. Award Criteria | - Subject to ITB Clause 29, the Purchaser will award the contract to the successful Bidder whose bid has been determined to be substantially responsive and to be the lowest evaluated bid, further provided that the Bidder is determined to be qualified to perform the contract satisfactorily. |
| 29. Purchaser’s Right to Accept Any Bid and to Reject Any or All Bids | 29.1 Under relevant provisions of SPP Rules the Purchaser reserves the right to accept or reject any bid, and to annul the bidding process and reject all bids at any time prior to award of contract without thereby incurring any liability to the affected Bidder or bidders or any obligation to inform the affected Bidder or bidders of the grounds for the Purchaser’s action. |
| 30. Notification of Award | - Prior to the expiration of the period of bid validity, the Purchaser will notify the successful Bidder in writing by registered letter that its bid has been accepted. The notification of award will constitute the formation of the contract. 30.2 Upon the successful Bidder’s furnishing of the performance security the Purchaser will promptly notify each unsuccessful Bidder and will discharge its bid security. |
| 31. Signing the Contract Agreement | 31.1 At the same time as the Purchaser notifies the successful Bidder that its bid has been accepted, the Purchaser will send the Bidder the Contract Agreement provided in the bidding documents, incorporating all agreements between the parties. 31.2 Within fifteen (15) days of receipt of the Contract Agreement, the successful Bidder shall sign the Contract Agreement and return it to the Purchaser. |
| 32. Performance Security | 32.1 Within fifteen (15) days after receipt of the notification of award, the successful Bidder shall furnish the performance security in the amount given in the Bid Data Sheet. 32.2 Failure of the successful Bidder to comply with the requirements of ITB Clause 32 or Clause 33 shall constitute sufficient grounds for the annulment of the award and forfeiture of the bid security, in which event the Purchaser may make the award to the next lowest evaluated Bidder or call for new bids. |

| 33.Corrupt or Fraudulent Practices | 33.1 The Purchaser as well as bidders/Suppliers/Contractors observes the highest standard of ethics during the procurement and execution of such contracts. In pursuance of this policy, the Purchaser (a) defines, for the purposes of this provision, the terms set forth below as follows: (i) “corrupt practice” means the offering, giving, receiving or soliciting of anything of value to influence the action of a public official in the procurement process or in contract execution; and (ii) “fraudulent practice” means a misrepresentation of facts in order to influence a procurement process or the execution of a contract to the detriment of the Borrower, and includes collusive practice among bidders (prior to or after bid submission) designed to establish bid prices at artificial non-competitive levels and to deprive the Borrower of the benefits of free and open competition; (b) will reject a proposal for award if it determines that the Bidder recommended for award has engaged in corrupt or fraudulent practices in competing for the contract in question; - will cancel the proposal and recover from the successful Bidders the amount of any loss resulting from such cancellation if the successful Bidder and person employed by the successful Bidder or acting on its behalf, offer or gives to any person any gift or consideration of any kind as an inducement or reward for doing, or forbearing to do , any action in the relation to obtaining or in the execution of the concession or any other contract with the purchaser if it at any time determine that the Bidder has , directly or through an agent, engaged in corrupt, fraudulent, collusive or coercive practices. |
| --- | --- |

# Section III.  Bid Data Sheet

The following bid-specific data for the plant and equipment to be procured shall amend and/or supplement the provisions in the Instructions to Bidders (ITB).  Whenever there is a conflict, the provisions herein shall prevail over those in the ITB.

| INTRODUCTION |  |
| --- | --- |
| ITB 1.1 | Nuclear Institute Of Medicine And Radiotherapy (NIMRA) Jamshoro |
| ITB 1.1 | Name of Project: Provision of Purchase of Linear Accelerator, MRI System and C.T. Simulator to be installed at Nuclear Institute Of Medicine And Radiotherapy (NIMRA) Jamshoro |
| ITB 2.1 | Name of Contractor _________________________________ __________________________________________________ |
| ITB 2.1 | Name of Purchaser: Health Department, Government of Sindh |
| ITB 6.1 | Address of Purchaser; Nuclear Institute of Medicine And Radiotherapy (NIMRA) Jamshoro through Health Department, Government of Sindh |
| ITB 8.1 | Language of bid is English |
| ITB 9.2 (c) | Qualification requirements for bidders Size of operation The Bidder has an average annual turnover RS. 100.00 Millions Turnover in Pak rupees (defined as total payment received by the Supplier for the contracts completed or under execution) over the three (03) years period throughout world. Technical Experience: - For at least last three years minimum of 20 units of similar capacity equipment has been sold in Pakistan / throughout the world. - Department reserves right to inspect/visit the locations of quoted machinery/equipments installed in Pakistan and Sindh Financial Position Soundness of the Bidder’s financial position showing long-term profitability demonstrated through audited annual financial statement (balance sheet, income statement) for the last three (03) year. Technical Staff / Workshop The bidder shall furnish certificate from the sole agent / manufacturer that the contractor has fully equipped workshop along with list of trained personnel / Engineers and equipments with the technical bid, same can be physically verified by the purchase committee at any time furnish complete details of trained engineer / Technical staff and well equipped work shop located in Pakistan especially at Sindh province for after sale & service conformation Litigation History. The Bidder’s should not be involved in any litigation with the Government / semi Government in Pakistan. |
| ITB 9.2 (d) (i) | Spares required for operation; number of years following completion |
| BID PRICE AND CURRENCY |  |
| ITB 11.3 | The price shall be fixed |
| ITB 12.1 | ITB Sub-Clause 12.1 shall apply to all plant and equipment and all service referred. |
| BID PREPARATION AND SUBMISSION |  |
| ITB 13.1 | Amount of bid security Equivalent to 21/2 % of the total value of bid. |
| ITB 14.1 | Bid validity period 90 days from the date of bid opening |
| ITB 15.1 | Original “Financial and Technical” Proposal. |
| ITB 17.1 | Address of bid submission “as ITB 20.1” Deadline for bid submission 04-05-2016 upto 12:00 noon. |
| ITB 20.1 | Location, time and date of bid opening In the Office of the I.I. Depot, Near JPMC, Rafique Shaheed Road, Karachi : 04-05-2016, at 01.00 p.m. |

| BID EVALUATION |  |
| --- | --- |
| ITB 23.1 | Currency chosen for the purpose of converting to a common currency Pak Rupees Date of Exchange Rate: The original deadline for receiving of Bid. |
| ITB 24.1 (c) | Optional: The following additional criteria will be used in the evaluation in the manner stated and to the extent indicated in the Bid Data Sheet. As mentioned at Technical Specifications |
| ITB 25.1 (a) | Time schedule Time to complete the facilities from the effective date specified in Article 3 of the Contract Agreement for determining time for completion of pre commissioning activities shall be three (03) months maximum. The adjustment rate in the event of completion beyond the maximum period shall be 0.5% for each week of delay from that maximum period. |
| CONTRACT AWARD |  |
| ITB 32 | Amount of performance security 10% of the total contract price in the form of Bank guarantee or Pay order / Bank draft. |

# Section IV.  General Conditions of Contract

## Table of Clauses

A.  Contract and Interpretation	24

1.	Definitions	24

2.	Contract Documents	25

3.  Interpretation	26

4.	Notices	27

5.	Governing Law	27

6.	Settlement of Disputes	28

B.  Subject Matter of Contract	29

7.	Scope of Facilities	29

8.	Time for Commencement and Completion	29

9.	Contractor’s Responsibilities	29

10.	Purchaser’s Responsibilities	30

C.  Payment	31

11.	Contract Price	31

12.	Terms of Payment	31

13.	Securities	31

14.	Taxes and Duties	32

E.   Execution of the Facilities	32

15.	Work Program	32

16.	Design and Engineering	33

17.	Procurement	33

18.	Installation	34

19.       Preliminary Inspection	35

20.	Completion of the Facilities	35

21.Commissioning and Operational Acceptance	35

F.  Guarantees and Liabilities	35

22.	Completion Time Guarantee	35

23.	Defect Liability	36

24.	Functional Guarantees	36

25.	Force Majeure	36

26.	Extension of Time for Completion	37

27.	Termination for Default	37

General Conditions of Contract

A.  Contract and Interpretation

| 1. Definitions | 1.1 The following words and expressions shall have the meanings hereby assigned to them: “Contract” means the Contract Agreement entered into between the Purchaser and the Contractor, together with the Contract Documents referred to therein; they shall constitute the Contract, and the term “the Contract” shall in all such documents be construed accordingly. “GCC” means the General Conditions of Contract hereof. “SCC” means the Special Conditions of Contract. “Day” means calendar day “Month” means calendar month “Purchaser” means the person named as such in the SCC. “Contractor” means the person(s) whose bid to perform the Contract has been accepted by the Purchaser and is named as such in the Contract Agreement, and includes the legal successors or permitted assigns of the Contractor. “Contract Price” means the sum specified in the Contract Agreement, as may be made pursuant to the Contract. “Facilities” means the Plant and Equipment to be supplied and installed, as well as all the Installation Services to be carried out by the Contractor under the Contract. “Plant and Equipment” means permanent plant, equipment, machinery, apparatus, articles and things of all kinds to be provided and incorporated in the Facilities by the Contractor under the Contract “Installation Services” means all those services ancillary to the supply of the Plant and Equipment for the Facilities, to be provided by the Contractor under the Contract; e.g. transportation and provision of marine or other similar insurance, inspection, expediting, site preparation works (including the provision and use of Contractor’s Equipment and the supply of all construction materials required), installation, testing, pre-commissioning, commissioning, operations, maintenance, the provision of operations and maintenance manuals, training, etc. “Country of Origin” means the countries and territories eligible under the rules “Site” means the land and other places upon which the Facilities are to be installed, and such other land or places as may be specified in the Contract as forming part of the Site. “Time for Completion” means the time within which Completion of the Facilities as a whole is to be attained in accordance with the specifications in the SCC and the relevant provisions of the Contract. “Completion” means that the Facilities have been completed operationally and structurally and put in a tight and clean condition, and that all work in respect of Pre-commissioning of the Facilities or such specific part thereof has been completed; in other words, that the Facilities or specific part thereof are ready for Commissioning. “Pre-commissioning” means the testing, checking and other requirements specified in the Technical Specifications that are to be carried out by the Contractor in preparation for Commissioning. “Commissioning” means operation of the Facilities which operation is to be carried out by the Contractor |
| --- | --- |
| 2. Contract Documents | 2.1 All documents forming part of the Contract (and all parts thereof) are intended to be correlative, complementary and mutually explanatory. The Contract shall be read as a whole. |
| 3. Interpretation | 3.1 Language 3.1.1 All Contract Documents, all correspondence and communications to be given and all other documentation to be prepared and supplied under the Contract shall be written in English. 3.1.2 If any of the Contract Documents, correspondence or communications are prepared in any language other than the governing language under GCC Sub-Clause 3.1.1 above, the English translation of such documents, correspondence or communications shall prevail in matters of interpretation. 3.2 Headings The headings and marginal notes in the General Conditions of Contract are included for ease of reference, and shall neither constitute a part of the Contract nor affect its interpretation. 3.3 Persons Words importing persons or parties shall include firms, corporations, and government entities. 3.4 Amendment No amendment or other variation of the Contract shall be effective. 3.5 Independent Contractor The Contractor shall be an independent contractor performing the Contract. The Contract does not create any agency, partnership, joint venture or other joint relationship between the parties hereto. Subject to the provisions of the Contract, the Contractor shall be solely responsible for the manner in which the Contract is performed 3.6 Country of Origin “Origin” means the place where the materials, equipment and other supplies for the Facilities are mined, grown, produced or manufactured, and from which the services are provided. |
| 4. Notices | 4.1 Unless otherwise stated in the Contract, all notices to be given under the Contract shall be in writing, and shall be sent by personal delivery, airmail post, special courier, telegraph, telex, facsimile to the address of the relevant party set out in the SCC, with the following provisions: 4.1.1 Any notice sent by telegraph, telex, facsimile shall be confirmed within two (2) days after dispatch by notice sent by airmail post or special courier, except as otherwise specified in the Contract. 4.1.2 Any notice sent by airmail post or special courier shall be deemed (in the absence of evidence of earlier receipt) to have been delivered ten (10) days after dispatch. In proving the fact of dispatch, it shall be sufficient to show that the envelope containing such notice was properly addressed, stamped and conveyed to the postal authorities or courier service for transmission by airmail or special courier. 4.1.3 Any notice delivered personally or sent by telegraph, telex, facsimile shall be deemed to have been delivered on date of its dispatch. - Either party may change its postal, telex, facsimile address or addressee for receipt of such notices by ten (10) days’ notice to the other party in writing. 4.2 Notices shall be deemed to include any approvals, consents, instructions, orders and certificates to be given under the Contract. |
| 5. Governing Law | 5.1 The Contract shall be governed by and interpreted in accordance with laws of the country specified in the SCC. |
| 6. Settlement of Disputes | 6.1 Adjudicator 6.1.1 If any dispute of any kind whatsoever shall arise between the Purchaser and the Bidder in connection with or arising out of the Contract, including without prejudice to the generality of the foregoing, any question regarding its existence, validity or termination, or the execution of the Facilities whether during the progress of the Facilities or after their completion and whether before or after the termination, abandonment or breach of the Contract the parties shall seek to resolve any such dispute or difference by mutual consultation. If the parties fail to resolve such a dispute or difference by mutual consultation, then the dispute shall be referred in writing by either party to the Adjudicator, with a copy to the other party. 6.1.2 The Adjudicator shall give its decision in writing to both parties within twenty-eight (28) days of a dispute being referred to it. 6.2 If, after thirty (30) days from the commencement of such informal negotiation, the Purchaser and the supplier have been unable to resolve amicably a Contract dispute, either party may require that the dispute be referred to the arbitrator for resolution through arbitration. 6.3 In case of any dispute concerning the interpretation and application of contract that shall be referred to the Law Department, Govt. of Sindh, who or his nominee shall act as sole arbitrator. The award rendered by the sole arbitrator shall be final and binding on the Parties |

B.  Subject Matter of Contract

| 7. Scope of Facilities | 7.1 Unless otherwise expressly limited in the Technical Specifications, the Contractor’s obligations cover the provision of all Plant and Equipment and the performance of all Installation Services. required for the design, the manufacture (including procurement, quality assurance, installation, Pre-commissioning and delivery) of the Plant and Equipment and the installation, completion and commissioning of the Facilities in accordance with the plans, procedures, specifications, codes and any other documents as specified in the Technical Specifications. 7.2 The Contractor shall, unless specifically excluded in the Contract, perform all such work and/or supply all such items and materials not specifically mentioned in the Contract but that can be reasonably inferred from the Contract as being required for attaining Completion of the Facilities as if such work and/or items and materials were expressly mentioned in the Contract. 7.3 In addition to the supply of Mandatory Spare Parts included in the Contract, the Contractor agrees to supply spare parts required for the operation and maintenance of the Facilities for the period specified in the Special Conditions of Contract |
| --- | --- |
| 8. Time for Commencement and Completion | 8.1 The Contractor shall commence work on the Facilities within the period specified in the SCC and without prejudice to GCC Sub-Clause 22.2 hereof, the Contractor shall thereafter proceed with the Facilities in accordance with the time schedule specified. |
| 9. Contractor’s Responsibilities | 9.1 The Contractor shall design, manufacture (including associated purchases and/or subcontracting), install and complete the Facilities with due care and diligence in accordance with the Contract. 9.2 The Contractor confirms that it has entered into this Contract on the basis of a proper examination of the data relating to the Facilities provided by the Purchaser, and on the basis of information that the Contractor could have obtained from a visual inspection of the Site. The Contractor acknowledges that any failure to acquaint itself with all such data and information shall not relieve its responsibility for properly estimating the difficulty or cost of successfully performing the Facilities. 9.3 The Contractor shall acquire in its name all permits, approvals or licenses from all local, state or national government authorities or public service undertakings in the country where the Site is located that are necessary for the performance of the Contract, including, without limitation, visas for the Contractor’s and Subcontractor’s personnel and entry permits for all imported Equipment. The Contractor shall acquire all other permits, approvals and/or licenses that are not the responsibility of the Purchaser and that are necessary for the performance of the Contract. 9.4 The Contractor shall comply with all laws in force in the country where the Facilities are installed and where the Installation Services are carried out. The laws will include all local, state, national or other laws that affect the performance of the Contract and bind upon the Contractor. The Contractor shall indemnify and hold harmless the Purchaser from and against any and all liabilities, damages, claims, fines, penalties and expenses of whatever nature arising or resulting from the violation of such laws by the Contractor or its personnel, including the Subcontractors and their personnel, but without prejudice to GCC Sub-Clause 10.1 hereof. |
| 10. Purchaser’s Responsibilities | 10.1 To submit the manufacturers undertaking that if his authorized distributor / contractor fail to carry out any assignment in total or in part, manufacturer will be responsible to carry out the same along with all the terms and conditions upto the warranty period of the bid. 10.2 The Purchaser shall ensure the accuracy of all information and/or data to be supplied by the Purchaser to the Contractor. 10.3 The Purchaser shall be responsible for acquiring and providing legal and physical possession of the Site and access thereto, and for providing possession of and access to all other areas reasonably required for the proper execution of the Contract, including all requisite rights of way, as specified. The Purchaser shall give full possession of and accord all rights of access thereto on or before the date(s) specified. 10.4 If requested by the Contractor, the Purchaser shall use its best endeavors to assist the Contractor in obtaining in a timely and expeditious manner all permits, approvals and/or licenses necessary for the execution of the Contract from all local, state or national government authorities or public service undertakings that such authorities or undertakings require the Contractor or Subcontractors or the personnel of the Contractor or Subcontractors, as the case may be, to obtain. |

C.  Payment

| 11. Contract Price | 11.1 The Contract Price shall be as specified in Article 2 (Contract Price) of the Form of Contract Agreement. 11.2 Unless indicated otherwise in the SCC, the Contract Price shall be a firm lump sum not subject to any alteration, except in the event of a Change in the Facilities or as otherwise provided in the Contract. |
| --- | --- |
| 12. Terms of Payment | 12.1 The Contract Price shall be paid as specified in the Contract Agreement. The procedures to be followed in making application for and processing payments shall be finalized at the time of Contract. 12.2 The currency or currencies in which payments are made to the Contractor under this Contract shall be specified in the Contract Agreement, subject to the general principle that payments will be made in the currency or currencies in which the Contract Price has been stated. |
| 13. Securities | 13.1 Issuance of Securities The Contractor shall provide the securities specified below in favor of the Purchaser at the times, and in the amount, manner and form specified below. 13.2 Performance Security 13.2.1 The Contractor shall, within fifteen (15) days of the notification of contract award, provide a security for the due performance of the Contract in the amount specified in the SCC. 13.2.2 The security shall be denominated in the currency or currencies of the Contract, or in a freely convertible currency acceptable to the Purchaser, and shall be in one of the forms, pay order / bank draft or bank guarantees. |
| 14. Taxes and Duties | 14.1 Except as otherwise specifically provided in the Contract, the Contractor shall bear and pay all taxes, duties, levies and charges assessed on the Contractor, its Subcontractors or their employees by all municipal, state or national government authorities in connection with the Facilities in and outside of the country where the Site is located. 14.2 If any tax exemptions, reductions, allowances or privileges may be available to the Contractor in the country where the Site is located, the Purchaser shall use its best endeavors to enable the Contractor to benefit from any such tax savings to the maximum allowable extent. 14.3 For the purpose of the Contract, it is agreed that the Contract Price specified in Article 2 of the Form of Contract Agreement is based on the taxes, duties, levies and charges prevailing at the date of bid submission in the country where the Site is located. If any rates of Tax are increased or decreased, a new Tax is introduced, an existing Tax is abolished, or any change in interpretation or application of any Tax occurs in the course of the performance of Contract, which was or will be assessed on the Contractor. |

E.   Execution of the Facilities

| 15. Work Program | 15.1 Work Procedures The Contract shall be executed in accordance with the Contract Documents and procedures given in the section on Sample Forms and Procedures of the Contract Documents. The Contractor may execute the Contract in accordance with its own standard project execution plans and procedures to the extent that they do not conflict with the provisions contained in the Contract. |
| --- | --- |
| 16. Design and Engineering | 16.1 Specifications The Contractor shall execute the basic and detailed design and the engineering work in compliance with the provisions of the Contract, or where not so specified, in accordance with good engineering practice. |
| 17. Procurement | 17.1 Plant and Equipment The Contractor shall transport all the Plant and Equipment from the manufacturers in an expeditious and orderly manner to the Site. 17.2 Transportation. 17.2.1 The Contractor shall at its own risk and expense transport all the Plant and Equipment to the Site by the mode of transport that the Contractor judges most suitable under all the circumstances. 17.2.2 Unless otherwise provided in the Contract, the Contractor shall be entitled to select any safe mode of transport to carry the Plant and Equipment. 17.2.3 Upon dispatch of each shipment of the Plant and Equipment, the Contractor shall notify the Purchaser by telex, facsimile of the description of the Plant and Equipment, the point and means of dispatch, and the estimated time and point of arrival in the country where the Site is located, if applicable, and at the Site. The Contractor shall furnish the Purchaser with relevant shipping documents to be agreed upon between the parties. 17.2.4 The Contractor shall be responsible for obtaining, if necessary, approvals from the authorities for transportation of the Plant and Equipment to the Site. The Purchaser shall use its best endeavors in a timely and expeditious manner to assist the Contractor in obtaining such approvals, if requested by the Contractor. The Contractor shall indemnify and hold harmless the Purchaser from and against any claim for damage that may be caused by the transport of the Plant and Equipment to the Site. 17.3 Customs Clearance The Contractor shall, at its own expense, handle all imported Plant and Equipment and Contractor’s Equipment at the point(s) of import and shall handle any formalities for customs clearance. |
| 18. Installation | 18.1 Setting Out/Supervision/Labor 18.1.1 Bench Mark: The Contractor shall be responsible for the true and proper setting-out of the Facilities in relation to benchmarks, reference marks and lines provided to it in writing by or on behalf of the Purchaser. If, at any time during the progress of installation of the Facilities, any error shall appear in the position, level or alignment of the Facilities, the Contractor shall forthwith notify the Purchaser of such error and, at its own expense, immediately rectify such error to the reasonable satisfaction of Purchaser. 18.1.2 Contractor’s Supervision: The Contractor shall give or provide all necessary superintendence during the installation of the Facilities, and the Construction Manager or its deputy shall be constantly on the Site to provide full-time superintendence of the installation. The Contractor shall provide and employ only technical personnel who are skilled and experienced in their respective callings and supervisory staff who are competent to adequately supervise the work at hand. 18.1.3Labor: The Contractor shall provide & employ on the Site in the installation of the Facilities such skilled, semi-skilled and unskilled labor as is necessary for the proper and timely execution of the Contract. |
| 19. Preliminary Inspection | 19.1 The Purchaser or its representative shall have the right to preliminary inspection of the Plant & Machinery before shipment to confirm their conformity to the Contract specification 19.2 The inspection will be conducted at the premises of the Manufacturer, at the time of shipment, all reasonable facilities and assistance shall be furnished to the inspectors at no charge to the Purchaser. |
| 20. Completion of the Facilities | 20.1 As soon as the Facilities or any part thereof has, in the opinion of the Contractor, been completed operationally and structurally and put in a tight and clean condition as specified in the Technical Specifications, excluding minor items not materially affecting the operation or safety of the Facilities, the Contractor shall so notify the Purchaser in writing. |
| 21.Commissioning and Operational Acceptance | 21.1 Commissioning Commissioning of the Facilities or any part thereof shall be commenced by the Contractor immediately after issue of the Completion Certificate by the Purchaser. |

F.  Guarantees and Liabilities

| 22. Completion Time Guarantee | 22.1 The Contractor guarantees that it shall attain Completion of the Facilities within the Time for Completion specified in the SCC 22.2 If the Contractor fails to attain Completion of the Facilities within the Time for Completion or any extension thereof under GCC, the Contractor shall pay to the Purchaser liquidated damages in the amount specified in the SCC as a percentage rate of the Contract Price, or the relevant part thereof. The aggregate amount of such liquidated damages shall in no event exceed the amount specified as “Maximum” in the SCC. Once the “Maximum” is reached, the Purchaser may consider termination of the Contract. Such payment shall completely satisfy the Contractor’s obligation to attain Completion of the Facilities within the Time for Completion or any extension thereof under GCC. The Contractor shall have no further liability whatsoever to the Purchaser in respect thereof. However, the payment of liquidated damages shall not in any way relieve the Contractor from any of its obligations to complete the Facilities or from any other obligations and liabilities of the Contractor under the Contract. |
| --- | --- |
| 23. Defect Liability | - The Contractor warrants that the Facilities or any part thereof shall be free from defects in the design, engineering, materials, and workmanship of the Plant and Equipment supplied and of the work executed. - In case of defect in any part at the time of installation it shall be replaced with new one instead of repair. |
| 24. Functional Guarantees | 24.1 The Contractor guarantees that during the Guarantee Test, the Facilities and all parts thereof shall attain the Functional, subject to and as specified in Technical Specification. |
| 25. Force Majeure | 25.1 Not with-standing the provisions of GCC Clauses 16,17 & 18 the supplier shall not be liable for forfeiture of its performance security / earnest money, or termination / blacklisting for default if and to the extent that it delay in performance or other failure to perform its obligation under the contract is the result of an event of Force Majeure means an act of God or an event beyond the control of the supplier and not involving the supplier’s fault or negligence directly or indirectly purporting to mis-planning, mismanagement and / or lack of foresight to handle the situation. Such events may include but are not restricted to acts of the purchaser in its sovereign capacity, wars or revolutions, fires, floods earthquakes. Epidemics, quarantine restrictions and freight embargoes. If a Force Majeure situation arises, the supplier shall promptly notify the purchaser in writing with sufficient and valid evidence of such condition and the cause thereof. The Force Majeure Committee will examine the pros and cons of the case and all reasonable alternative means for completion of supply orders under this contract and will submit its recommendations to the competent authority. However, unless otherwise directed by the purchaser in writing, the supplier shall continue to perform its obligations under the contract as far as is reasonably practical and shall seek reasonable alternative means for performance not prevented by the Force Majeure event. |

H.  Change in Contract Elements

| 26. Extension of Time for Completion | 26.1 The Time(s) for Completion specified in the SCC shall be extended if the Contractor is delayed or impeded in the performance of any of its obligations under the Contract by reason of any of the following: (a) any Change in the Facilities as provided in GCC (b) any occurrence of Force Majeure as provided in GCC unforeseen conditions or other occurrence of any of the matters (c) any other matter specifically mentioned in the Contract |
| --- | --- |
| 27. Termination for Default | 27.1 The purchaser, without prejudice to any other remedy for breach of contract by written notice of default sent to the supplier, may terminate this contract in whole or in part: - If the supplier fails to deliver any or all of the Plant and Machinery within the period(s) specified in the Contract, or within any extension thereof granted by the Purchaser. - If the Supplier fails to perform any other obligation(s) under the Contract. - If the supplier, in the judgment of the Purchaser has engaged in corrupt or fraudulent practices in competing for or in executing the Contract. For the purpose of this clause: “corrupt practice” means the offering, giving, receiving or soliciting of any thing of value to influence the action of a public official in the procurement process or in contract execution “Fraudulent Practice” means a misrepresentation of facts in order to influence a procurement process or the execution of a contract to the detriment of the Purchaser, and includes collusive practice among bidders (prior to or after bid submission) designed to establish bid prices at artificial non-competitive levels and to deprive the Purchase of the benefits of free and open competition. |

# Section V.  Special Conditions of Contract

## Table of Clauses

1.	Definitions (GCC Clause 1)	40

2.	Notices (GCC Clause 4)	40

3.	Governing Law (GCC Clause 5)	40

4.	Scope of Facilities [Spare Parts] (GCC Clause 7)	40

5.	Time for Commencement and Completion	41

6. 	Contract Price (GCC Clause 11)	41

7.	Securities (GCC Clause 13)	41

8.	Completion Time Guarantee (GCC Clause 22)	41

9.	Specification …………………………………………………………………37

Special Conditions of Contract

The following Special Conditions of Contract (SCC) shall supplement the General Conditions of Contract (GCC).  Whenever there is a conflict, the provisions herein shall prevail over those in the GCC.  The corresponding clause number of the GCC is indicated in parentheses.

1.	Definitions (GCC Clause 1)

The Purchaser is: Nuclear Institute Of Medicine And Radiotherapy (NIMRA) Jamshoro.

The Project Manager is: Nuclear Institute Of Medicine And Radiotherapy (NIMRA) Jamshoro

The Contractor is:  ___________________________________________________

Country of Origin:  __________________________________________________

2.	Notices (GCC Clause 4)

GCC 4.1—Purchaser’s address for notice purposes:

Nuclear Institute Of Medicine And Radiotherapy (NIMRA) Jamshoro

Contractor’s address for notice purposes_____________________

______________________________________________________

3.	Governing Law (GCC Clause 5)

Sample Provision

GCC 5.1—The Contract shall be interpreted in accordance with the laws of the Purchaser’s country.

4.	Scope of Facilities [Spare Parts] (GCC Clause 7)

GCC 7.3—The Contractor agrees to supply spare parts for a period of 10 (Ten) years.

The Contractor shall carry sufficient inventories to ensure an ex-stock supply of consumable spares for the plant and equipment.  Other spare parts and components shall be supplied as promptly as possible, but at the most within three (3) months of placing the order and opening the letter of credit.  In addition, in the event of termination of the production of spare parts, advance notification will be made to the Purchaser of the pending termination, with sufficient time to permit the Purchaser to procure the needed requirement.  Following such termination, the Contractor will furnish to the extent possible and at no cost to the Purchaser the blueprints, drawings and specifications of the spare parts, if requested.

5.	Time for Commencement and Completion

The Contractor shall commence work on the Facilities from the Effective Date for determining Time for Completion as specified in the Contract Agreement.

The Completion of the Facilities shall be attained within three (03) months

6. 	Contract Price (GCC Clause 11)

The Contract price shall be fixed

7.	Securities (GCC Clause 13)

GCC 13.2.1— Amount of performance security, as a percentage 10% (Ten Percent) of the Contract Price for the Facility is provided, shall be:

GCC 13.2.2— The performance security shall be in the form of the: Pay order /  Bank Draft / Bank Guarantee  attached hereto in the section on Sample Forms and Procedures.

8.	Completion Time Guarantee (GCC Clause 22)

GCC 22.2—Applicable rate for liquidated damages:

Maximum deduction for liquidated damages: 0.5% per week. (Maximum up to 10%)

The above rate and maximum apply to the price of the part of the Facilities, as quoted in the Price Schedule, for that part for which the Contractor fails to achieve Completion within the particular Time for Completion.

# Section VI. Schedule of Requirements / Technical Specification

| Item No. | Description of Goods | Quantity |
| --- | --- | --- |
| 01 | Linear Accelerator | 01 Unit |
| 02 | MRI System | 01 Unit |
| 03 | C.T. Simulator | 01 Unit. |

Radiotherapy Linear Accelerator

| Sr #. | REQUIRED SPECIFICATION |
| --- | --- |
|  | The dual mode medical linear accelerator with dual photon and five electron energies, should be fully digital and further upgradable for IMRT IGRT cone beam CT Arc Therapy possible. Technical specifications are as follows: |
|  | - Country of Origin: |
|  | US, EU, Japan system should be FDA 510K/ CE approved |
|  | - 05 year Comprehensive warranty of the system. 10 year warranty of RF source. (Subject to contract with vendor) 20 year warranty of wave guide. (Subject to contract with vendor) |
|  | - Reliability: |
|  | Must provide guarantee and provision for at least 95% up time during Warranty period |
|  | - Energies and Dose rate: |
|  | Photon: Dual Energy, 6 MV and 10 MV or more. Electron: 5 distinct Energies shall be 6,9,12,15,18 MeV. Dose rate should be selectable independent upon energy, but variable for total body irradiation, between 50/100 MU/min up to minimum of 600 MU/min in high dose rate photon mode and upto 1000 MU/min in high dose rate electron mode. Independent dose chambers for x-ray and electrons. Field flatness, symmetry, and uniformity should be specified and must be in compliance with IEC/AAPM/ICRU |
|  | - Mechanical: |
|  | Gantry : rotation of 360 degrees, with an isocentre accuracy of an atleast 1 mm, including digital displays of the gantry angle inside and outside the treatment room. Head (MLC): rotation of at least 270 deg. With an accuracy of at least 1 mm at the isocentric height should be at a convenient level and indicated by fixed wall mounted lasers. Isocentric height should be at a convenient level and indicated by fixed wall mounted lasers. Low noise operation is to be ensured through adequate sound proofing for patient comfort. A spherical radial resolution of less than or equal to 1mm is required for high precision. |
|  | - Wedges: |
|  | Cross plane wedges should be supplied to compliment the virtual wedge functionality. The machine should be capable of delivering variable wedge angles 15, 30, 45 and 60ₒ. Constraint wedge factor for virtual wedges. Automatic calculation for virtual wedge resumption in case of wedge treatment interruption. |
|  | - Electron Applicators: |
|  | Electron applicators should be supplied both circular and rectangular fields. Electron beam shaping system/trimmer to be supplied for irregular electron fields. |
|  | - Head/MLC: |
|  | Equipment must be fitted with an integrated Multi Leaf Collimator (MLC) with minimum120 leaves with physical leaf width to get the best available resolution (≤ 1 cm) at the isocentre. Field size s max 40x40 cm. Transmission of MLC/collimators should be in tolerance limit for protection of healthy tissues. The spatial resolution in the direction of movement of leaves should be better than 1 mm. Machine should be supplied with an accessory holder with the best clearance. Clearance from accessory holder to isocentre should be as large as possible to enable automatic delivery of complex treatment protocols. The penumbra of the field should not exceed 9mm in any direction measured in 10cm of water at 100cm FAD between the 20 & 80 % intensity points. Radiation leakage through and between the leaves as well as for closed leaves must not exceed 2% (please specify). Total leaf travel should be as much as possible (Please specify range of travel). The MLC shape must be displayed on the in-room and control room monitors. MLC must be integrated in Record and Verify system (R/V) and operate from a single keyboard shared with the Linac. Leaf adjustment should also be possible with the hand pendant in the treatment room. The movement of each jaw should be independent of the other jaws. Accessory holder able to accept all applicator and trays for beam limitation and be coded accordingly. Trays to be supplied with the system. |
|  | - Portal Imaging (EPI): |
|  | The system must be equipped with amorphous silicon, flat panel (Electronic Based) portal imaging module for patient setup and treatment verification. The portal imaging module should be able to send and receive DICOM Images and to be integrated into the Linac user console with single screen and keyboard operation and supply software on the console. The Amorphous silicon flat panel should have good resolution (please specify) The imaging system should have a high bit pixel depth (please specify) The spatial resolution should be min of 1024x768x8 bit or more The module should be automatically deployed and retracted and be incorporated into auto sequencing protocols. The portal imaging device should not cause any obstruction to the automatic treatment functionality and operation of the machine. |

|  | - Leakage Radiation: |
| --- | --- |
|  | The system must comply with the following limits or with the most recent IEC relevant limits whichever is better. |
|  | X-ray leakage in the patient plane should be less than 0.1% |
|  | X-ray leakage outside patient plane should be less than 0.1% |
|  | X-ray leakage transmission through the diaphragms should be less than 0.5% of the unattenuated beam. |
|  | - Patient Table: |
|  | The table should be able to withstand the weight of 200 kg or more. The table should be motor driven with lateral and longitudinal with variable speed motor. Tabletop should be Indexed carbon fiber, compatible for portal imaging, conformal therapy and intensity modulated radiation settings The couch top shall have side rails with attached universal clamps and immobilization straps. Patient support panels in the couch shall be provided to facilitate large posterior treatments at extended distances without moving the patient. Hand pendant for the table should be provided. Table should be controlled using hand pendant, through table panel control inside the room and from control console The table must be able to rotate around the isocentre (motorized and free movement; please specify rotation angle range) The table should have positional accuracy and rigidity. Treatment should be possible from any angle with no obstruction by the table. |
|  | - Cooling System: |
|  | Cooling system (with fall safe option) with all adequate cooling capacity as per linac unit requirement must be installed by the supplier/manufacturer (please give specifications including brand name and warranty time in case the unit is third party chiller unit). |
|  | - Treatment Planning System |
|  | The Radiotherapy Treatment Planning System should be based on advanced computer system having latest hardware and software to perform all kinds of 3D Teletherapy and virtual simulation. One server for data base management integrated with the planning terminals. TPS with two work stations (two calculation licenses) should be provided for conventional photons, electrons, 3D CRT and inverse planning ( DICOM -3RT interface to network with simulator, CT scanner, dosimetry system with flat-bed scanner) with at least one license that includes IMRT. It should include fusion and registration software, virtual simulation software and beam modeling software. Also it should have real time DRR / DCR functionality, and plan approval facility and should provide Multiple Algorithms (Monte carlo, etc). A minimum of 3 contouring licenses / work stations should be included. |

|  | - Networking: |
| --- | --- |
|  | Data network between 3D Treatment Planning Systems (available in the department) with DICOM and DICOM RT export and import licenses, Radiotherapy simulator, Portal image analysis system, CT Scanner, MRI and all the terminals supplied along with the accelerator. |
|  | - Hardware requirements: Latest i7 PC server computer, should include at least 4 terminals workstation with latest computer and 21" LCD monitor medical grade appropriate color printers should be included. |
|  | - Immobilization accessories:- Cost may be quoted separately Two sets of standard adult/pediatric immobilization accessories (head, neck, chest, thorax, pelvis) to be supplied compatible and index table with Linear accelerator table top |
|  | - Record & Verify System: compatible with LINAC & TPS |
|  | - Accessories: |
| 1 | The equipment must be supplied with complete accessory kit to enable the machine to treat all known clinical setups/procedures |
| 2 | Complete set of electron applicator. |
| 3 | All available shadow trays/holders should be offered |
| 4 | Shielding blocks tray should be accepted by the collimator. |
| 5 | Special applicators like Total Skin Electron applicator. |
| 6 | Set of mechanical front pointers for mechanical isocentre/ODI check |
| 7 | Accessory spares |
| 8 | Close circuit room cameras (Four) |
| 9 | Camera Monitors (Two) |
| 10 | Patient Intercom System |
|  | - Construction of bunker Room Shielding and motorizsed Door: |
| 1 | Complete construction of bunker with all aspects, motorized door to be quoted separately. |
| 2 | The supplier is expected to provide advice on the radiation shielding requirements of the room where the installation of Linac will take place. |
|  | - Un-interrupted Power Supply (UPS) and voltage stabilizer: |
|  | UPS (including batteries) and voltage stabilizer with overvoltage/surge protection as per linac power requirements with backup of 10 min at full load for whole LINAC systems including associated TPS, server etc. Resettable over-current breaker shall be fitted for protection. |
|  | - Laser Patient Setup system: |
|  | The system must be supplied with red laser system to ensure accurate patient setup reference to isocentre. Width of laser light at isocentre should be very fine (please specify) compatible with CT simulator. |
|  | - Safety Compliance |
|  | Compliance with safety requirements in the international basic safety standard for protection against Ionizing Radiation and for the Safety Of Radiation Sources (BSS) and the International Electro Technical Commission Standards (IECC) |

|  | - Accompanying Documents and Software |
| --- | --- |
|  | The accompanying document shall comply with BSS and IEC standards, The performance specifications and operating and maintenance instructions shall be provided in English language. The documents/software etc. shall include: |
|  | - Performance specifications |
|  | - Operating instructions and manual |
|  | - Preventive maintenance instructions - service manual. |
|  | - Schematics diagrams |
|  | - Any other software, if needed by the user/maintenance personnel. |
|  | - Service password(s) if any, should not be supplier specific, but should be users defined. |
|  | - Acceptance Tests - Supplier must include full installation, commissioning acceptance testing safety checks initial radiation survey, including photon and neutron leakage measurement. - All the equipment to be used in step cited above shall be provided by the supplier. Spare parts: Provide list and cost of all replaceable spare parts in the machine |
|  | - Out of Warranty service Shall provide copies of post warranty service contract that are available each of following conditions: - Full parts and labour for a 10 year term post warranty period (PWP) (Mention parts list ) - Only service service contract without parts for 10 years |
|  | - TRAINING, EDUCATION AND CLINICAL SUPPORT Two weeks factory site Engineer training and clinical educational courses for Oncologist and physicist, onsite as well as abroad, provided in order to let hospital staff smooth and comprehensive clinical start-up with new equipment and modern technologies. |

WARRANTY & POST WARRANTY AGREEMENT & MAINTENANCE CONTRACT

WARRANTY PERIOD SHALL BE OF FIVE YEARS STARTING FROM DATE OF FULL FUNCTIONAL COMMISSIONING

MAINTENANCE OF THE MACHINE/EQUIPMENT DURING THE WHOLE PERIOD OF WARRANTY SHALL BE RESPONSIBILITY OF THE VENDOR FAILING WHICH, THE VENDOR WILL HAVE TO PAY DOWN TIME PENALTY AS INDICATED IN TERMS AND CONDITIONS

AN ANNUAL OPTIMAL UPTIME OF 95% IS CONSIDERED AS AN ACCEPTABLE LEVEL OF PERFORMANCE. UPTIME SHALL BE DEFINED AS TIME AVAILABLE TO USER FOR DATA ACQUISITION AND PROCESSING ON ALL WORKING DAYS THROUGHOUT THE YEAR

IF THE UPTIME PERCENTAGE FOR THE MEASUREMENT PERIOD SHALL FALL SHORT OF 95% ANNUALLY, THE FOLLOWING FORMULA WILL BE APPLIED TO DETERMINE ADDITIONAL DAYS IN THE WARRANTY/SERVICE CONTRACT PERIOD - 100 - 96% NO PENALTY - 95 - 90% THE WARRANTY PERIOD WILL BE EXTENDED BY 01 TIMES THE NUMBER OF DAYS AS EXTRA DOWN TIME. - 89-80% THE WARRANTY PERIOD WILL BE EXTENDED BY 02 TIMES THE NUMBER OF DAYS AS EXTRA DOWN TIME. - BELOW 80% THE WARRANTY PERIOD WILL BE EXTENDED BY 03 TIMES THE NUMBER OF DAYS AS EXTRA DOWN TIME.

DOWN TIME IS DEFINED AS THE FAILURE IN THE EQUIPMENT OPERATION TO ACQUIRE OR PROCESS THE DATA, RESULTING IN INABILITY TO CARRY OUT THE REQUIRED PROCEDURE PROPERLY.

DOWN TIME WILL START AS THE END USER NOTIFIED THE DESIGNATED SERVICE/FACILITY VERBALLY OR IN WRITING. THE RESPONSE TIME OF 12 TO 24 HOURS WILL BE CONSIDERED AS NORMAL DEPENDING UPON THE LOCATION OF UNIT.

DOWN TIME WILL END WHEN SATISFACTORY REPAIRS HAVE BEEN AFFECTED AND THE SYSTEM IS AGAIN FULLY FUNCTIONAL WITH ALL PARAMETERS AVAILABLE FOR CLINICAL USE.

THE SCHEDULED PREVENTIVE MAINTENANCE SHALL NOT BE CONSIDERED AS DOWN TIME AND SHALL NOT BE MORE THAN THREE DAYS

MANUFACTURER/SUPPLIER WILL BE RESPONSIBLE FOR PREVENTIVE MAINTENANCE OF THE SYSTEM AS PER SERVICE MANUAL.

VENDOR WILL MENTION THE RATE OF SERVICE CHARGES WITH AND WITHOUT PARTS FOR THE COMPLETE SYSTEM AFTER THE EXPIRY OF FIVE YEARS WARRANTY PERIOD ON YEARLY BASIS ALONG WITH THE BID. THEY SHOULD ALSO INCLUDE PRICE MAJOR BREAKABLE PARTS.

MANUFACTURER/SUPPLIER WILL GUARANTEE THE AVAILABILITY OF SPARE PARTS AND ACCESSORIES FOR THE SYSTEM FOR AT LEAST TEN YEARS.

MANUFACTURER/SUPPLIER SHOULD ALSO CONFIRM THAT IN EVENT OF THE SERVICES COSTS FAILING OR REDUCING IN LINE WITH THE MANUFACTURER POLICY WORLDWIDE DUE TO TECHNICAL INNOVATIONS OR OTHER REVIEWS, THEY WILL REDUCE THE MAINTENANCE COST ACCORDINGLY.

| TECHNICAL SPECIFICATIONS OF 0.35 TESLA OR MORE MRI SYSTEM |  |  |
| --- | --- | --- |
| LATEST GENERATION 0.35 OR MORE PERMANENT MAGNET MRI SYSTEM WITH CLINICAL APPLICATION SOFTWARE WITH COILS TO COVER THE WHOLE BODY. THE QUOTED MODEL MUST BE THE LATEST AND SUCCESSFUL MRI HAVING LARGE INSTALLED BASE AROUND THE WORLD. THE BIDDER MUST INDICATE THE NUMBER OF MRI SCANNERS INSTALLED WITHIN PAKISTAN AND LETTER OF SATISFACTION FROM ATLEAST THREE HOPISTALS. COMPANY SHOULD PROVIDE THE NAME AND QULAUTIFACTION DETAILS OF ITS TRAINED ENGNEERING STAFF. |  |  |
| S.NO | ITEM SPECIFICARIONS | DESCRIPTION OF SPECS |
| 1) | MAGNET |  |
| 1.1 | TYPE | PERMANENT MAGNET OPEN MRI |
| 1.2 | SHIELDING METHOD | ACTIVE/PASSIVE |
| 1.3 | SHIMMING METHOD | ACTIVE/PASSIVE |
| 1.4 | FIELD STRENGTH | 0.35TESLA OR MORE |
| 1.5 | MAGNET DIAMETER | OPEN ON 3 SIDES OR BETTER |
| 1.6 | MEGNET HOMOGENEITY/STABILITY | 3.0 PPM AT 36 CM DSV OR SIMILAR. |
| 2) | RF SYSTEM |  |
| 2.1 | TYPE | DIGITAL |
| 2.2 | TRANSMITTING POWER | 2.5 KW OR MORE |
| 2.3 | RECEIVER CHANNELS | MINIMUM 04 RF RECEIVE CHANNELS OR MORE |
| 3) | COILS |  |
| 3.1 | NECK COIL |  |
| 3.2 | HEAD COIL |  |
| 3.3 | FLEXIBLE BODY COIL (L) OR COIL TO DO SPINE |  |
| 3.4 | FLEXIBLE BODY COIL (M) |  |
| 3.5 | SHOULDER COIL |  |
| 3.6 | WRIST COIL |  |
| 3.7 | BILATERAL BREAST COIL |  |
| 3.8 | KNEE JOINT COIL |  |
| 4) | GRADIENT SUB SYSTEM |  |
| 4.1 | GRADIENT STRENGTH | 20 MT/M OR MORE, TRUE AXIS VALUES OR SIMILAR |
| 4.2 | SLEW RATE | 50MT/M/S OR MORE, TRUE AXIS VALUES OR SIMILAR |
| 4.3 | FOV | FROM 6MM TO 350MM OR MORE. |
| 5) | RESOLUTION PARAMETEER |  |
| 5.1 | MINIMUM SLICE THICKNESS - 2DFT | 2.0 MM OR BETTER |
| 5.2 | MINIMUM THICKNESS - 3DFT | 0.1 MM OR BETTER |
| 6) | PATIENT COUCH |  |
| 6.1 | TO COVER/SCAN WHOLE BODY | 1.5 METER OR MORE |
| 7) | OPERATING CONSOLE |  |
| 7.1 | STANDARD CONSOLE |  |
| 7.2 | DISPLAY MONITOR | 18” OR MORE HIGH RESOLUTION LCD/TFT MONITOR (NON INTERLACED AND FLICKER FREE) |
| 7.3 | KEY BOARD & MOUSE |  |
| 8) | COMPUTER |  |
| 8.1 | TYPE | MULTITASKING, REAL TIME INTERACTIVE CONTROL |
| 8.2 | PERFORMANCE | SIMULTANEOUS SCANNING, IMAGE ACQUISITION, IMAGAE RECONSTRUCTION, PROCESSING AND FILMING ETC ORCHIVING, NETWORKING ETC. IMAGE PROCESSING SHOULD NOT BE LESS THAN 500 IMAGES/SECOND. |
| 8.3 | HARD DISK CAPABILITY | 30 GB OR MORE |
| 8.4 | DICOM PACKAGE | DICOM 3.0 COMPLIANT |
| 9) | IMAGING TECHNIQUES OR EQUIVALENT |  |
| 9.1 | SPINE ECHO 2D (SE 2D), SPINE ECHO 3D (SE 3D), FAST FIELD ECHO (FFE), T1/T2/N-FAST FIELD ECHO (FFE), B-FAST FIELD ECHO (B-FFE), T1-FAST FIELD ECHO 3D, N/B-FAST FIELD ECHO 3D, TURBO SPINE ECHO (TSE), INVERSION RECOVERY (IR), IRSE, FAST IR, FLAIR, FAST FLAIR, MRA 2D (TOF), MRA 3D(TOF), ECHO PLANER IMAGING (SE-EPI), DIFFUSION WEIGHTED ECHO PLANER IMAGING (DW-EPI), DIFFUSION WEIGHTED IMAGING BASED ON SPINE ECHO (DWISE), DUAL ECHO (DE), DUAL SPIN ECHO (DSE), DRIVE2D: FAST RECOVERY BASED TSE2D, DRIVE3D: FAST RECOVERY BASED TSE3D, ACQUISITION MODE: SINGLE SLICE (2DFT), MULTISLICE (2D FT), MULTISLICE EXCITATION SIMULTANEOUSLY (2/4), VOLUME (3DFT), MULTIVOLUME (3DFT), FAT/WATER SEPARATION, MULTIPLANER RECONSTRUCTION (MRP), REAL TIME ZOOM, SELECTION OF IMAGE FILTERS, ROI ANALYSIS TOOLS, BATCH MODE PROCESSING, ANALYSIS OF CONTRAST DYNAMIC EXAMINATIONS, ADC MAP CALCULATION INCLUDING COMBINED ADC MAP, ISOTROPIC DW IMAGE CALCULATION, MIP. |  |
| 9.2 | CARDIAC GATING |  |
| 9.3 | RESPIRATORY GATING |  |
| 10) | ADVANCED SOFTWARE/APPLICATIONS: THE SYSTEM SHOULD OFFER A LARGE SELECTION OF ROUTINE AND ADVANCED APPLICATIONS. - NEURO - SPINE - BODY - MUSCULOSKELEETAL - VASCULAR - MAMMOGRAPHY - KINEMATIC JOINTS |  |
| 11) | MRI SUITE standard and safety |  |
| 11.1 | VENDOR WILL BE RESPONSIBLE FOR RENOVATION OF MRI SUITE AS PER INTERNATIONAL STANDARDS INCLUDING FURNITURE, AIR CONDITIONING AND ATLEAST FOUR CCTV CAMERAS TO COVER ALL AREAS WITH A DVR TO HAVE ATLEAST THREE MONTHS RECORDING OF ALL FOUR CAMERAS. CCTV SHOULD ALSO HAVE UPS SUPPLY. |  |
| 12) | ACCESSORIES TO BE OFFERED WITH SYSTEM: |  |
| 12.1 | ONLINE SINE WAVE UPS FOR WHOLE SYSTEM WITH A MINIMUM BACK UP OF 10-MINUTES ON FULL LOAD |  |
| 12.2 | RF CAGE/SHIELDING WITH INTERIOR DÉCOR OUTLETS OF OXYGEN/AND SUCTION WITH NON-MAGNET SUCTION UNIT AND OXYGEN FLOW METER INLCLUDING ELECTRICAL AND AIR CONDITIONING NEED ACCORDING TO SYSTEM REQUIREMENT |  |
| 12.3 | LATEST DRY/LASER IMAGER DICOM COMPATIBLE UP TO 14 X 17 INCH FILM SIZE |  |
| 12.4 | NON-FERROUS PATIENT TROLLEY WITH I/V STAND |  |
| 12.5 | NON-FERROUS WHEEL CHAIR WITH IV POLE |  |
| 12.6 | NON-FERROUS HAND HELD METAL DETECTOR |  |
| 12.7 | NON-FERROUS FIRE EXTINGUISHER |  |
| 13) | POWER REQUIREMENT |  |
| 13.1 | SINGLE THREE PHASE LINE VOLTAGE OF 220V, 50HZ |  |
| 14) | TRAINING |  |
| 14.1 | ONE WEEK FACTORY TRAINING OF ENGINEER,TRAINIG OF RADIOLOGIST AT A REPUTABLE CENTRE AND FOR TWO WEEKS ON SITE AVAILABILITY OF AN APPLICATION SPECIALIST FOR TECHNOLOGIST TRAINING. |  |
| 15) | WARRANTY & POST WARRANTY AGREEMENT & MAINTENANCE CONTRACT |  |
| 15.1 | WARRANTY PERIOD SHALL BE OF FIVE YEARS STARTING FROM DATE OF FULL FUNCTIONAL COMMISSIONING |  |
| 15.2 | MAINTENANCE OF THE MACHINE/EQUIPMENT DURING THE WHOLE PERIOD OF WARRANTY SHALL BE RESPONSIBILITY OF THE VENDOR FAILING WHICH, THE VENDOR WILL HAVE TO PAY DOWN TIME PENALTY AS INDICATED IN TERMS AND CONDITIONS |  |
| 15.3 | AN ANNUAL OPTIMAL UPTIME OF 95% IS CONSIDERED AS AN ACCEPTABLE LEVEL OF PERFORMANCE. UPTIME SHALL BE DEFINED AS TIME AVAILABLE TO USER FOR DATA ACQUISITION AND PROCESSING ON ALL WORKING DAYS THROUGHOUT THE YEAR |  |
| 15.4 | IF THE UPTIME PERCENTAGE FOR THE MEASUREMENT PERIOD SHALL FALL SHORT OF 95% ANNUALLY, THE FOLLOWING FORMULA WILL BE APPLIED TO DETERMINE ADDITIONAL DAYS IN THE WARRANTY/SERVICE CONTRACT PERIOD - 100 - 96% NO PENALTY - 95 - 90% THE WARRANTY PERIOD WILL BE EXTENDED BY 01 TIMES THE NUMBER OF DAYS AS EXTRA DOWN TIME. - 89-80% THE WARRANTY PERIOD WILL BE EXTENDED BY 02 TIMES THE NUMBER OF DAYS AS EXTRA DOWN TIME. - BELOW 80% THE WARRANTY PERIOD WILL BE EXTENDED BY 03 TIMES THE NUMBER OF DAYS AS EXTRA DOWN TIME. |  |
| 15.5 | DOWN TIME IS DEFINED AS THE FAILURE IN THE EQUIPMENT OPERATION TO ACQUIRE OR PROCESS THE DATA, RESULTING IN INABILITY TO CARRY OUT THE REQUIRED PROCEDURE PROPERLY. |  |
| 15.6 | DOWN TIME WILL START AS THE END USER NOTIFIED THE DESIGNATED SERVICE/FACILITY VERBALLY OR IN WRITING. THE RESPONSE TIME OF 12 TO 24 HOURS WILL BE CONSIDERED AS NORMAL DEPENDING UPON THE LOCATION OF UNIT. |  |
| 15.7 | DOWN TIME WILL END WHEN SATISFACTORY REPAIRS HAVE BEEN AFFECTED AND THE SYSTEM IS AGAIN FULLY FUNCTIONAL WITH ALL PARAMETERS AVAILABLE FOR CLINICAL USE. |  |
| 15.8 | THE SCHEDULED PREVENTIVE MAINTENANCE SHALL NOT BE CONSIDERED AS DOWN TIME AND SHALL NOT BE MORE THAN THREE DAYS |  |
| 15.9 | MANUFACTURER/SUPPLIER WILL BE RESPONSIBLE FOR PREVENTIVE MAINTENANCE OF THE SYSTEM AS PER SERVICE MANUAL. |  |
| 15.10 | VENDOR WILL MENTION THE RATE OF SERVICE CHARGES WITH AND WITHOUT PARTS FOR THE COMPLETE SYSTEM AFTER THE EXPIRY OF FIVE YEARS WARRANTY PERIOD ON YEARLY BASIS ALONG WITH THE BID. THEY SHOULD ALSO INCLUDE PRICE MAJOR BREAKABLE PARTS. |  |
| 15.11 | MANUFACTURER/SUPPLIER WILL GUARANTEE THE AVAILABILITY OF SPARE PARTS AND ACCESSORIES FOR THE SYSTEM FOR AT LEAST TEN YEARS. |  |
| 15.12 | MANUFACTURER/SUPPLIER SHOULD ALSO CONFIRM THAT IN EVENT OF THE SERVICES COSTS FAILING OR REDUCING IN LINE WITH THE MANUFACTURER POLICY WORLDWIDE DUE TO TECHNICAL INNOVATIONS OR OTHER REVIEWS, THEY WILL REDUCE THE MAINTENANCE COST ACCORDINGLY. |  |
| 16) | OPTIONAL ITEMS - IF AVAILABLE WITH VENDOR (NON-AVAILABILITY WILL NOT BE THE BASIS FOR TECHNICAL ACCEPTANCE/REJECTION) | MENTION SEPARATE PRICE OF EACH COMPONENT |
| 16.1 | CIRCULAR COILS (MAXIMUM AVAILABLE SIZE) |  |
| 16.2 | FLEXIBLE BODY COIL (MAXIMUM AVAILABLE SIZE) |  |
| 16.3 | NON FERROUS ANESTHESIA MACHINE |  |
| 16.4 | CONNECTIVITY WITH LOCAL HMIS/RIS SYSTEMS |  |
| 16.5 | TWO DESKTOP COMPUTERS WITH 64 BIT DUAL PROCESSOR OF LATEST TECHNOLOGY,02MB CACHE RAM AT 01 GHZ, HD 100GB OR MORE, DVD/CD RW, 10/100/1000 BASE T ETHERNET LAN, 56K MODEM, GRAPHICS CARE, NETWORKING CARD, 17” TFT LCD DISPLAY WITH OTHER STANDARD ACCESSORIES. OPERATING SYSTEM SOFTWARE WITH CT VIEWER DICOM COMPATIBLE SOFTWARE SHOULD BE INCLUDED. VENDOR WILL INSTALL/NETWORK THESE COMPUTERS FROM CT SCAN ROOMS TO CONSULTANT’S ROOMS. |  |
| 16.6 | TWO LASER BLACK & WHITE PRINTERS, CAPABLE OF PRINTING ON MULTI-SIZE PAPERS, 1200 DPI OR BETTER. |  |
| 16.7 | TWO LED/LCD FILM VIEWERS/ILLUMINATORS, 14 X 17” SIZE, WITH VARIABLE LIGHT CONTROL. |  |
| 16.8 | MANUFACTURER AND SUPPLIER SHALL BE RESPONSIBLE FOR RECTIFYING ALL PROBLEMS AT THEIR OWN EXPENSE AND MANAGE ANY DEFECT OR FAULT IN THE SYSTEM DEVELOPING AT ANY TIME DURING INSTALLATION AND COMMISSIONING PERIOD. |  |
| 17) | ELECTRICAL REQUIREMENT OF M.R.I SITES |  |
| 17.1 | PROVISION OF DB IN MRI SUITE WILL BE RESPONSIBILITY OF THE VENDOR |  |
| 17.2 | M.R.I SUITE AIR-CONDITIONERS IN THE EQUIPMENT ROOM, AND THE CONSOLE ROOM THE VENDOR WILL BE RESPONSIBLE FOR COMPLETE FINISHING OF THE MR SUITE INCLUDES THE AIR-CONDITIONING AND ALL WIRING/WALL SHIELDING AND NECESSARY FURNITURE MRI SUITE. |  |
| 18.0 | OTHER TERMS AND CONDITIONS FOR MRI SCANNER |  |
| 18.1 | TENDER IS INVITED AS PER SPP RULE 46 (2) -D TWO-BID ENVELOPE SYSTEM |  |
| 18.2 | FOR TECHNICAL BID EACH VENDOR IS REQUESTED TO FILL UP THE ABOVE MENTIONED PERFORMA WITH THE DESIRED INFORMATION IN THE LAST COLUMN, ON THE SAME MS WORD FORMAT IN YES OR NO FORM WITH VERIFIABLE REFERENCES (ALONG WITH PROPER PAGINATION, PARAGRAPH & LINE DESCRIPTION) |  |
| 18.3 | VENDOR MUST ALSO ENCLOSE A SOFT COPY ON CD/USB IN ADDITION TO HARD COPY FOR TECHNICAL EVALUATION COMMITTEE |  |
| 18.4 | ANY WRONG INFORMATION/FIGURE OR HIDING OF FACTS WILL LEAD TO DISQUALIFICATION FROM THE VENDOR. |  |
| 18.5 | ONLY SOLE DISTRIBUTORS IN PAKISTAN OR MANUFACTURER ARE ELIGIBLE TO PARTICIPATE |  |
| 18.6 | PRICING OF OPTIONAL CONFIGURATION AND POST WARRANTY TERMS SHOULD BE QUOTED SEPARATELY. |  |
| 18.7 | PRICING OF OPTIONAL CONFIGURATION WITH ACCESSORIES AND SERVICES SHULD BE QUOTED SEPARATELY. |  |
| 18.8 | TECHNICAL EVALUATION WILL BE BASED ON THE STANDARD CONFIGURATION WITH ACCESSOREIS OF THE WHOLE UNIT. |  |
| 18.9 | COMMERCIAL EVALUATION WILL BE BASED ON THE STANDARD CONFIGURATION WITH ACCESSOREIS OF THE WHOLE UNIT. |  |
| 18.10 | COST ON CIF BASIS LC WILL BE OPENED BY THE VENDOR THE CHARGES OF THE LC, INSURANCE WILL ALSO BE BORNE BY THE VENDOR. |  |
| 18.11 | COST ON F.O.R BASIS |  |
| 19 | QUALITY & SAFETY STANDARDS: |  |
| 19.1 | COUNTRY OF ORIGIN SHOULD BE USA/EUROPE/JAPAN WITH THEIR RESPECTIVE STANDARDS MDD (CE) COMPLIANCE / MHLW JAPAN / FDA 510K. IN CASE EQUIPMENT ORIGIN IS FROM OTHER COUNTRY THAN FDA 510K USA APPROVAL IS MANDATORY. |  |
| CIVIL AND ELECTRICAL WORK FOR MRI SCANNER INSTALLATION AND COMMISSIONING THE CONTRACTING FIRM SHALL VISIT THE SITE BEFORE SUBMISION OF TNEDER TO ASSES THE NOISE LEVELS OF THE SITE FOR MAGNET OF MRI AND ALSO TO ENSURE THAT COST OF ALL NECESSARY ELECTRICAL BOARDS/CABLES/EARTHLING NEEDED TO INSTALL AND COMMISION THE COMPLETE SYSTEM ARE INCLUDED IN THE SUPPLY SCAN IS FIXED WITHIN SIX MONTHS FROM THE DATE OF AWARDING OF THE TENDER. |  |  |

TECHNICAL SPECIFICATION OF CT SIMULATOR

| TECHNICAL SPECIFICATIONS OF 16 SLICE OR MORE CT SIMULATOR |  |
| --- | --- |
| LATEST GENERATION WHOLE BODY 16 SLICES OR MORE CT SIMULATOR WITH CLINICAL APPLICATION SOFTWARE. THE QUOTED MODEL MUST BE THE LATEST AND SUCCESSFUL SIMULATOR HAVING LARGE INSTALLED BASE AROUND THE WORLD. THE BIDDER MUST HAVE SOLD OUT EQUIPMENT IN THE SINDH / PAKISTAN AND INDICATE THE NUMBER OF CT SIMULATOR INSTALLED WITHIN PAKISTAN AND LETTER OF SATISFACTION FROM ATLEAST THREE HOPISTALS. COMPANY SHOULD PROVIDE THE NAME AND QULAUTIFACTION DETAILS OF ITS TRAINED ENGNEERING STAFF. |  |
| S.NO. | ITEM SPECIFICATIONS. |
| 1) | GANTRY |
| 1.1 | GANTRY BORE TO BE AT LEAST 80cm OR MORE |
| 1.2 | THREE DIMENSIONAL MOVING LASER POSITIONING LIGHTS ALONG WITH CONTROLLER WITH COMPATIBILTY CT HARDWARE/SOFTWARE |
| 1.3 | GANTRY ROTATION SPEED TO BE LSS THAN 0.8 SEC FOR GENERATION OF 16-SLICES PER 360 DEGREE ROTATIONS OR BETTER. |
| 2) | TUBE |
| 2.1 | HEAT STORAGE CAPACITY OF 7.5 MHU OR MORE |
| 2.2 | ANODE HEAT DISSIPATION OF AT LEAST 700 KHU OR ABOVE. |
| 3) | GENERATOR & DETECTOR |
| 3.1 | HIGH FREQUENCY TYPE MAXIMUM POWER OF 50 KW OR MORE |
| 3.2 | 80 TO 135 KVp RANGE OR SIMILAR |
| 3.3 | DOSE REDUCTION HARDWARE/SOFT WARE WITH AUTOMATIC MA ADJUSTMENT ACCORDING TO BODY WEIGHT. |
| 3.5 | CALCULATE PATIENT DOSE IN MILLI-GRAY PREFERABLY BEFORE AXIAL ACQUISITION. |
| 3.6 | LOW CONTRAST DETECTABILITY (LCD) CALCULATED ON A CATPHAN CT PHANTOM OF 5MM RESOLUTION WITH A CT NUMBER OF 5HU (0.3%) CONTRAST DIFFERENCE OR BETTER |
| 3.7 | SCAN LENGTH OF AT LEAST 1.5 METERS IN A SINGLE ACQUISITION. |
| 3.8 | SPATIAL RESOLUTION 16 LP/CM OR BETTER AT CUT OFF |
| 4) | TABLE |
| 4.1 | The scanning couch top must be carbon fiber flat table top with minimum dimensions of 235 cmX40cm |
| 4.2 | SINGLE ACQUISITION SCANS RANGE OF AT LEAST 1.5 METER. |
| 4.3 | WEIGHT BEARING CAPACITY OF 180 KG OR MORE |
| 5) | CONSOLE COMPUTER |
| 5.1 | SYSTEM ARCHITECTURE AND OPERATING SYSTEM MUST BE BASED ON LATEST CAPACITY |
| 5.2 | MULTITASKING AND PARALLEL PROCESSING CPU SYSTEM, DUAL XEON PROCESSOR OR BETTER |
| 5.3 | ATLEAST 100GB OF STORAGE SPACE OR MORE/ 50,000 IMAGE STORAGE |
| 5.4 | REAL TIME SPIRAL/HELICAL RECONSTRUCTION OF 6 IMAGES PER SECOND OR BETTER AT 512 X 512 MATRIXES. |
| 5.5 | CONSOLE MONITOR COLOR 19” OR MORE TFT TYPE |
| 5.6 | ITERATIVE DOSE REDUCTION TECHNOLOGY SOFTWARE/HARDWARE TO BE PROVIDED IN ADDITION TO OTHER DOSE REDUCTION TECHNIQUES. (like AIDR / CARE / I-DOSE/ ASIR or equivalent) |
| 5.7 | DVD/CDR |
| 6) | ADDITIONAL MULTI-MODALITY WORK STATION: (AS RECOMMENDED BY MANUFACTURER AND SUPPLIED WORLD-WIDE BY THEM) |
| 6.1 | HIGH-SPEED LINK TO OPERATE CONSOLE ON DICOM NETWORK |
| 6.2 | SYSTEM ARCHITECTURE AND OPERATING SYSTEM MUST BE BASED ON LATEST DUAL XEON PROCESSOR OR BIT RISC. |
| 6.3 | SHOULD HAVE HI-RESOLUTION LCD/TFT SCREEN OF 18-INCHES OR MORE. |
| 6.4 | CD/DVD WRITER |
| 6.5 | DICOM VIEWER FOR VIEWING CT IMAGES |
| 7) | SOFTWARE (INCLUDED AS STANDARD) |
| 7.1 | FULL COLOR VOLUME RENDERING 3D AT CONSOLE AND WORKSTATION. |
| 7.2 | ISOTROPIC VOLUME ACQUISITION AT CONSOLE. |
| 7.3 | CT ANGIO BOTH AT CONSOLE AND WORKSTATION. |
| 7.4 | VESSEL STENOSIS ANALYSIS AT BOTH CONSOLE AND WORKSTATION. |
| 7.5 | 3D SURFACE RENDERING AT BOTH CONSOLE AND WORKSTATION. |
| 7.6 | CURVED PLANER REFORMATION AT CONSOLE AND WORKSTATION. |
| 7.7 | CONTRAST MEDIA BASED SYNCHRONIZING SOFT WARE AT CONSOLE. |
| 7.8 | VIRTUAL ENDOSCOPY / COLONOSCOPY / BRONCHOSCOPY AT WORKSTATION AND CONSOLE. |
| 7.9 | CEREBRAL BLOOD FLOW / HEAD PERFUSION AT CONSOLE AND WORK STATION. |
| 7.10 | LUNG VOLUME ANALYSIS SOFTWARE ON CONSOLE AND WORKSTATION. |
| 7.11 | PEDIATRIC SCANNING PACKAGE WITH SMALL FOV AND LOW KVP SCANNING |
| 7.12 | FLY THROUGH SOFTWARE AT CONSOLE AND WORKSTATION |
| 7.13 | CT FLUOROSCOPY WITH REAL TIME IMAGING AND DISPLAY WITH REQUIRED HARD WARE/ SOFT WARE COMPLETE WITH ONE HIGH-RESOLUTION MONITOR IN CT ROOM ON MOBILE BASE/CEILING SUSPENDED. |
| NOTE | CONSOLE AND WORK STATION SHOULD BE CAPABLE OF INDEPENDENT WORKING. |
| 8) | DICOM 3 CAPABILITY |
| 8.1 | DICOM 3 CAPABILITY FOR SEND, RECEIVE, ARCHIVE, RETRIEVE AND PRINT |
| 9) | POWER REQUIREMENT |
| 9.1 | THREE PHASE, WITH LINE VOLTAGE OF 220V, 50 HZ. |
| 10) | ACCESSORIES (INCLUDED AS STANDARD) |
| 10.1 | DICOM COMPATIBLE DRY LASER IMAGER FOR FILM SIZE 14” X 17”. |
| 10.2 | ONLINE SINE WAVE UPS FOR WHOLE SYSTEM WITH A MINIMUM BACKUP OF 10 MINUTES ON FULL LOAD. |
| 10.3 | DUAL HEAD PROGRAMMABLE, POWER INJECTOR WITH FLOW/VOLUME AND TEMPERATURE CONTROL FOR BOTH HEAD INDEPENDENTLY, MOUNTED ON MOBILE BASE/CEILING SUSPENDED WITH 100 SYRINGES OF 150 ML CAPACITY AND CONNECTING TUBES. |
| 10.4 | STANDARD SET OF PHANTOMS FOR CALIBRATION OF CT. |
| 10.5 | TABLE ACCESSORIES: TABLE PAD, ARM REST, PATIENT RESTRAINT KIT, IV POLE, INFANT CRADLE, FLAT HEAD HOLDER |
| 10.6 | TWO DESKTOP COMPUTERS WITH 64BIT DUAL PROCESSORS OF LATEST TECHNOLOGY, 02MB CACHE RAM AT 01 GHZ, HD 100GB OR MORE, DVD/CD RW, 10/100/1000 BASE T ETHERNET LAN, 56K MODEM, GRAPHICS CARD, NETWORKING CARD, 17” TFT LCD DISPLAY WITH OTHER STANDARD ACCESSORIES. OPERATING SYSTEM SOFTWARE WITH CT VIEWER DICOM COMPATIBLE SOFTWARE SHOULD BE INCLUDED. VENDOR WILL INSTALL / NETWORK THESE COMPUTERS FROM CT SCAN ROOMS TO CONSULTANT’S ROOMS. |
| 10.7 | TWO LASER BLACK & WHITE PRINTERS, CAPABLE OF PRINTING ON MULTI-SIZE PAPERS, 1200 DPI OR BETTER. |
| 10.8 | TWO FILM VIEWERS/ ILLUMINATORS, 14”X 17” SIZE, WITH VARIABLE LIGHT CONTROL, 4X1 FORMATS. |
| 10.9 | MANUFACTURER AND SUPPLIER SHALL BE RESPONSIBLE FOR RECTIFYING ALL PROBLEMS AT THEIR OWN EXPENSE AND MANAGE ANY DEFECT OR FAULT IN THE SYSTEM DEVELOPING AT ANY TIME DURING INSTALLATION AND COMMISSIONING PERIOD. |
| 11) | TRAINING |
| 11.2 | ONE WEEK FACTORY TRAINING OF ENGINEER AND RADIOLOGIST ABROAD AT A REUTABLE CENTRE AND TECHNOLOGISTS TRAINING BY AN APPLICATION SPECIALIST FOR TWO WEEK AT SITE. |
| 12) | WARRANTY & POST WARRANTY AGREEMENT & MAINTENANCE CONTRACT |
| 12.1 | WARRANTY PERIOD SHALL BE OF THREE YEARS STARTING FROM DATE OF FULL FUNCTIONAL COMMISSIONING INCLUDING CT SCAN TUBE (UNLIMITED EXPOSURES IN THREE YEARS) |
| 12.2 | MAINTENANCE OF THE MACHINE/EQUIPMENT DURING THE WHOLE PERIOD OF WARRANTY SHALL BE RESPONSIBILITY OF THE VENDOR FAILING WHICH, THE VENDOR WILL HAVE TO PAY DOWN TIME PENALTY AS INDICATED IN TERMS AND CONDITIONS. |
| 12.3 | AN ANNUAL OPTIMAL UPTIME OF 95% IS CONSIDERED AS AN ACCEPTABLE LEVEL OF PERFORMANCE. UPTIME SHALL BE DEFINED AS TIME AVAILABLE TO USER FOR DATA ACQUISITION AND PROCESSING ON ALL WORKING DAYS THROUGHOUT THE YEAR. |
| 12.4 | IF THE UPTIME PERCENTAGE FOR THE MEASUREMENT PERIOD SHALL FALL SHORT OF 95% ANNUALLY, THE FOLLOWING FORMULA WILL BE APPLIED TO DETERMINE ADDITIONAL DAYS IN THE WARRANTY/SERVICE CONTRACT PERIOD. - 100- 96% NO PENALTY - 95-90% THE WARRANTY PERIOD WILL BE EXTENDED BY 01 TIMES THE NUMBER OF DAYS AS EXTRA DOWN TIME. - 89-80% THE WARRANTY PERIOD WILL BE EXTENDED BY 02 TIMES THE NUMBER OF DAYS AS EXTRA DOWN TIME. - BELOW 80% THE WARRANTY PERIOD WILL BE EXTENDED BY 03 TIMES THE NUMBER OF DAYS AS EXTRA DOWN TIME. |
| 12.5 | DOWN TIME IS DEFINED AS THE FAILURE THE EQUIPMENT OPERATION TO ACQUIRE OR PROCESS THE DATA, RESULTING IN INABILITY TO CARRY OUT THE REQUIRED PROCEDURE PROPERLY. |
| 12.6 | DOWN TIME WILL START AS THE END USER NOTIFIES THE DESIGNATED SERVICE/FACILITY VERBALLY OR IN WRITING. THE RESPONSE TIME OF 12 TO 24 HOURS WILL BE CONSIDERED AS NORMAL DEPENDING UPON THE LOCATION OF UNIT. |
| 12.7 | DOWN TIME WILL END WHEN SATISFACTORY REPAIRS HAVE BEEN AFFECTED AND THE SYSTEM IS AGAIN FULLY FUNCTIONAL WITH ALL PARAMETERS AVAILABLE FOR CLINICAL USE. |
| 12.8 | THE SCHEDULED PREVENTIVE MAINTENANCE SHALL NOT BE CONSIDERED AS DOWN TIME AND SHALL NOT BE MORE THAN TWO DAYS. |
| 12.9 | MANUFACTURER /SUPPLIER WILL BE RESPONSIBLE FOR PREVENTIVE MAINTENANCE OF THE SYSTEM AS PER SERVICE MANUAL SHOULD KEEP A CHECK ON ELECTRICAL, MAGNETIC, TEMPERATURE AND HUMIDITY CONDITIONS. SUCH A CHECK SHOULD BE MADE EVERY FOUR WEEKS AND RECORD SHOULD BE PROVIDED TO THE RADIOLOGY DEPARTMENT. |
| 12.10 | VENDOR WILL MENTION THE RATE OF SERVICE CHARGES WITH AND WITHOUT PARTS FOR THE COMPLETE SYSTEM, AFTER THE EXPIRY OF FIVE YEARS WARRANTY PERIOD ON YEARLY BASIS ALONG WITH THE BID. THIS SHOULD ALSO MENTION SEPRATE PRICE OF CT SCAN TUBE |
| 12.11 | MANUFACUTRER/SUPPLIER WILL GUARANTEE THE AVAILABILITY OF SPARE PARTS AND ACCESSORIES FOR THE SYSTEM FOR AT LEAST TEN YEARS. |
| 12.12 | MANUFACTURER/SUPPLIER SHOULD ALSO CONFIRM THAT IN EVENT OF THE SERVICES COSTS FAILING OR REDUCING IN LINE WITH THE MANUFACTURER POLICY WORLDWIDE DUE TO TECHNICAL INNOVATIONS OR OTHER REVIEWS, THEY WILL REDUCE THE MAINTENANCE COST ACCORDINGLY. |
| 13) | INSTALLATION AND COMMISSIONING |
| 13.1 | THE CONTRACTING FIRM WILL INSTALL COMPLETE SYSTEM INCLUDING ALL ACCESSOREIS WITH SIX MONTHS STARTING FROM DATE OF ESTABLISHING L/C |
| 14) | RADIATION PROTECTION |
| 14.1 | PROTECTION DEVICES (LEAD APRONS 04 WITH HANGERS, LEAD GLOVES 04 PAIRS, 0.5MM PB EQUIVALENT LIGHT WEIGHT) |
| 14.2 | LEAD GLASS FOR CONTROL ROOM 5X3 FEET 0.5MM PB.EQUIVALENT |
| 15. | MANUFACTUER AND SUPPLIER SHALL BE RESPONSIBLE FOR RECTIFYING ALL PROBLEMS AT THEIR OWN EXPENSE AND MANAGE ANY DEFECT OR FAULT IN THE SYSTEM DEVELOPING AT ANY TIME DURING ISNTLALATION AND COMMISSIONING PERIOD. |
| 16) | OTHER TERMS AND CONDITIONS FOR CT SIMULATOR |
| 16.1 | TENDER IS INVITED AS PER SPP RULE 46 (2)-d TWO-BID TWO ENVELOPE SYSTEM |
| 16.2 | FOR TECHNICAL BID EACH VENDOR IS REQUESTED TO FILL UP THE ABOVE MENTIONED PERFORMA WITH THE DESIRED INFORMATION IN THE LAST COLUMN, ON THE SAME MS WORD FORMAT IN YES OR NO FORM WITH VERIFIABLE REFERENCES (ALONG WITH PROPER PAGINATION, PARAGRAPH & LINE DESCRIPTION) |
| 16.3 | VENDOR MUST ALSO ENCLOSE A SOFT COPY ON CD/USB IN ADDITION TO HARD COPY FOR TECHNICAL EVALUATION COMMITTEE |
| 16.4 | ANY WRONG INFORMATION/FIGURE OR HIDING OF FACTS WILL LEAD TO DISQUALIFICATION FROM THE VENDOR. |
| 16.5 | ONLY SOLE DISTRIBUTORS IN PAKISTAN OR MANUFACTURER ARE ELIGIBLE TO PARTICIPATE |
| 16.6 | PRICING OF OPTIONAL CONFIGURATION AND POST WARRANTY TERMS SHOULD BE QUOTED SEPARATELY. |
| 16.7 | PRICING OF OPTIONAL CONFIGURATION WITH ACCESSORIES AND SERVICES SHULD BE QUOTED SEPARATELY. |
| 16.8 | TECHNICAL EVALUATION WILL BE BASED ON THE STANDARD CONFIGURATION WITH ACCESSOREIS AND SERVICES OF THE WHOLE UNIT. |
| 16.9 | COMMERCIAL EVALUATION WILL BE BASED ON THE STANDARD CONFIGURATION WITH ACCESSOREIS OF THE WHOLE UNIT. |
|  | COST ON CIF BASIS LC WILL BE OPENED BY THE VENDOR THE CHARGES OF THE LC, INSURANCE WILL ALSO BE BORN BY THE VENDOR. |
|  | COST ON F.O.R BASIS |
| 17) | QUALITY & SAFETY STANDARDS: |
| 17.1 | COUNTRY OF ORIGIN SHOULD BE USA/ EUROPE/JAPAN WITH THEIR RESPECTIVE STANDARDS MDD (CE) COMPLIANCE / MHLW JAPAN / FDA 510K. |

CIVIL AND ELECTRICAL WORK FOR CT SIMULATOR

INSTALLATION AND COMMISSIONING WILL BE RESPONSIBILITY OF VENDOR:

THE CONTRACTING FIRM SHALL VISIT THE SITE BEFORE SUBMMISION OF TENDER AND ENSURE THAT COST OF ALL NECESSARY ELECTRICAL BOARDS / CABLES / EARTHLING NEEDED TO INSTALL AND COMMISION THE COMPLETE SYSTEM IS INCLUDED IN THE TENDER COST, INCLUDING SUPPLY AND INSTALLATION OF ELECTRICAL DISTRIBUTION BOX FOR THE CT SIM IN THE CT SCAN ROOM ETC.

RENOVATION OF THE CT SIM SUITE FOR THE INSTALLATION AS PER INTERNATIONAL AND PNRA STANDARS INCLUDING FURNITURE AND AIR CONDITIONERS SHALL BE RESPONSIBLY OF VENDOR.

Section VIII.  Sample Forms and Procedures

Table of Sample Forms and Procedures

1. Bid Forms and Price Schedules	43

2. Form of Bid Security / Earnest Money (Bank Guarantee)	50

3. Form of Contract Agreement	51

4. Performance Security Forms	743

5. Form of Completion Certificate	754

6. Form of Operational Acceptance Certificate	765

7. Integrity Pact	56

### 1. Bid Forms and Price Schedules

1.1  Bid Form—Single Stage Bidding

Date:

[Name of Contract]

To:  [Name and address of Purchaser]

Ladies and / or Gentlemen,

Having examined the bidding documents, including Addendum Nos. [insert numbers], the receipt of which is hereby acknowledged, we, the undersigned, offer to design, manufacture, test, deliver, install, pre-commission and commission the Facilities under the above-named Contract in full conformity with the said bidding documents for the sum of:  [amount of foreign currency in words], [amount in figures], and [amount of local currency in words], [amount in figures] or such other sums as may be determined in accordance with the terms and conditions of the Contract.  The above amounts are in accordance with the Price Schedules attached herewith and are made part of this bid.

We undertake, if our bid is accepted, to commence the Facilities and to achieve Completion within the respective times stated in the bidding documents.

If our bid is accepted, we undertake to provide an advance payment security and a performance security in the form, in the amounts, and within the times specified in the bidding documents.

We agree to abide by this bid, which consists of this letter and Attachments 1 through [number] hereto, for a period of [number] days from the date fixed for submission of bids as stipulated in the bidding documents, and it shall remain binding upon us and may be accepted by you at any time before the expiration of that period.

Commissions or gratuities, if any, paid or to be paid by us to agents relating to this Bid, and to contract execution if we are awarded the contract, are listed below

| Name and address of agent |  | Amount and Currency |  | Purpose of Commission or gratuity |
| --- | --- | --- | --- | --- |
| (if none, state “none”). |  |  |  |  |

Until a formal contract is prepared and executed between us, this bid, together with your written acceptance thereof and your notification of award, shall constitute a binding contract between us.

We understand that you are not bound to accept the lowest or any bid you may receive.

Dated this 	 day of 	, 19

[signature]

In the capacity of

[position]

Duly authorized to sign this bid for and on behalf of

[name of Bidder]

Attachments:  [In accordance with ITB Sub-Clause 9.3, plus any additional attachment either specified by the Purchaser in the Bid Data Sheet or included by the Bidder.  In this regard, the Purchaser may include in the bidding document, as an Attachment to the Bid Form, a list of “Bidding Data,” including the most important data to be provided by the Bidder (e.g., Time for Completion, Functional Guarantees, etc.).]

1.3 Price Schedules

In order to receive consistent and responsive bids, it is recommended that Purchasers include a PREAMBLE to the Price Schedules indicating exactly what is required of bidders when completing and pricing their bids.

The following Preamble is given as an example only.  Purchasers are responsible for ensuring that the Preamble included in the bidding documents is complete and appropriate for the contract in question.

PREAMBLE

General

- The Price Schedules are divided into separate Schedules as follows:

Schedule No. 1	Plant and Equipment Supplied from Abroad

Schedule No. 2	Plant and Equipment

Supplied from within the Purchaser’s Country

Schedule No. 3	Local Transportation

Schedule No. 4	Installation Services

(Civil work, Electrical work and other services)

Schedule No. 5	Grand Summary

- The Schedules do not generally give a full description of the plant and equipment to be supplied and the services to be performed under each item.  Bidders shall be deemed to have read the Technical Specifications and other sections of the bidding documents to ascertain the full scope of the requirements included in each item prior to filling in the rates and prices.  The entered rates and prices shall be deemed to include for the full scope as aforesaid, including overheads and profit.

- If bidders are unclear or uncertain as to the scope of any item, they shall seek clarification in accordance with the Instructions to Bidders in the bidding documents prior to submitting their bid.

Pricing

- Prices shall be filled in indelible ink, and any alterations necessary due to errors, etc., shall be initialed by the bidder authorized to sign the bid form

As specified in the Bid Data Sheet and Special Conditions of Contract, prices shall be fixed and firm for the duration of the Contract.

- Bid prices shall be quoted in the manner indicated and in the currencies specified in the Instructions to Bidders in the bidding documents.

For each item, bidders shall complete each appropriate column in the respective Schedules, giving the price breakdown as indicated in the Schedules.

Prices given in the Schedules against each item shall be for the scope covered by that item as detailed in the Technical Specification in the bidding documents

- Where there are errors between the total of the amounts given under the column for the price breakdown and the amount given under the Total Price, the former shall prevail and the latter will be corrected accordingly.

Where there are errors between the total of the amounts of Schedule Nos. 1 to 4 and the amount given in Schedule No. 5 (Grand Summary), the former shall prevail and the latter will be corrected accordingly.

Where there are discrepancies between amounts stated in figures and amounts stated in words, the amounts stated in words shall prevail.

- Items left blank will be deemed to have been included in other items. The TOTAL for each Schedule and the TOTAL of the Grand Summary shall be deemed to be the total price for executing the Facilities and sections thereof in complete accordance with the Contract, whether or not each individual item has been priced.

Schedules of Rates and Prices

Schedule No. 1.  Plant, Equipment, Supplied From Abroad

| Item |  | Description | Code1 |  | Qty. |  |  | Unit Price2 | Total Price | Unit Price | Total Price |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  |  |  | CIF | CIF | DDP | DDP |
|  |  |  |  |  | (1) |  |  | (2) | (1)x(2) |  |  |
|  |  |  |  |  |  |  | Total |  |  |  |  |
| Code | Country |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  | Name of Bidder |  |  |  |  |  |
|  |  |  |  |  |  | Signature of Bidder |  |  |  |  |  |

Schedule No. 2.  Plant, Equipment Supplied from Within the Purchaser’s Country

| Item | Description |  |  | Qty. |  | EXW Unit Price |  | EXW Total Price |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | (1) |  | (2) |  | (1) x (2) |  |
| TOTAL (to Schedule No. 5. Grand Summary) |  |  |  |  |  |  |  |  |  |
|  |  |  | Name of Bidder |  |  |  |  |  |  |
|  |  |  | Signature of Bidder |  |  |  |  |  |  |

Schedule No. 3.  Local Transportation

| Item | Description |  | Qty. | Unit Price |  |  | Total Price |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | Local Currency Portion |  |  |  |
|  |  |  | (1) | (2) |  |  | (1) x (2) |
| TOTAL (to Schedule No. 5. Grand Summary) |  |  |  |  |  |  |  |
|  |  |  |  | Name of Bidder |  |  |  |
|  |  |  |  | Signature of Bidder |  |  |  |

Schedule No. 4.  Installation Services (Civil work, Electric work and other services)

| Item | Description | Qty. | Unit Price |  |  |  | Total Price |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Foreign Currency Portion |  | Local Currency Portion |  | Foreign |  | Local |
|  |  | (1) | (2) |  | (3) |  | (1) x (2) |  | (1) x (3) |
| TOTAL (to Schedule No. 5. Grand Summary) |  |  |  |  |  |  |  |  |  |
|  |  |  | Name of Bidder |  |  |  |  |  |  |
|  |  |  | Signature of Bidder |  |  |  |  |  |  |

Schedule No. 5.  Grand Summary

| Item | Description |  |  |  | Total Price |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | Foreign |  | Local |  |
|  | Total Schedule No. 1. Plant, Equipment, Supplied from Abroad |  |  |  |  |  |  |  |
|  | Total Schedule No. 2. Plant, Equipment Supplied from Within the Purchaser’s Country |  |  |  |  |  |  |  |
|  | Total Schedule No. 3. Local Transportation, |  |  |  |  |  |  |  |
|  | Total Schedule No. 4. Installation Services (Civil work , Electric work, and other services) |  |  |  |  |  |  |  |
| TOTAL (to Bid Form) |  |  |  |  |  |  |  |  |
|  |  |  | Name of Bidder |  |  |  |  |  |
|  |  |  | Signature of Bidder |  |  |  |  |  |

### 2. Form of Bid Security / Earnest Money

_____________________________________________
[Bank’s Name, and Address of Issuing Branch or Office]

Beneficiary:	___________________ [Name and Address of Purchaser]

Date:	________________

BID GUARANTEE No.:	_________________

We have been informed that [name of the Bidder] (hereinafter called "the Bidder") has submitted to you its bid dated (hereinafter called "the Bid") for the execution of [name of contract] under Invitation for Bids.

Furthermore, we understand that, according to your conditions, bids must be supported by a bid guarantee.

At the request of the Bidder, we [name of Bank] hereby irrevocably undertake to pay you any sum or sums not exceeding in total an amount of [amount in figures]  ([amount in words]) upon receipt by us of your first demand in writing accompanied by a written statement stating that the Bidder is in breach of its obligation(s) under the bid conditions, because the Bidder:

(a) 	has withdrawn its Bid during the period of bid validity specified by the Bidder in the Form of Bid; or

(b) 	does not accept the correction of errors in accordance with the Instructions to Bidders.

(c) 	having been notified of the acceptance of its Bid by the Purchaser during the period of bid validity, (i) fails or refuses to execute the Contract Form, if required, or (ii) fails or refuses to furnish the performance security, in accordance with the ITB.

This guarantee will expire:  (a) if the Bidder is the successful bidder, upon our receipt of copies of the contract signed by the Bidder and the performance security issued to you upon the instruction of the Bidder; or (b) if the Bidder is not the successful bidder, upon the earlier of (i) our receipt of a copy of your notification to the Bidder of the name of the successful bidder; or (ii) twenty-eight days after the expiration of the Bidder’s Bid.

Consequently, any demand for payment under this guarantee must be received by us at the office on or before that date.

_______________________

[signature(s)]

### 3. Form of Contract Agreement

THIS CONTRACT AGREEMENT is made the 	 day of 	, 19	.

BETWEEN

(1) [Name of Purchaser], a corporation incorporated under the laws of [country of Purchaser] and having its principal place of business at [address of Purchaser] (hereinafter called “the Purchaser”), and (2) [name of Contractor], a corporation incorporated under the laws of [country of Contractor] and having its principal place of business at [address of Contractor] (hereinafter called “the Contractor”).

WHEREAS the Purchaser desires to engage the Contractor to design, manufacture, test, deliver, install, complete and commission certain Facilities, viz. [list of facilities] (“the Facilities”) and the Contractor have agreed to such engagement upon and subject to the terms and conditions hereinafter appearing.

NOW IT IS HEREBY AGREED as follows:

| Article 1. Contract Documents | 1.1 Contract Documents The following documents shall constitute the Contract between the Purchaser and the Contractor, and each shall be read and construed as an integral part of the Contract: (a) This Contract Agreement and Price Schedule submitted by the contractor. (b) Bid and Price Schedules submitted by the Contractor (c) Special Conditions of Contract (d) General Conditions of Contract (e) Technical Specifications (f) Any other documents shall be added here |
| --- | --- |
| Article 2. Contract Price | 2.1 Contract Price The Purchaser hereby agrees to pay to the Contractor the Contract Price in consideration of the performance by the Contractor of its obligations hereunder. The Contract Price shall be the aggregate of: [amount in words], [amount in figures] as specified in Price Schedule No. 5 (Grand Summary), or such other sums as may be determined in accordance with the terms and conditions of the Contract. |
| Article 3. Effective Date for Determining Time for Completion | - Effective Date (Reference GCC Clause 1) The Time of Completion of the Facilities shall be determined from the date when all of the following conditions have been fulfilled: (a) This Contract Agreement has been duly executed or signed for and on behalf of the Purchaser and the Contractor; (b) The Contractor has submitted to the Purchaser the performance security. Each party shall use its best efforts to fulfill the above conditions for which it is responsible as soon as practicable. |

IN WITNESS WHEREOF the Purchaser and the Contractor have caused this Agreement to be duly executed by their duly authorized representatives the day and year first above written.

Signed by, for and on behalf of the Purchaser

[Signature]

[Title]

in the presence of

Signed by, for and on behalf of the Contractor

[Signature]

[Title]

in the presence of

### 4. Performance Security Forms

Bank Guarantee

________________________________ [Bank’s Name, and Address of Issuing Branch or Office]

Beneficiary:	[Name and Address of Purchaser]

Date:	________________

PERFORMANCE GUARANTEE No.:	_________________

WhereAs [name of Contractor] (hereinafter called "the Contractor") has undertaken in pursuance of  Contract No. [reference number of the contract] dated ____________  to supply [name of contract and brief description of Facilities] (hereinafter called "the Contract").

And Whereas it has been stipulated by you in the said Contract that the Supplier shall furnish you with a bank guarantee by a reputable bank for the sum specified therein as security for compliance with the Supplier’s performance obligations in accordance with the contract.

And WhereAS We have to give the Supplier a guarantee:

Therefore We hereby affirm that we are Guarantors are responsible to you, on behalf of the Supplier , up to a total of [amount of the guarantee in words and figures] and we undertake to pay you , upon your first written demand declaring the supplier to be in default under the contract and without cavil or argument, any some or sums within the limits of [amount of guarantee] as aforesaid with your needing to prove or to show grounds or reason for your demand or the sum specified therein .

This guarantee is valid un till 	the ____ day of _____, 2___

____________________ 
[signature(s) and seal]

____________________________________________________________ 
[name of bank]

____________________________________________________________
[address]

____________________

[date]

### 5. Form of Completion Certificate

Date:

Loan/Credit No:

IFB No:

[Name of Contract]

To:  [Name and address of Contractor]

Dear Ladies and/or Gentlemen,

Pursuant to GCC Clause 20 (Completion of the Facilities) of the General Conditions of the Contract entered into between yourselves and the Purchaser dated [date], relating to the [brief description of the Facilities], we hereby notify you that the following part(s) of the Facilities was (were) complete on the date specified below, and that, in accordance with the terms of the Contract, the Purchaser hereby takes over the said part(s) of the Facilities, together with the responsibility for care and custody and the risk of loss thereof on the date mentioned below.

1.	Description of the Facilities or part thereof:  [description]

2.	Date of Completion:  [date]

However, you are required to complete the outstanding items listed in the attachment hereto as soon as practicable.

This letter does not relieve you of your obligation to complete the execution of the Facilities in accordance with the Contract nor of your obligations during the Defect Liability Period.

Very truly yours,

____________________________________________________________________________

Medical Superintendent

____________________

____________________

### 6. Form of Operational Acceptance Certificate

Date:

Loan/Credit No:

IFB No:

[Name of Contract]

To:  [Name and address of Contractor]

Dear Ladies and/or Gentlemen,

Pursuant to GCC Sub-Clause 21 (Operational Acceptance) of the General Conditions of the Contract entered into between yourselves and the Purchaser dated [date], relating to the [brief description of the Facilities], we hereby notify you that the Functional Guarantees of the following part(s) of the Facilities were satisfactorily attained on the date specified below.

1.	Description of the Facilities or part thereof:  [description]

2.	Date of Operational Acceptance:  [date]

This letter does not relieve you of your obligation to complete the execution of the Facilities in accordance with the Contract nor of your obligations during the Defect Liability Period.

Very truly yours,

____________________________________________________________________________

Medical Superintendent

____________________

____________________

### 7. Integrity Pact

DECLARATION OF FEES, COMMISSIONS AND BROKERAGE ETC. PAYABLE BY THE

SUPPLIERS / CONTRACTORS OF GOODS, SERVICES & WORKS

_________________________ [the Seller/Supplier/Contractor] hereby declares its intention not to

obtain or induce the procurement of any contract, right, interest, privilege or other obligation or benefit from Government of Pakistan or any administrative subdivision or agency thereof or any other entity owned or controlled by it (Government of Pakistan) through  any corrupt business practice.

Without limiting the generality of the foregoing, [the Seller/Supplier/Contractor] represents and warrants that it has fully declared the brokerage, commission, fees etc. paid or payable to

anyone and not given or agreed to give and shall not give or agree to give to anyone within or outside Pakistan either directly or indirectly through any natural or juridical person, including its affiliate, agent, associate, broker, consultant, director, promoter, shareholder, sponsor or subsidiary, any  commission, gratification, bribe, finder’s fee or kickback, whether described as consultation fee or otherwise, with the object of obtaining or including the procurement of contract, right, interest, privilege or other obligation or benefit in whatsoever form from Government of Pakistan, except that which has been expressly declared pursuant hereto.

[The Seller/Supplier/Contractor] certifies that it has made and will make full disclosure of all agreements and arrangements with all persons in respect of or related to the transaction with Government of Pakistan and has not taken any action or will not take any action to circumvent the above declaration, representation or warranty.

[The Seller/Supplier/Contractor] accepts full responsibility and strict liability for making any false declaration, not making full disclosure, misrepresenting facts or taking any action likely to defeat the purpose of this declaration, representation and warranty. It agrees that any contract, right, interest, privilege or other obligation or benefit obtained or procured as aforesaid shall, without prejudice to any other right and remedies available to Government of Pakistan under any law, contract or other instrument, be voidable at the option of Government of Pakistan.

Notwithstanding any rights and remedies exercised by Government of Pakistan in this regard, [the Seller/Supplier/Contractor] agrees to indemnify Government of Pakistan for any loss or damage incurred by it on account of its corrupt business practices and further pay compensation to Government of Pakistan in an amount equivalent to ten time the sum of any commission, gratification, bribe, finder’s fee or kickback given by [the Seller/Supplier/Contractor] as aforesaid for the purpose of obtaining or inducing the procurement of any contract, right, interest, privilege or other  obligation or benefit in whatsoever form from Government of Pakistan.